## Commentary
2020-07-29
Worksheet 'pyUBreadings' requires the user to enter all the dial readings and settings of the UB as well as estimates of the uncertainty in temperature and in the temperature coefficients for resistance and inductance. In general it is set up not to apply any temperature corrections, just to evaluate the uncertainty.
Use the UniversalBridge3 project to process the input data. This uses the latest ubdict.csv file for corrections to the Universal Bridge. Running 'calculate_imp.py' produces a new 'indResults.xlsx'. Paste these results into the 'pyUBresults' worksheet of this spreadsheet and, if correctly pasted, should immediately produce the report in the'pyReport' worksheet. Note that there are two named blocks in 'pyUBresults' and so the new data should just be pasted over the old data in order to preserve the inductor1 and inductor2 named blocks.
You should check that the input data listed in 'pyUBresults' matches the input data in 'pyUBreadings', just in case the wrong file names are enterd in 'calculate_imp.py'.

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | Muirhead A-517 |  | 4: Client; 10: Enter, say, 5 ppm as 5e-6
Date from | 2024-12-16 |  | 4: Job number; 5: 96330136; 10: 1 sigma uncertainties from specifications
to | 2024-12-16 |  | 4: S Number; 5: S22075
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
L0 | 1591.5 | 4Z | 2 | 110 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 4Z | 101476 | 194950 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 4Z | 199848 | 280100 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 4Z | 292588 | 344778 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 4Z | 393378 | 405721 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 4Z | 490373 | 459176 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 4Z | 590778 | 510272 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 4Z | 689189 | 557351 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 4Z | 788515 | 603071 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 4Z | 889520 | 648391 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 4Z | 1010587 | 698940 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L0 | 1591.5 | 5Z | 0 | 0 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 5Z | 103874 | 168179 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 5Z | 207639 | 243857 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 5Z | 304242 | 301271 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 5Z | 398848 | 350041 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 5Z | 494609 | 395561 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 5Z | 592259 | 438801 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 5Z | 689707 | 480181 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 5Z | 782991 | 518311 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 5Z | 878170 | 556590 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 5Z | 984940 | 597390 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
L0 | 1591.5 | 4Z | 2 | 110 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-07 | 0.0012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 4Z | 101476 | 194950 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.01015 | 9.0477e-07 | 2.0133 | 19.4852 | 0.11559 | 2.084 | 1.9072e-07 | 0.0016386 | 4.744 | 70.5404 | L11591.5 | 1000 | 19485.1611 | 115.585 | 2.084 | 1000000 | 10149.5029 | 0.90477 | 2.0133 | 1591.5
L2 | 1591.5 | 4Z | 199848 | 280100 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.019989 | 1.6328e-06 | 2.0348 | 28.0007 | 0.16609 | 2.084 | 3.0806e-07 | 0.0023937 | 5.3003 | 69.3862 | L21591.5 | 1000 | 28000.6686 | 166.0868 | 2.084 | 1000000 | 19988.7416 | 1.6328 | 2.0348 | 1591.5
L3 | 1591.5 | 4Z | 292588 | 344778 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.029265 | 2.3489e-06 | 2.0412 | 34.4689 | 0.20448 | 2.0838 | 4.2994e-07 | 0.0032214 | 5.4632 | 63.476 | L31591.5 | 1000 | 34468.854 | 204.4829 | 2.0838 | 1000000 | 29264.6635 | 2.3489 | 2.0412 | 1591.5
L4 | 1591.5 | 4Z | 393378 | 405721 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.039346 | 3.1359e-06 | 2.0441 | 40.5635 | 0.24069 | 2.0837 | 5.6646e-07 | 0.0041694 | 5.5359 | 57.7278 | L41591.5 | 1000 | 40563.5172 | 240.6906 | 2.0837 | 1000000 | 39345.7523 | 3.1359 | 2.0441 | 1591.5
L5 | 1591.5 | 4Z | 490373 | 459176 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.049047 | 3.8968e-06 | 2.0454 | 45.9093 | 0.27247 | 2.0835 | 6.9958e-07 | 0.0051039 | 5.5701 | 53.3845 | L51591.5 | 1000 | 45909.3357 | 272.4707 | 2.0835 | 1000000 | 49047.2624 | 3.8968 | 2.0454 | 1591.5
L6 | 1591.5 | 4Z | 590778 | 510272 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.05909 | 4.6862e-06 | 2.0462 | 51.0192 | 0.30287 | 2.0833 | 8.3831e-07 | 0.0060833 | 5.59 | 49.7866 | L61591.5 | 1000 | 51019.2402 | 302.8683 | 2.0833 | 1000000 | 59089.8432 | 4.6862 | 2.0462 | 1591.5
L7 | 1591.5 | 4Z | 689189 | 557351 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.068933 | 5.461e-06 | 2.0466 | 55.7274 | 0.33089 | 2.0832 | 9.7481e-07 | 0.0070503 | 5.6021 | 46.933 | L71591.5 | 1000 | 55727.4208 | 330.8926 | 2.0832 | 1000000 | 68932.9827 | 5.461 | 2.0466 | 1591.5
L8 | 1591.5 | 4Z | 788515 | 603071 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.078868 | 6.2436e-06 | 2.0469 | 60.2997 | 0.35812 | 2.083 | 1.1129e-06 | 0.0080308 | 5.6101 | 44.5931 | L81591.5 | 1000 | 60299.6932 | 358.1205 | 2.083 | 1000000 | 78867.6411 | 6.2436 | 2.0469 | 1591.5
L9 | 1591.5 | 4Z | 889520 | 648391 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.08897 | 7.0399e-06 | 2.0472 | 64.832 | 0.38512 | 2.0829 | 1.2536e-06 | 0.0090312 | 5.6158 | 42.6435 | L91591.5 | 1000 | 64831.9633 | 385.12 | 2.0829 | 1000000 | 88970.2343 | 7.0399 | 2.0472 | 1591.5
L10 | 1591.5 | 4Z | 1010587 | 698940 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10108 | 7.9947e-06 | 2.0473 | 69.8872 | 0.41526 | 2.0827 | 1.4224e-06 | 0.010233 | 5.6206 | 40.5814 | L101591.5 | 1000 | 69887.1646 | 415.2582 | 2.0827 | 1000000 | 101079.4433 | 7.9947 | 2.0473 | 1591.5
L0 | 1591.5 | 5Z | 0 | 0 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-06 | 0.012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 5Z | 103874 | 168179 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.1039 | 9.209e-06 | 2.0143 | 168.189 | 0.99775 | 2.0837 | 1.9298e-06 | 0.016478 | 4.7719 | 60.552 | L11591.5 | 1 | 168.189 | 0.99775 | 2.0837 | 1000000 | 103895.5268 | 9.209 | 2.0143 | 1591.5
L2 | 1591.5 | 5Z | 207639 | 243857 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.20768 | 1.6917e-05 | 2.0358 | 243.8715 | 1.4468 | 2.0837 | 3.177e-06 | 0.024515 | 5.3249 | 59.0158 | L21591.5 | 1 | 243.8715 | 1.4468 | 2.0837 | 1000000 | 207682.031 | 16.9174 | 2.0358 | 1591.5
L3 | 1591.5 | 5Z | 304242 | 301271 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.30431 | 2.4389e-05 | 2.0418 | 301.289 | 1.7879 | 2.0835 | 4.4521e-06 | 0.033194 | 5.4781 | 53.8621 | L31591.5 | 1 | 301.289 | 1.7879 | 2.0835 | 1000000 | 304305.051 | 24.389 | 2.0418 | 1591.5
L4 | 1591.5 | 5Z | 398848 | 350041 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.39893 | 3.178e-05 | 2.0443 | 350.0619 | 2.0779 | 2.0833 | 5.7354e-06 | 0.042113 | 5.541 | 49.3423 | L41591.5 | 1 | 350.0619 | 2.0779 | 2.0833 | 1000000 | 398930.6571 | 31.7803 | 2.0443 | 1591.5
L5 | 1591.5 | 5Z | 494609 | 395561 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.49471 | 3.9293e-05 | 2.0455 | 395.5846 | 2.3489 | 2.0831 | 7.0502e-06 | 0.051343 | 5.5734 | 45.7493 | L51591.5 | 1 | 395.5846 | 2.3489 | 2.0831 | 1000000 | 494711.5026 | 39.2934 | 2.0455 | 1591.5
L6 | 1591.5 | 5Z | 592259 | 438801 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.59238 | 4.6971e-05 | 2.0462 | 438.8271 | 2.6065 | 2.0829 | 8.3995e-06 | 0.060868 | 5.5921 | 42.8219 | L61591.5 | 1 | 438.8271 | 2.6065 | 2.0829 | 1000000 | 592381.7396 | 46.9715 | 2.0462 | 1591.5
L7 | 1591.5 | 5Z | 689707 | 480181 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.68985 | 5.4643e-05 | 2.0467 | 480.2096 | 2.8531 | 2.0827 | 9.7512e-06 | 0.070443 | 5.6038 | 40.5031 | L71591.5 | 1 | 480.2096 | 2.8531 | 2.0827 | 1000000 | 689849.9346 | 54.6435 | 2.0467 | 1591.5
L8 | 1591.5 | 5Z | 782991 | 518311 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.78315 | 6.1993e-05 | 2.047 | 518.3419 | 3.0806 | 2.0825 | 1.1048e-05 | 0.079649 | 5.6112 | 38.6766 | L81591.5 | 1 | 518.3419 | 3.0806 | 2.0825 | 1000000 | 783153.2668 | 61.9935 | 2.047 | 1591.5
L9 | 1591.5 | 5Z | 878170 | 556590 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.87835 | 6.9497e-05 | 2.0472 | 556.6232 | 3.3089 | 2.0823 | 1.2374e-05 | 0.089072 | 5.6166 | 37.1491 | L91591.5 | 1 | 556.6232 | 3.3089 | 2.0823 | 1000000 | 878351.9916 | 69.4967 | 2.0472 | 1591.5
L10 | 1591.5 | 5Z | 984940 | 597390 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.98514 | 7.7917e-05 | 2.0474 | 597.4256 | 3.5525 | 2.0821 | 1.3862e-05 | 0.099665 | 5.6209 | 35.6448 | L101591.5 | 1 | 597.4256 | 3.5525 | 2.0821 | 1000000 | 985144.1186 | 77.917 | 2.0474 | 1591.5

## pyReport
| 2025 Coax | [6 empty] | 2020 S22075 Coax!
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | delta L /U | delta R /U
L10 | 10 | 101.0794 | 0.0079947 | 2.0473 | 69.8872 | 0.41526 | 2.0827 | 10 | 101.1231 | 0.0079574 | 2.0495 | 69.7374 | 0.41445 | 2.0827 | -5.4667 | 0.36062
L9 | 9 | 88.9702 | 0.0070399 | 2.0472 | 64.832 | 0.38512 | 2.0829 | 9 | 89.0051 | 0.0070067 | 2.0494 | 64.6747 | 0.38427 | 2.0829 | -4.9597 | 0.40822
L8 | 8 | 78.8676 | 0.0062436 | 2.0469 | 60.2997 | 0.35812 | 2.083 | 8 | 78.8959 | 0.006214 | 2.0491 | 60.1387 | 0.35724 | 2.083 | -4.5205 | 0.44948
L7 | 7 | 68.933 | 0.005461 | 2.0466 | 55.7274 | 0.33089 | 2.0832 | 7 | 68.9557 | 0.0054349 | 2.0488 | 55.5799 | 0.3301 | 2.0832 | -4.1629 | 0.44584
L6 | 6 | 59.0898 | 0.0046862 | 2.0462 | 51.0192 | 0.30287 | 2.0833 | 6 | 59.107 | 0.0046637 | 2.0483 | 50.8781 | 0.30211 | 2.0833 | -3.6581 | 0.46615
L5 | 5 | 49.0473 | 0.0038968 | 2.0454 | 45.9093 | 0.27247 | 2.0835 | 5 | 49.0604 | 0.003878 | 2.0475 | 45.7702 | 0.27172 | 2.0835 | -3.3752 | 0.51068
L4 | 4 | 39.3458 | 0.0031359 | 2.0441 | 40.5635 | 0.24069 | 2.0837 | 4 | 39.355 | 0.0031207 | 2.0461 | 40.4068 | 0.23984 | 2.0837 | -2.9534 | 0.65109
L3 | 3 | 29.2647 | 0.0023489 | 2.0412 | 34.4689 | 0.20448 | 2.0838 | 3 | 29.2704 | 0.0023376 | 2.0432 | 34.3478 | 0.20384 | 2.0838 | -2.4569 | 0.59215
L2 | 2 | 19.9887 | 0.0016328 | 2.0348 | 28.0007 | 0.16609 | 2.084 | 2 | 19.9916 | 0.001625 | 2.0365 | 27.9029 | 0.16558 | 2.084 | -1.7638 | 0.58862
L1 | 1 | 10.1495 | 0.00090477 | 2.0133 | 19.4852 | 0.11559 | 2.084 | 1 | 10.1508 | 0.00090104 | 2.0143 | 19.4083 | 0.11521 | 2.084 | -1.4256 | 0.66501
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 | 10 | 985.1441 | 0.077917 | 2.0474 | 597.4256 | 3.5525 | 2.0821 | 10 | 988.3118 | 0.077768 | 2.0496 | 599.6681 | 3.5659 | 2.0821 | -40.654 | -0.63124
L9 | 9 | 878.352 | 0.069497 | 2.0472 | 556.6232 | 3.3089 | 2.0823 | 9 | 880.8806 | 0.069341 | 2.0494 | 558.6212 | 3.3208 | 2.0823 | -36.3847 | -0.60382
L8 | 8 | 783.1533 | 0.061993 | 2.047 | 518.3419 | 3.0806 | 2.0825 | 8 | 785.1729 | 0.061836 | 2.0492 | 519.6584 | 3.0884 | 2.0825 | -32.5778 | -0.42735
L7 | 7 | 689.8499 | 0.054643 | 2.0467 | 480.2096 | 2.8531 | 2.0827 | 7 | 691.4255 | 0.054489 | 2.0489 | 480.7646 | 2.8565 | 2.0827 | -28.8344 | -0.1945
L6 | 6 | 592.3817 | 0.046971 | 2.0462 | 438.8271 | 2.6065 | 2.0829 | 6 | 593.5664 | 0.046825 | 2.0484 | 439.0086 | 2.6076 | 2.0829 | -25.2203 | -0.069601
L5 | 5 | 494.7115 | 0.039293 | 2.0455 | 395.5846 | 2.3489 | 2.0831 | 5 | 495.5522 | 0.03916 | 2.0477 | 395.4624 | 2.3482 | 2.0831 | -21.3945 | 0.051999
L4 | 4 | 398.9307 | 0.03178 | 2.0443 | 350.0619 | 2.0779 | 2.0833 | 4 | 399.4904 | 0.031663 | 2.0463 | 349.5711 | 2.0751 | 2.0833 | -17.6117 | 0.23616
L3 | 3 | 304.3051 | 0.024389 | 2.0418 | 301.289 | 1.7879 | 2.0835 | 3 | 304.6368 | 0.024292 | 2.0438 | 300.8116 | 1.7851 | 2.0835 | -13.6027 | 0.26698
L2 | 2 | 207.682 | 0.016917 | 2.0358 | 243.8715 | 1.4468 | 2.0837 | 2 | 207.8499 | 0.016847 | 2.0376 | 243.2545 | 1.4432 | 2.0837 | -9.92 | 0.42649
L1 | 1 | 103.8955 | 0.009209 | 2.0143 | 168.189 | 0.99775 | 2.0837 | 1 | 103.9444 | 0.0091732 | 2.0154 | 167.6371 | 0.9945 | 2.0837 | -5.3104 | 0.55321

## re2019
Prepare for full recalculation of the 2019 measurement
On 4 December 2019 Minyu did comparison of twisted vs coax using the UB
|  | Coax | [7 empty] | For python processing
|  |  | L | R | [6 empty] | xdial | rdial
5Z |  | 10 | 988108 | 599631 | 599625 | 2660 | 3225 | 3198 | 0.3236% | 0 | 0 | 6
|  | 9 | 880699 | 558587 | 558581 | 2112 | 3018 | 2991 | 0.3396% | 1 | 103923 | 167631
|  | 8 | 785011 | 519627 | 519621 | 1678 | 2304 | 2277 | 0.2901% | 2 | 207807 | 243243
|  | 7 | 691283 | 480736 | 480730 | 1296 | 1803 | 1776 | 0.2569% | 3 | 304574 | 300796
|  | 6 | 593444 | 438983 | 438977 | 969 | 1102 | 1075 | 0.1811% | 4 | 399408 | 349552
|  | 5 | 495450 | 395440 | 395434 | 672 | 934 | 907 | 0.1831% | 5 | 495450 | 395440
|  | 4 | 399408 | 349552 | 349546 | 439 | 575 | 548 | 0.1372% | 6 | 593444 | 438983
|  | 3 | 304574 | 300796 | 300790 | 251 | 452 | 425 | 0.1395% | 7 | 691283 | 480736
|  | 2 | 207807 | 243243 | 243237 | 116 | 260 | 233 | 0.1121% | 8 | 785011 | 519627
|  | 1 | 103923 | 167631 | 167625 | 28 | 95 | 68 | 0.06543% | 9 | 880699 | 558587
|  | 0 | 0 | 6 | [5 empty] | 10 | 988108 | 599631
|  | Twisted | [8 empty] | xdial | rdial
|  | 10 | 985448 | 596406 | 596427 |  |  |  |  | 0 | 0 | -21
|  | 9 | 878587 | 555569 | 555590 |  |  |  |  | 1 | 103895 | 167536
|  | 8 | 783333 | 517323 | 517344 |  |  |  |  | 2 | 207691 | 242983
|  | 7 | 689987 | 478933 | 478954 |  |  |  |  | 3 | 304323 | 300344
|  | 6 | 592475 | 437881 | 437902 |  |  |  |  | 4 | 398969 | 348977
|  | 5 | 494778 | 394506 | 394527 |  |  |  |  | 5 | 494778 | 394506
|  | 4 | 398969 | 348977 | 348998 |  |  |  |  | 6 | 592475 | 437881
|  | 3 | 304323 | 300344 | 300365 |  |  |  |  | 7 | 689987 | 478933
|  | 2 | 207691 | 242983 | 243004 |  |  |  |  | 8 | 783333 | 517323
|  | 1 | 103895 | 167536 | 167557 |  |  |  |  | 9 | 878587 | 555569
|  | 0 | 0 | -21 | [5 empty] | 10 | 985448 | 596406
4Z |  | Coax
|  |  | L | R | [6 empty] | xdial | rdial
|  | 10 | 1011026 | 697686 | 697324 | 277 | 1131 | 831 | 0.08219% | 0 | 3 | 362
|  | 9 | 889871 | 647063 | 646701 | 212 | 960 | 660 | 0.07417% | 1 | 101490 | 194431
|  | 8 | 788799 | 601706 | 601344 | 166 | 816 | 516 | 0.06542% | 2 | 199878 | 279371
|  | 7 | 689418 | 556121 | 555759 | 126 | 796 | 496 | 0.07194% | 3 | 292647 | 343815
| dislexic swap? | 6 | 590951 | 509106 | 508744 | 89 | 641 | 341 | 0.0577% | 4 | 393472 | 404401
|  | 5 | 490506 | 458031 | 457669 | 58 | 556 | 256 | 0.05219% | 5 | 490506 | 458031
|  | 4 | 393472 | 404401 | 404039 | 38 | 368 | 68 | 0.01728% | 6 | 590951 | 509106
|  | 3 | 292647 | 343815 | 343453 | 17 | 322 | 22 | 0.007518% | 7 | 689418 | 556121
|  | 2 | 199878 | 279371 | 279009 | 6 | 296 | -4 | -0.002001% | 8 | 788799 | 601706
|  | 1 | 101490 | 194431 | 194069 | 1 | 296 | -4 | -0.003941% | 9 | 889871 | 647063
|  | 0 | 3 | 362 | [5 empty] | 10 | 1011026 | 697686
|  | Twisted | [8 empty] | xdial | rdial
|  | 10 | 1010749 | 696555 | 696493 |  |  |  |  | 0 | 0 | 62
|  | 9 | 889659 | 646103 | 646041 |  |  |  |  | 1 | 101490 | 194135
|  | 8 | 788633 | 600890 | 600828 |  |  |  |  | 2 | 199872 | 279075
|  | 7 | 689292 | 555325 | 555263 |  |  |  |  | 3 | 292630 | 343493
|  | 6 | 590862 | 508465 | 508403 |  |  |  |  | 4 | 393434 | 404033
|  | 5 | 490448 | 457475 | 457413 |  |  |  |  | 5 | 490448 | 457475
|  | 4 | 393434 | 404033 | 403971 |  |  |  |  | 6 | 590862 | 508465
|  | 3 | 292630 | 343493 | 343431 |  |  |  |  | 7 | 689292 | 555325
|  | 2 | 199872 | 279075 | 279013 |  |  |  |  | 8 | 788633 | 600890
|  | 1 | 101490 | 194135 | 194073 |  |  |  |  | 9 | 889659 | 646103
|  | 0 | 0 | 62 | 0 |  |  |  |  | 10 | 1010749 | 696555

## coax_v_twist
On 3 December 2019 the effect of twisted lead vs coaxial connection were investigated.
Measurements in AI VIII
LCR
Twisted lead variability
|  |  |  | normal/position ratio |  | ppm
| normal | 984.9926 |  | 1 |  | 0
| below | 984.9485 |  | 1 |  | 44.7739
| horizontal | 984.9939 |  | 1 |  | -1.3198
| above | 984.9565 |  | 1 |  | 36.6514
| change twist |  |  | low/high twist ratio
| below | 985.4645 |  | 0.99948 |  | -523.6109
| horizontal | 985.4827 |  | 0.9995 |  | -496.0006
| above | 985.4838 |  | 0.99946 |  | -535.0671
Above data suggests that dependence of angle of lead entry is 40 ppm
Change in twist has a much larger 500 ppm shift!
| Coax connection |  |  | horizontal/postion |  |  |  | twisted horizontal/coax |  | ppm
| below | 988.2161 |  | 1 |  | 4.4525 |  | 0.99723 |  | -2765.9942
| horizontal | 988.2205 |  | 1 |  | 0 |  | 0.99723 |  | -2770.4343
| above | 988.2231 |  | 1 |  | -2.631 |  | 0.99723 |  | -2773.058
The coax connection is not sensitive to lead position.
The coax connection gives a higher inductance!
The very large change seen here might have been caused by some other change?
|  |  |  | twisted/coax |  | ppm
UB | twisted | 590859 |  | 0.99985 |  | -152.2974
|  | 1010745 |  | 0.99972 |  | -275.9578
| coax | 590949 |  |  |  |  | coax reading is higher
1011024
Note that the zero dial setting was not included in the checks.
28.04
On 4 December 2019 Minyu did comparison of twisted vs coax using the UB | [9 empty] | 1.9604
Coax
| Raw dial readings |  |  | coax/twisted |  | ppm |  | coax-twisted |  |  | calc-meas
10 | 988108 |  |  | 1.0027 |  | 2699.2799 |  | 2660 |  | 2559.6348 | -100.3652
9 | 880699 |  |  | 1.0024 |  | 2403.8598 |  | 2112 |  | 2081.9727 | -30.0273
8 | 785011 |  |  | 1.0021 |  | 2142.1286 |  | 1678 |  | 1652.7061 | -25.2939
7 | 691283 |  |  | 1.0019 |  | 1878.2963 |  | 1296 |  | 1272.0618 | -23.9382
6 | 593444 |  |  | 1.0016 |  | 1635.512 |  | 969 |  | 940.2985 | -28.7015
5 | 495450 |  |  | 1.0014 |  | 1358.1849 |  | 672 |  | 657.7166 | -14.2834
4 | 399408 |  |  | 1.0011 |  | 1100.3361 |  | 439 |  | 424.6747 | -14.3253
3 | 304574 |  |  | 1.0008 |  | 824.7816 |  | 251 |  | 241.6165 | -9.3835
2 | 207807 |  |  | 1.0006 |  | 558.522 |  | 116 |  | 109.1232 | -6.8768
1 | 103923 |  |  | 1.0003 |  | 269.5029 |  | 28 |  | 28.04 | 0.04
0 | 0
[6 empty] | 10.0158 |  | 95
Twisted
10 | 985448
9 | 878587
8 | 783333
7 | 689987
6 | 592475
5 | 494778
4 | 398969
3 | 304323
2 | 207691
1 | 103895
0 | 0
This pattern of difference might be modelled as a mutual inductance per section.
Why does the ppm difference scale with dial setting rather than being constant?
Note that coil inductance is proportional to the square of the number of turns
| 100 | turns max | 2700 | delta at max turns
| 0.133 | effective fractional turn of lead |  |  |  | 0.131
| 1.5 | power |  |  |  | 1.5
Inductance | N |  |  | N+ | total L | calculated | measured | calc-meas | N^2 | (N+)^2 | ratio
100 | 100 |  | 1 | 100.133 | 100.2662 | 0.26618 | 0.266 | 0.00017689 | 10000 | 10026.6177 | 1.0027
90 | 94.8683 |  | 0.94868 | 94.9819 | 90.2156 | 0.21559 | 0.2112 | 0.004389 | 9000 | 9021.5589 | 1.0024
80 | 89.4427 |  | 0.89443 | 89.5379 | 80.1703 | 0.17033 | 0.1678 | 0.0025306 | 8000 | 8017.0331 | 1.0021
70 | 83.666 |  | 0.83666 | 83.7439 | 70.1304 | 0.1304 | 0.1296 | 0.00080067 | 7000 | 7013.0401 | 1.0019
60 | 77.4597 |  | 0.7746 | 77.5215 | 60.0958 | 0.095798 | 0.0969 | -0.0011018 | 6000 | 6009.5798 | 1.0016
50 | 70.7107 |  | 0.70711 | 70.7577 | 50.0665 | 0.066522 | 0.0672 | -0.00067789 | 5000 | 5006.6522 | 1.0013
40 | 63.2456 |  | 0.63246 | 63.2792 | 40.0426 | 0.042571 | 0.0439 | -0.0013287 | 4000 | 4004.2571 | 1.0011
30 | 54.7723 |  | 0.54772 | 54.7941 | 30.0239 | 0.023945 | 0.0251 | -0.0011552 | 3000 | 3002.3945 | 1.0008
20 | 44.7214 |  | 0.44721 | 44.7333 | 20.0106 | 0.010641 | 0.0116 | -0.00095858 | 2000 | 2001.0641 | 1.0005
10 | 31.6228 |  | 0.31623 | 31.627 | 10.0027 | 0.0026602 | 0.0028 | -0.00013982 | 1000 | 1000.266 | 1.0003
In essence fewer turns are needed for the final step than for the first step.
This would suggest that M had proportionally more impact on the lower values.
0.01
|  |  | 100 | 100.01 | 1.0002 | [7 empty] | plot against turns rather than inductance
|  |  | 50 | 50.01 | 1.0004
|  |  | 10 | 10.01 | 1.002
Inductance | N |  |  | N+ | L
100 | 100 |  | 1 | 100.133 | 100.2662
90 | 94.8683 |  | 0.94868 | 95.0013 | 90.2525
80 | 89.4427 |  | 0.89443 | 89.5757 | 80.2381
70 | 83.666 |  | 0.83666 | 83.799 | 70.2227
60 | 77.4597 |  | 0.7746 | 77.5927 | 60.2062
50 | 70.7107 |  | 0.70711 | 70.8437 | 50.1883
40 | 63.2456 |  | 0.63246 | 63.3786 | 40.1684
30 | 54.7723 |  | 0.54772 | 54.9053 | 30.1459
20 | 44.7214 |  | 0.44721 | 44.8544 | 20.1191
10 | 31.6228 |  | 0.31623 | 31.7558 | 10.0843
|  |  | 101 | 1001
|  |  | 100 | 1000
|  |  | 1.0201 | 1.002
Now understand capacitance to be the main cause of change when using the coax cable.
| Readings |  |  | Zero correcte
| C | G |  | C | G
Short coax | 3.442e-11 | 3.91e-10 |  | 3.4418e-11 | 3.9475e-10 |  | For 1 H unit assuming nominal L and 600 ohm
twisted leads | 9.6947e-13 | 6.417e-11 |  | 9.6746e-13 | 6.792e-11
LCR zero | 2.01e-15 | -3.75e-12 |  |  |  |  | cap induced L change |  |  | coax-twisted
[7 empty] | short coax | twisted | [5 empty] | error / %
|  |  | R | [6 empty] | calc | measured
|  |  | 168.4377 | 0.28003 | 168.0187 | 1 | 3.3459e-05 | 9.402e-07 |  | 3.2519e-05 | 3.1285e-05 | 0.96207 | 31.2855 | 0.003346%
|  |  | 244.4601 | 0.40642 | 243.8521 | 2 | 0.00013572 | 3.8125e-06 |  | 0.00013191 | 0.0001305 | 0.98931 | 130.5014 | 0.006786%
|  |  | 302.212 | 0.50243 | 301.4603 | 3 | 0.00030696 | 8.6198e-06 |  | 0.00029834 | 0.00028345 | 0.95009 | 283.4527 | 0.01023%
|  |  | 351.3596 | 0.58414 | 350.4856 | 4 | 0.00054723 | 1.5362e-05 |  | 0.00053187 | 0.00049125 | 0.92362 | 491.2475 | 0.01368%
|  |  | 397.4338 | 0.66074 | 396.4452 | 5 | 0.00085655 | 2.4037e-05 |  | 0.00083251 | 0.00075915 | 0.91188 | 759.1475 | 0.01713%
|  |  | 441.3742 | 0.73379 | 440.2763 | 6 | 0.001235 | 3.4644e-05 |  | 0.0012003 | 0.0010919 | 0.90964 | 1091.8631 | 0.02058%
|  |  | 483.0297 | 0.80305 | 481.8282 | 7 | 0.0016826 | 4.7186e-05 |  | 0.0016354 | 0.001485 | 0.908 | 1484.962 | 0.02404%
|  |  | 521.6111 | 0.86719 | 520.3137 | 8 | 0.0021996 | 6.1663e-05 |  | 0.0021379 | 0.0019185 | 0.89739 | 1918.53 | 0.02749%
|  |  | 560.1833 | 0.93132 | 558.7899 | 9 | 0.0027858 | 7.8072e-05 |  | 0.0027078 | 0.0024168 | 0.89256 | 2416.8473 | 0.03095%
|  |  | 601.4962 | 1 | 600 | 10 | 0.0034414 | 9.6411e-05 |  | 0.003345 | 0.0030448 | 0.91028 | 3044.8386 | 0.03441%
[9 empty] | 10 | 0.003345
[9 empty] | 9 | 0.0027078
[9 empty] | 8 | 0.0021379
[9 empty] | 7 | 0.0016354
[9 empty] | 6 | 0.0012003
[9 empty] | 5 | 0.00083251
[9 empty] | 4 | 0.00053187
[9 empty] | 3 | 0.00029834
[9 empty] | 2 | 0.00013191
[9 empty] | 1 | 3.2519e-05

## LHistory
History
40585
[19 empty] | as ratio | Typical U
|  | S20803 | S21193 | S21431 | S21637 | S21784 | S21960 | S22075 | [9 empty] | twist
| Nominal | 1992-09-03 | 1997-04-03 | 2001-09-14 | 2006-06-15 | 2011-02-11 | 2015-08-20 | 2019-12-04 | 2024-12-20 | 1992-09-03 | 1997-04-03 | 2001-09-14 | 2006-06-15 | 2011-02-11 | 2015-08-20 | 2019-12-04 | 2024-12-20 | 2025-11-30
| 100 | 101.1 | 101.0931 | 101.0775 | 101.0768 | 101.0934 | 101.1144 | 101.1231 | 101.0794 | 1 | 0.99993 | 0.99978 | 0.99977 | 0.99993 | 1.0001 | 1.0002 | 0.9998 |  | 6.9238e-05 | 0.0070
| 90 | 89 | 88.9737 | 88.9669 | 88.9616 | 88.9860 | 89.0063 | 89.0051 | 88.9702 | 1 | 0.9997 | 0.99963 | 0.99957 | 0.99984 | 1.0001 | 1.0001 | 0.99967 |  | 7.3034e-05 | 0.0065
| 80 | 78.9 | 78.8686 | 78.8643 | 78.8583 | 78.8826 | 78.9009 | 78.8959 | 78.8676 | 1 | 0.9996 | 0.99955 | 0.99947 | 0.99978 | 1 | 0.99995 | 0.99959 |  | 7.4778e-05 | 0.0059
| 70 | 68.9 | 68.933 | 68.9293 | 68.9246 | 68.9465 | 68.9631 | 68.9557 | 68.933 | 1 | 1.0005 | 1.0004 | 1.0004 | 1.0007 | 1.0009 | 1.0008 | 1.0005 |  | 7.5472e-05 | 0.0052
| 60 | 59.1 | 59.0888 | 59.0861 | 59.082 | 59.1017 | 59.1163 | 59.107 | 59.0898 | 1 | 0.99981 | 0.99976 | 0.9997 | 1 | 1.0003 | 1.0001 | 0.99983 |  | 7.6142e-05 | 0.0045
| 50 | 49.05 | 49.0463 | 49.0445 | 49.0405 | 49.0579 | 49.0702 | 49.0604 | 49.0473 | 1 | 0.99993 | 0.99989 | 0.99981 | 1.0002 | 1.0004 | 1.0002 | 0.99994 |  | 7.7472e-05 | 0.0038
| 40 | 39.35 | 39.3448 | 39.3434 | 39.3402 | 39.3546 | 39.3650 | 39.355 | 39.3458 | 1 | 0.99987 | 0.99983 | 0.99975 | 1.0001 | 1.0004 | 1.0001 | 0.99989 |  | 7.878e-05 | 0.0031
| 30 | 29.26 | 29.2635 | 29.2625 | 29.2602 | 29.2715 | 29.2794 | 29.2704 | 29.2647 | 1 | 1.0001 | 1.0001 | 1 | 1.0004 | 1.0007 | 1.0004 | 1.0002 |  | 7.8606e-05 | 0.0023
| 20 | 19.99 | 19.9873 | 19.9867 | 19.985 | 19.9931 | 19.9988 | 19.9916 | 19.9887 | 1 | 0.99987 | 0.99983 | 0.99975 | 1.0002 | 1.0004 | 1.0001 | 0.99994 |  | 8.004e-05 | 0.0016
| 10 | 10.15 | 10.1489 | 10.1486 | 10.1477 | 10.1520 | 10.1550 | 10.1508 | 10.1495 | 1 | 0.9999 | 0.99986 | 0.99977 | 1.0002 | 1.0005 | 1.0001 | 0.99995 |  | 8.867e-05 | 0.0009
| 1000 | 985 | 986.0623 | 985.2113 | 985.1001 | 984.939 | 985.244 | 988.3118 | 985.1441 | 1 | 1.0011 | 1.0002 | 1.0001 | 0.99994 | 1.0002 | 1.0034 | 1.0001 | 0 | 7.0051e-05 | 0.069
| 900 | 879 | 879.06 | 878.3959 | 878.2876 | 878.232 | 878.380 | 880.8806 | 878.352 | 1 | 1.0001 | 0.99931 | 0.99919 | 0.99913 | 0.99929 | 1.0021 | 0.99926 | 0 | 6.8259e-05 | 0.060
| 800 | 783 | 783.7116 | 783.1934 | 783.0962 | 783.086 | 783.228 | 785.1729 | 783.1533 | 1 | 1.0009 | 1.0002 | 1.0001 | 1.0001 | 1.0003 | 1.0028 | 1.0002 | 0 | 6.7688e-05 | 0.053
| 700 | 690 | 690.2875 | 689.8752 | 689.8051 | 689.813 | 689.929 | 691.4255 | 689.8499 | 1 | 1.0004 | 0.99982 | 0.99972 | 0.99973 | 0.9999 | 1.0021 | 0.99978 | 0 | 6.8116e-05 | 0.047
| 600 | 592 | 592.7167 | 592.4074 | 592.3583 | 592.360 | 592.471 | 593.5664 | 592.3817 | 1 | 1.0012 | 1.0007 | 1.0006 | 1.0006 | 1.0008 | 1.0026 | 1.0006 | 0 | 6.7568e-05 | 0.040
| 500 | 494.8 | 494.9729 | 494.7435 | 494.7064 | 494.710 | 494.803 | 495.5522 | 494.7115 | 1 | 1.0003 | 0.99989 | 0.99981 | 0.99982 | 1 | 1.0015 | 0.99982 | 0 | 6.8715e-05 | 0.034
| 400 | 399 | 399.1135 | 398.9589 | 398.9299 | 398.931 | 399.008 | 399.4904 | 398.9307 | 1 | 1.0003 | 0.9999 | 0.99982 | 0.99983 | 1 | 1.0012 | 0.99983 | 0 | 6.7669e-05 | 0.027
| 300 | 304.4 | 304.4292 | 304.3255 | 304.3075 | 304.305 | 304.371 | 304.6368 | 304.3051 | 1 | 1.0001 | 0.99976 | 0.9997 | 0.99969 | 0.9999 | 1.0008 | 0.99969 | 0 | 6.8988e-05 | 0.021
| 200 | 207.7 | 207.7596 | 207.6978 | 207.6929 | 207.691 | 207.733 | 207.8499 | 207.682 | 1 | 1.0003 | 0.99999 | 0.99997 | 0.99996 | 1.0002 | 1.0007 | 0.99991 | 0 | 6.7405e-05 | 0.014
| 100 | 103.9 | 103.9278 | 103.9029 | 103.9039 | 103.900 | 103.923 | 103.9444 | 103.8955 | 1 | 1.0003 | 1 | 1 | 1 | 1.0002 | 1.0004 | 0.99996 | 0 | 7.6997e-05 | 0.008
Note on D16: Christian: Actually UB read 985.254

## exampleU
Budget for L8 1591.5
ureal(60.08704790830853,0.1713140491986457,20.451673221652594)
uuttempL8 | 0.12049
temperature | 0.12043
Amp_phase_err | 0.013329
rtempcoL8 | 0.012019
Adialresureal(62.0,10.0,10.0) | 0.0011776
Adialresureal(600890.0,10.0,10.0) | 0.0011776
Adial_lin62 | 0.0010001
Adial_lin600890 | 0.0010001
Amperr | 0.00084121
R4Aacdc | 0.00015022
G1acdc | 0.00015022
G2acdc | 0.00015022
G1alpha | 6.0087e-05
G1stab | 6.0087e-05
R4Aalpha | 6.0087e-05
G2alpha | 6.0087e-05
R4Astab | 6.0087e-05
G2stab | 6.0087e-05
cal_temp | 4.2488e-05
G2err | 3.0043e-05
G1err | 3.0043e-05
R4Aerr | 3.0042e-05
uuttempL0 | 1.2432e-05
rtempcoL0 | 1.2401e-06
C1err | 0
C1alpha | 0
C1stab | 0
Bdial_lin0 | 0
Bdialresureal(0.0,1.0,10.0, label='Bdial_lin0') | 0
Bdial_lin788633 | 0
Bdialresureal(788633.0,1.0,10.0) | 0
ureal(0.0788795386314938,3.031875053076101e-06,27.783955120984952)
uuttempL8 | 1.9769e-06
temperature | 1.9211e-06
Amperr | 1.1043e-06
C1stab | 3.944e-07
R4Aacdc | 1.972e-07
G2acdc | 1.972e-07
xtempcoL8 | 1.972e-07
Bdialres0.0 | 1.1778e-07
Bdialres788633.0 | 1.1778e-07
Amp_phase_err | 1.0161e-07
Bdial_lin0 | 1.0002e-07
Bdial_lin788633 | 1.0002e-07
C1alpha | 7.888e-08
R4Aalpha | 7.888e-08
G2alpha | 7.888e-08
R4Astab | 7.888e-08
G2stab | 7.888e-08
C1err | 7.8868e-08
cal_temp | 5.5776e-08
G2err | 3.9439e-08
R4Aerr | 3.9438e-08
G1err | 0
G1acdc | 0
G1alpha | 0
G1stab | 0
Adial_lin62 | 0
Adialres62.0 | 0
uuttempL0 | 0
xtempcoL0 | 0
Adial_lin600890 | 0
Adialres600890.0 | 0

## Room Conditions
2017-12-04 16:30 | 20.9 | 53.6 |  | RH correction | -2.128 | from Y:\Quality\TECHPROC\E058-Lab Air Cond\loggercalibration_2017.xls
2017-12-04 16:35 | 20.9 | 53.8 | Temperature | RH
2017-12-04 16:40 | 20.9 | 54.8 | #REF! | 52.2805 | 50.1525
2017-12-04 16:45 | 20.9 | 54.4 | #REF! | 1.1136
2017-12-04 16:50 | 20.9 | 54.3
2017-12-04 16:55 | 20.8 | 54.2 | from GMH
2017-12-04 17:00 | 20.8 | 53.6
2017-12-04 17:05 | 20.9 | 53.4
2017-12-04 17:10 | 20.9 | 53.3
2017-12-04 17:15 | 20.8 | 53.7
2017-12-04 17:20 | 20.8 | 53
2017-12-04 17:25 | 20.8 | 52.6
2017-12-04 17:30 | 20.9 | 52.4
2017-12-04 17:35 | 20.9 | 52
2017-12-04 17:40 | 20.9 | 52
2017-12-04 17:45 | 20.9 | 52.1
2017-12-04 17:50 | 20.9 | 52
2017-12-04 17:55 | 20.9 | 51.8
2017-12-04 18:00 | 20.9 | 52.1
2017-12-04 18:05 | 20.9 | 52.1
2017-12-04 18:10 | 20.9 | 52
2017-12-04 18:15 | 20.9 | 52.3
2017-12-04 18:20 | 20.9 | 52.3
2017-12-04 18:25 | 20.9 | 52
2017-12-04 18:30 | 21 | 51.9
2017-12-04 18:35 | 21 | 52
2017-12-04 18:40 | 20.9 | 51.9
2017-12-04 18:45 | 20.9 | 51.9
2017-12-04 18:50 | 20.9 | 51.8
2017-12-04 18:55 | 20.9 | 51.8
2017-12-04 19:00 | 20.9 | 51.8
2017-12-04 19:05 | 20.9 | 51.8
2017-12-04 19:10 | 21 | 51.6
2017-12-04 19:15 | 21 | 51.7
2017-12-04 19:20 | 21 | 51.6
2017-12-04 19:25 | 21 | 51.7
2017-12-04 19:30 | 21 | 51.7
2017-12-04 19:35 | 21 | 51.7
2017-12-04 19:40 | 21 | 51.9
2017-12-04 19:45 | 21 | 51.9
2017-12-04 19:50 | 21 | 51.6
2017-12-04 19:55 | 21 | 51.7
2017-12-04 20:00 | 21 | 51.7
2017-12-04 20:05 | 21 | 51.7
2017-12-04 20:10 | 21 | 51.6
2017-12-04 20:15 | 21 | 51.9
2017-12-04 20:20 | 21 | 51.9
2017-12-04 20:25 | 21 | 51.9
2017-12-04 20:30 | 21 | 51.8
2017-12-04 20:35 | 21 | 51.7
2017-12-04 20:40 | 21.1 | 51.6
2017-12-04 20:45 | 21.1 | 51.8
2017-12-04 20:50 | 21 | 52
2017-12-04 20:55 | 21 | 52.2
2017-12-04 21:00 | 21 | 52.2
2017-12-04 21:05 | 21 | 52
2017-12-04 21:10 | 21 | 51.9
2017-12-04 21:15 | 21 | 52.1
2017-12-04 21:20 | 21.1 | 52
2017-12-04 21:25 | 21 | 52.2
2017-12-04 21:30 | 21 | 51.9
2017-12-04 21:35 | 21 | 52
2017-12-04 21:40 | 21 | 51.8
2017-12-04 21:45 | 21.1 | 51.5
2017-12-04 21:50 | 21 | 51.7
2017-12-04 21:55 | 21.1 | 51.6
2017-12-04 22:00 | 21.1 | 51.5
2017-12-04 22:05 | 21.1 | 51.3
2017-12-04 22:10 | 21.1 | 51.5
2017-12-04 22:15 | 21 | 51.7
2017-12-04 22:20 | 21.1 | 51.6
2017-12-04 22:25 | 21.1 | 52
2017-12-04 22:30 | 21.1 | 51.5
2017-12-04 22:35 | 21.1 | 51.4
2017-12-04 22:40 | 21.1 | 51.2
2017-12-04 22:45 | 21.1 | 51.2
2017-12-04 22:50 | 21.1 | 51.3
2017-12-04 22:55 | 21.1 | 51.4
2017-12-04 23:00 | 21.1 | 51.3
2017-12-04 23:05 | 21.1 | 51.3
2017-12-04 23:10 | 21.2 | 51.4
2017-12-04 23:15 | 21.2 | 51.3
2017-12-04 23:20 | 21.2 | 51.5
2017-12-04 23:25 | 21.1 | 51.7
2017-12-04 23:30 | 21.1 | 51.7
2017-12-04 23:35 | 21.1 | 51.3
2017-12-04 23:40 | 21.1 | 51.5
2017-12-04 23:45 | 21.1 | 51.4
2017-12-04 23:50 | 21.1 | 51.4
2017-12-04 23:55 | 21.1 | 51.3
2017-12-05 | 21.1 | 51.4
2017-12-05 00:05 | 21.1 | 51.3
2017-12-05 00:10 | 21.1 | 51
2017-12-05 00:15 | 21.1 | 51.3
2017-12-05 00:20 | 21.1 | 51.1
2017-12-05 00:25 | 21.2 | 51.2
2017-12-05 00:30 | 21.1 | 51.4
2017-12-05 00:35 | 21.2 | 51.3
2017-12-05 00:40 | 21.1 | 51.5
2017-12-05 00:45 | 21.2 | 51.3
2017-12-05 00:50 | 21.2 | 51.3
2017-12-05 00:55 | 21.1 | 51.3
2017-12-05 01:00 | 21.2 | 51.3
2017-12-05 01:05 | 21.1 | 51.4
2017-12-05 01:10 | 21.2 | 51.2
2017-12-05 01:15 | 21.2 | 51.1
2017-12-05 01:20 | 21.1 | 51.3
2017-12-05 01:25 | 21.2 | 51.9
2017-12-05 01:30 | 21.2 | 51.9
2017-12-05 01:35 | 21.1 | 52
2017-12-05 01:40 | 21.1 | 51.8
2017-12-05 01:45 | 21.2 | 51.8
2017-12-05 01:50 | 21.1 | 51.8
2017-12-05 01:55 | 21.1 | 51.9
2017-12-05 02:00 | 21.1 | 52.1
2017-12-05 02:05 | 21.1 | 51.9
2017-12-05 02:10 | 21.2 | 51.8
2017-12-05 02:15 | 21.2 | 51.7
2017-12-05 02:20 | 21.2 | 51.9
2017-12-05 02:25 | 21.2 | 51.9
2017-12-05 02:30 | 21.2 | 51.8
2017-12-05 02:35 | 21.1 | 51.6
2017-12-05 02:40 | 21.2 | 51.7
2017-12-05 02:45 | 21.2 | 51.6
2017-12-05 02:50 | 21.1 | 51.7
2017-12-05 02:55 | 21.1 | 51.6
2017-12-05 03:00 | 21.1 | 51.6
2017-12-05 03:05 | 21.1 | 51.6
2017-12-05 03:10 | 21.1 | 51.3
2017-12-05 03:15 | 21.2 | 51.1
2017-12-05 03:20 | 21.1 | 51.3
2017-12-05 03:25 | 21.1 | 51.6
2017-12-05 03:30 | 21.1 | 51.7
2017-12-05 03:35 | 21.1 | 51.6
2017-12-05 03:40 | 21.2 | 51.5
2017-12-05 03:45 | 21.1 | 51.3
2017-12-05 03:50 | 21.1 | 51.2
2017-12-05 03:55 | 21.1 | 51.2
2017-12-05 04:00 | 21.2 | 51
2017-12-05 04:05 | 21.2 | 51.1
2017-12-05 04:10 | 21.2 | 51.4
2017-12-05 04:15 | 21.2 | 51.6
2017-12-05 04:20 | 21.1 | 51.5
2017-12-05 04:25 | 21.1 | 51.8
2017-12-05 04:30 | 21.2 | 51.7
2017-12-05 04:35 | 21.1 | 52
2017-12-05 04:40 | 21.1 | 51.9
2017-12-05 04:45 | 21.2 | 52
2017-12-05 04:50 | 21.1 | 52
2017-12-05 04:55 | 21.2 | 51.8
2017-12-05 05:00 | 21.2 | 51.8
2017-12-05 05:05 | 21.1 | 51.9
2017-12-05 05:10 | 21.1 | 51.8
2017-12-05 05:15 | 21.1 | 51.6
2017-12-05 05:20 | 21.2 | 51.7
2017-12-05 05:25 | 21.2 | 51.8
2017-12-05 05:30 | 21.2 | 51.7
2017-12-05 05:35 | 21.1 | 51.8
2017-12-05 05:40 | 21.1 | 51.5
2017-12-05 05:45 | 21.1 | 51.4
2017-12-05 05:50 | 21.1 | 51.5
2017-12-05 05:55 | 21.1 | 51.4
2017-12-05 06:00 | 21.1 | 51.5
2017-12-05 06:05 | 21.1 | 51.4
2017-12-05 06:10 | 21.2 | 51.3
2017-12-05 06:15 | 21.2 | 51.2
2017-12-05 06:20 | 21.2 | 51.2
2017-12-05 06:25 | 21.2 | 51.2
2017-12-05 06:30 | 21.2 | 51.3
2017-12-05 06:35 | 21.1 | 51.3
2017-12-05 06:40 | 21.2 | 51.1
2017-12-05 06:45 | 21.2 | 51.4
2017-12-05 06:50 | 21.2 | 51.6
2017-12-05 06:55 | 21.1 | 51.4
2017-12-05 07:00 | 21.2 | 51.4
2017-12-05 07:05 | 21.1 | 51.3
2017-12-05 07:10 | 21.2 | 51.3
2017-12-05 07:15 | 21.2 | 51.3
2017-12-05 07:20 | 21.1 | 51.4
2017-12-05 07:25 | 21.1 | 51.4
2017-12-05 07:30 | 21.1 | 51.3
2017-12-05 07:35 | 21.1 | 51.3
2017-12-05 07:40 | 21.1 | 51.3
2017-12-05 07:45 | 21.1 | 51.4
2017-12-05 07:50 | 21.2 | 51.3
2017-12-05 07:55 | 21.1 | 51.3
2017-12-05 08:00 | 21.1 | 51.4
2017-12-05 08:05 | 21.1 | 51.4
2017-12-05 08:10 | 21.1 | 51.4
2017-12-05 08:15 | 21.2 | 51.5
2017-12-05 08:20 | 21.2 | 51.6
2017-12-05 08:25 | 21 | 51.8
2017-12-05 08:30 | 21 | 51.8
2017-12-05 08:35 | 21 | 51.8
2017-12-05 08:40 | 21.1 | 51.6
2017-12-05 08:45 | 21 | 51.8
2017-12-05 08:50 | 21 | 51.7
2017-12-05 08:55 | 21 | 51.9
2017-12-05 09:00 | 21.1 | 52.1
2017-12-05 09:05 | 21 | 52.2
2017-12-05 09:10 | 21 | 52.1
2017-12-05 09:15 | 21 | 52
2017-12-05 09:20 | 21 | 52
2017-12-05 09:25 | 21 | 52.1
2017-12-05 09:30 | 21 | 52.1
2017-12-05 09:35 | 21 | 52.3
2017-12-05 09:40 | 20.9 | 52
2017-12-05 09:45 | 21 | 52
2017-12-05 09:50 | 20.9 | 51.8
2017-12-05 09:55 | 21 | 51.6
2017-12-05 10:00 | 20.9 | 51.7
2017-12-05 10:05 | 21 | 51.8
2017-12-05 10:10 | 21 | 51.7
2017-12-05 10:15 | 20.9 | 51.9
2017-12-05 10:20 | 20.9 | 51.9
2017-12-05 10:25 | 20.9 | 51.7
2017-12-05 10:30 | 20.8 | 51.9
2017-12-05 10:35 | 20.9 | 51.9
2017-12-05 10:40 | 20.9 | 52.3
2017-12-05 10:45 | 20.9 | 52.2
2017-12-05 10:50 | 20.8 | 52.4
2017-12-05 10:55 | 20.8 | 52.8
2017-12-05 11:00 | 20.9 | 52.4
2017-12-05 11:05 | 20.8 | 52.5
2017-12-05 11:10 | 20.9 | 52.8
2017-12-05 11:15 | 20.8 | 52.6
2017-12-05 11:20 | 20.8 | 52.7
2017-12-05 11:25 | 20.8 | 52.2
2017-12-05 11:30 | 20.8 | 52.5
2017-12-05 11:35 | 20.8 | 52.4
2017-12-05 11:40 | 20.8 | 52.4
2017-12-05 11:45 | 20.8 | 52.9
2017-12-05 11:50 | 20.8 | 52.7
2017-12-05 11:55 | 20.8 | 52.7
2017-12-05 12:00 | 20.8 | 52.9
2017-12-05 12:05 | 20.9 | 52.9
2017-12-05 12:10 | 20.9 | 53.1
2017-12-05 12:15 | 20.9 | 52.8
2017-12-05 12:20 | 20.9 | 52.9
2017-12-05 12:25 | 20.9 | 53.5
2017-12-05 12:30 | 20.9 | 53.2
2017-12-05 12:35 | 20.9 | 53.2
2017-12-05 12:40 | 20.8 | 53.4
2017-12-05 12:45 | 20.8 | 53.2
2017-12-05 12:50 | 20.9 | 53.7
2017-12-05 12:55 | 20.9 | 53.7
2017-12-05 13:00 | 20.9 | 53.2
2017-12-05 13:05 | 21 | 53.4
2017-12-05 13:10 | 20.9 | 53.8
2017-12-05 13:15 | 20.8 | 54.1
2017-12-05 13:20 | 20.8 | 54
2017-12-05 13:25 | 20.8 | 54
2017-12-05 13:30 | 20.8 | 54.2
2017-12-05 13:35 | 20.8 | 54.2
2017-12-05 13:40 | 20.8 | 54.5
2017-12-05 13:45 | 20.8 | 54.5
2017-12-05 13:50 | 20.8 | 54.4
2017-12-05 13:55 | 20.8 | 54.6
2017-12-05 14:00 | 20.7 | 54.7
2017-12-05 14:05 | 20.8 | 54.9
2017-12-05 14:10 | 20.8 | 54.9
2017-12-05 14:15 | 20.7 | 55.1
2017-12-05 14:20 | 20.7 | 55.6
2017-12-05 14:25 | 20.7 | 55.2
2017-12-05 14:30 | 20.8 | 55.3
2017-12-05 14:35 | 20.8 | 54.6
2017-12-05 14:40 | 20.7 | 54.8
2017-12-05 14:45 | 20.7 | 54.9
2017-12-05 14:50 | 20.7 | 54.7
2017-12-05 14:55 | 20.7 | 54.8
2017-12-05 15:00 | 20.8 | 54.7
2017-12-05 15:05 | 20.8 | 54.4
2017-12-05 15:10 | 20.7 | 54.5
2017-12-05 15:15 | 20.8 | 54.6
2017-12-05 15:20 | 20.8 | 54.6
2017-12-05 15:25 | 20.8 | 54.2
2017-12-05 15:30 | 20.8 | 54.8
2017-12-05 15:35 | 20.8 | 54.2
2017-12-05 15:40 | 20.7 | 54
2017-12-05 15:45 | 20.7 | 54.8
2017-12-05 15:50 | 20.7 | 54.3
2017-12-05 15:55 | 20.7 | 53.8
2017-12-05 16:00 | 20.7 | 53.8
2017-12-05 16:05 | 20.6 | 54.6
2017-12-05 16:10 | 20.6 | 54
2017-12-05 16:15 | 20.7 | 54.2
2017-12-05 16:20 | 20.7 | 54.4
2017-12-05 16:25 | 20.6 | 54.3
2017-12-05 16:30 | 20.7 | 54.2
2017-12-05 16:35 | 20.7 | 54.1
2017-12-05 16:40 | 20.7 | 53.7
2017-12-05 16:45 | 20.7 | 53.9

## Conditions
| Temp |  | RH
All GMH5I readings | #REF! | #REF!
Room | #REF! | #REF! | #REF! | #REF!