## Commentary
2020-07-29
Worksheet 'pyUBreadings' requires the user to enter all the dial readings and settings of the UB as well as estimates of the uncertainty in temperature and in the temperature coefficients for resistance and inductance. In general it is set up not to apply any temperature corrections, just to evaluate the uncertainty.
Use the UniversalBridge3 project to process the input data. This uses the latest ubdict.csv file for corrections to the Universal Bridge. Running 'calculate_imp.py' produces a new 'indResults.xlsx'. Paste these results into the 'pyUBresults' worksheet of this spreadsheet and, if correctly pasted, should immediately produce the report in the'pyReport' worksheet. Note that there are two named blocks in 'pyUBresults' and so the new data should just be pasted over the old data in order to preserve the inductor1 and inductor2 named blocks.
You should check that the input data listed in 'pyUBresults' matches the input data in 'pyUBreadings', just in case the wrong file names are enterd in 'calculate_imp.py'.

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | Muirhead A-517 |  | 4: Client; 10: Enter, say, 5 ppm as 5e-6
Date from | 2019-12-04 |  | 4: Job number; 5: 96330136; 10: 1 sigma uncertainties from specifications
to | 2019-12-05 |  | 4: S Number; 5: S22075
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
L0 | 1591.5 | 4Z | 3 | 419 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 4Z | 101493 | 195221 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 4Z | 199887 | 280480 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 4Z | 292655 | 345191 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 4Z | 393464 | 405970 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 4Z | 490515 | 459475 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 4Z | 590969 | 510603 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 4Z | 689431 | 557663 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 4Z | 788816 | 603296 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 4Z | 889889 | 648646 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 4Z | 1011051 | 699208 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L0 | 1591.5 | 5Z | 0 | 10 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 5Z | 103928 | 168318 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 5Z | 207818 | 244214 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 5Z | 304591 | 301974 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 5Z | 399423 | 351053 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 5Z | 495475 | 396953 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 5Z | 593476 | 440533 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 5Z | 691330 | 482103 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 5Z | 785060 | 521293 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 5Z | 880750 | 560533 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 5Z | 988170 | 601673 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
L0 | 1591.5 | 4Z | 3 | 419 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-07 | 0.0012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 4Z | 101493 | 195221 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.010151 | 9.0488e-07 | 2.0133 | 19.4814 | 0.11565 | 2.084 | 1.9074e-07 | 0.0016387 | 4.7441 | 70.5756 | L11591.5 | 1000 | 19481.3609 | 115.6546 | 2.084 | 1000000 | 10151.1033 | 0.90488 | 2.0133 | 1591.5
L2 | 1591.5 | 4Z | 199887 | 280480 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.019993 | 1.6331e-06 | 2.0348 | 28.0078 | 0.16622 | 2.084 | 3.0811e-07 | 0.0023941 | 5.3003 | 69.4302 | L21591.5 | 1000 | 28007.769 | 166.2208 | 2.084 | 1000000 | 19992.5424 | 1.6331 | 2.0348 | 1591.5
L3 | 1591.5 | 4Z | 292655 | 345191 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.029271 | 2.3494e-06 | 2.0412 | 34.4793 | 0.20464 | 2.0838 | 4.3004e-07 | 0.0032221 | 5.4632 | 63.5099 | L31591.5 | 1000 | 34479.2546 | 204.6365 | 2.0838 | 1000000 | 29271.2649 | 2.3494 | 2.0412 | 1591.5
L4 | 1591.5 | 4Z | 393464 | 405970 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.039354 | 3.1365e-06 | 2.0441 | 40.5575 | 0.24075 | 2.0837 | 5.6658e-07 | 0.0041703 | 5.5359 | 57.7295 | L41591.5 | 1000 | 40557.5168 | 240.7471 | 2.0837 | 1000000 | 39354.2541 | 3.1365 | 2.0441 | 1591.5
L5 | 1591.5 | 4Z | 490515 | 459475 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.049061 | 3.8979e-06 | 2.0454 | 45.9083 | 0.27256 | 2.0835 | 6.9978e-07 | 0.0051053 | 5.5701 | 53.3866 | L51591.5 | 1000 | 45908.3357 | 272.5569 | 2.0835 | 1000000 | 49061.3654 | 3.8979 | 2.0454 | 1591.5
L6 | 1591.5 | 4Z | 590969 | 510603 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.059109 | 4.6877e-06 | 2.0462 | 51.0214 | 0.30297 | 2.0833 | 8.3858e-07 | 0.0060852 | 5.5901 | 49.7882 | L61591.5 | 1000 | 51021.4404 | 302.9734 | 2.0833 | 1000000 | 59108.8472 | 4.6877 | 2.0462 | 1591.5
L7 | 1591.5 | 4Z | 689431 | 557663 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.068957 | 5.4629e-06 | 2.0466 | 55.7277 | 0.33099 | 2.0832 | 9.7515e-07 | 0.0070527 | 5.6021 | 46.9303 | L71591.5 | 1000 | 55727.7208 | 330.9867 | 2.0832 | 1000000 | 68957.0877 | 5.4629 | 2.0466 | 1591.5
L8 | 1591.5 | 4Z | 788816 | 603296 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.078898 | 6.246e-06 | 2.0469 | 60.2913 | 0.35816 | 2.083 | 1.1133e-06 | 0.0080338 | 5.6101 | 44.5818 | L81591.5 | 1000 | 60291.2927 | 358.1632 | 2.083 | 1000000 | 78897.6473 | 6.246 | 2.0469 | 1591.5
L9 | 1591.5 | 4Z | 889889 | 648646 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.089007 | 7.0428e-06 | 2.0472 | 64.8266 | 0.38518 | 2.0829 | 1.2541e-06 | 0.0090348 | 5.6158 | 42.6329 | L91591.5 | 1000 | 64826.563 | 385.1806 | 2.0829 | 1000000 | 89007.042 | 7.0428 | 2.0472 | 1591.5
L10 | 1591.5 | 4Z | 1011051 | 699208 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10113 | 7.9984e-06 | 2.0473 | 69.8831 | 0.41533 | 2.0827 | 1.4231e-06 | 0.010237 | 5.6206 | 40.5697 | L101591.5 | 1000 | 69883.0643 | 415.3268 | 2.0827 | 1000000 | 101125.7529 | 7.9984 | 2.0473 | 1591.5
L0 | 1591.5 | 5Z | 0 | 10 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-06 | 0.012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 5Z | 103928 | 168318 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10395 | 9.2128e-06 | 2.0144 | 168.318 | 0.99854 | 2.0837 | 1.9304e-06 | 0.016481 | 4.7725 | 60.5867 | L11591.5 | 1 | 168.318 | 0.99854 | 2.0837 | 1000000 | 103949.538 | 9.2128 | 2.0144 | 1591.5
L2 | 1591.5 | 5Z | 207818 | 244214 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.20786 | 1.6931e-05 | 2.0358 | 244.2186 | 1.4489 | 2.0837 | 3.1794e-06 | 0.024531 | 5.3253 | 59.0626 | L21591.5 | 1 | 244.2186 | 1.4489 | 2.0837 | 1000000 | 207861.0681 | 16.9311 | 2.0358 | 1591.5
L3 | 1591.5 | 5Z | 304591 | 301974 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.30465 | 2.4416e-05 | 2.0418 | 301.982 | 1.792 | 2.0835 | 4.4568e-06 | 0.033227 | 5.4784 | 53.9326 | L31591.5 | 1 | 301.982 | 1.792 | 2.0835 | 1000000 | 304654.1233 | 24.4162 | 2.0418 | 1591.5
L4 | 1591.5 | 5Z | 399423 | 351053 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.39951 | 3.1825e-05 | 2.0443 | 351.0639 | 2.0839 | 2.0833 | 5.7433e-06 | 0.042169 | 5.5413 | 49.4181 | L41591.5 | 1 | 351.0639 | 2.0839 | 2.0833 | 1000000 | 399505.7763 | 31.8254 | 2.0443 | 1591.5
L5 | 1591.5 | 5Z | 495475 | 396953 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.49558 | 3.9362e-05 | 2.0455 | 396.9667 | 2.3571 | 2.0831 | 7.0622e-06 | 0.051428 | 5.5735 | 45.8329 | L51591.5 | 1 | 396.9667 | 2.3571 | 2.0831 | 1000000 | 495577.6821 | 39.3615 | 2.0455 | 1591.5
L6 | 1591.5 | 5Z | 593476 | 440533 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.5936 | 4.7067e-05 | 2.0462 | 440.5493 | 2.6167 | 2.0829 | 8.4165e-06 | 0.060989 | 5.5923 | 42.9047 | L61591.5 | 1 | 440.5493 | 2.6167 | 2.0829 | 1000000 | 593598.9918 | 47.0673 | 2.0462 | 1591.5
L7 | 1591.5 | 5Z | 691330 | 482103 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.69147 | 5.4771e-05 | 2.0467 | 482.1217 | 2.8645 | 2.0827 | 9.7738e-06 | 0.070604 | 5.6039 | 40.5714 | L71591.5 | 1 | 482.1217 | 2.8645 | 2.0827 | 1000000 | 691473.271 | 54.7714 | 2.0467 | 1591.5
L8 | 1591.5 | 5Z | 785060 | 521293 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.78522 | 6.2157e-05 | 2.047 | 521.3141 | 3.0982 | 2.0825 | 1.1077e-05 | 0.079856 | 5.6113 | 38.797 | L81591.5 | 1 | 521.3141 | 3.0982 | 2.0825 | 1000000 | 785222.6956 | 62.1567 | 2.047 | 1591.5
L9 | 1591.5 | 5Z | 880750 | 560533 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.88093 | 6.97e-05 | 2.0472 | 560.5564 | 3.3323 | 2.0823 | 1.241e-05 | 0.089331 | 5.6166 | 37.3023 | L91591.5 | 1 | 560.5564 | 3.3323 | 2.0823 | 1000000 | 880932.5263 | 69.7004 | 2.0472 | 1591.5
L10 | 1591.5 | 5Z | 988170 | 601673 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.98837 | 7.8172e-05 | 2.0474 | 601.6989 | 3.5779 | 2.0821 | 1.3907e-05 | 0.099989 | 5.621 | 35.7826 | L101591.5 | 1 | 601.6989 | 3.5779 | 2.0821 | 1000000 | 988374.788 | 78.172 | 2.0474 | 1591.5

## pyReport
| 2025 Coax | [6 empty] | 2020 S22075 Coax!
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | delta L /U | delta R /U
L10 | 10 | 101.1258 | 0.0079984 | 2.0473 | 69.8831 | 0.41533 | 2.0827 | 10 | 101.1231 | 0.0079574 | 2.0495 | 69.7374 | 0.41445 | 2.0827 | 0.32566 | 0.35069
L9 | 9 | 89.007 | 0.0070428 | 2.0472 | 64.8266 | 0.38518 | 2.0829 | 9 | 89.0051 | 0.0070067 | 2.0494 | 64.6747 | 0.38427 | 2.0829 | 0.26866 | 0.39414
L8 | 8 | 78.8976 | 0.006246 | 2.0469 | 60.2913 | 0.35816 | 2.083 | 8 | 78.8959 | 0.006214 | 2.0491 | 60.1387 | 0.35724 | 2.083 | 0.28525 | 0.42598
L7 | 7 | 68.9571 | 0.0054629 | 2.0466 | 55.7277 | 0.33099 | 2.0832 | 7 | 68.9557 | 0.0054349 | 2.0488 | 55.5799 | 0.3301 | 2.0832 | 0.25103 | 0.44662
L6 | 6 | 59.1088 | 0.0046877 | 2.0462 | 51.0214 | 0.30297 | 2.0833 | 6 | 59.107 | 0.0046637 | 2.0483 | 50.8781 | 0.30211 | 2.0833 | 0.39706 | 0.47325
L5 | 5 | 49.0614 | 0.0038979 | 2.0454 | 45.9083 | 0.27256 | 2.0835 | 5 | 49.0604 | 0.003878 | 2.0475 | 45.7702 | 0.27172 | 2.0835 | 0.24392 | 0.50685
L4 | 4 | 39.3543 | 0.0031365 | 2.0441 | 40.5575 | 0.24075 | 2.0837 | 4 | 39.355 | 0.0031207 | 2.0461 | 40.4068 | 0.23984 | 2.0837 | -0.24218 | 0.62602
L3 | 3 | 29.2713 | 0.0023494 | 2.0412 | 34.4793 | 0.20464 | 2.0838 | 3 | 29.2704 | 0.0023376 | 2.0432 | 34.3478 | 0.20384 | 2.0838 | 0.35343 | 0.64253
L2 | 2 | 19.9925 | 0.0016331 | 2.0348 | 28.0078 | 0.16622 | 2.084 | 2 | 19.9916 | 0.001625 | 2.0365 | 27.9029 | 0.16558 | 2.084 | 0.56384 | 0.63086
L1 | 1 | 10.1511 | 0.00090488 | 2.0133 | 19.4814 | 0.11565 | 2.084 | 1 | 10.1508 | 0.00090104 | 2.0143 | 19.4083 | 0.11521 | 2.084 | 0.34317 | 0.63176
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 | 10 | 988.3748 | 0.078172 | 2.0474 | 601.6989 | 3.5779 | 2.0821 | 10 | 988.3118 | 0.077768 | 2.0496 | 599.6681 | 3.5659 | 2.0821 | 0.80632 | 0.56758
L9 | 9 | 880.9325 | 0.0697 | 2.0472 | 560.5564 | 3.3323 | 2.0823 | 9 | 880.8806 | 0.069341 | 2.0494 | 558.6212 | 3.3208 | 2.0823 | 0.74488 | 0.58076
L8 | 8 | 785.2227 | 0.062157 | 2.047 | 521.3141 | 3.0982 | 2.0825 | 8 | 785.1729 | 0.061836 | 2.0492 | 519.6584 | 3.0884 | 2.0825 | 0.80152 | 0.53441
L7 | 7 | 691.4733 | 0.054771 | 2.0467 | 482.1217 | 2.8645 | 2.0827 | 7 | 691.4255 | 0.054489 | 2.0489 | 480.7646 | 2.8565 | 2.0827 | 0.8713 | 0.47379
L6 | 6 | 593.599 | 0.047067 | 2.0462 | 440.5493 | 2.6167 | 2.0829 | 6 | 593.5664 | 0.046825 | 2.0484 | 439.0086 | 2.6076 | 2.0829 | 0.69302 | 0.58879
L5 | 5 | 495.5777 | 0.039362 | 2.0455 | 396.9667 | 2.3571 | 2.0831 | 5 | 495.5522 | 0.03916 | 2.0477 | 395.4624 | 2.3482 | 2.0831 | 0.64825 | 0.63816
L4 | 4 | 399.5058 | 0.031825 | 2.0443 | 351.0639 | 2.0839 | 2.0833 | 4 | 399.4904 | 0.031663 | 2.0463 | 349.5711 | 2.0751 | 2.0833 | 0.48436 | 0.71634
L3 | 3 | 304.6541 | 0.024416 | 2.0418 | 301.982 | 1.792 | 2.0835 | 3 | 304.6368 | 0.024292 | 2.0438 | 300.8116 | 1.7851 | 2.0835 | 0.70927 | 0.6531
L2 | 2 | 207.8611 | 0.016931 | 2.0358 | 244.2186 | 1.4489 | 2.0837 | 2 | 207.8499 | 0.016847 | 2.0376 | 243.2545 | 1.4432 | 2.0837 | 0.66248 | 0.66539
L1 | 1 | 103.9495 | 0.0092128 | 2.0144 | 168.318 | 0.99854 | 2.0837 | 1 | 103.9444 | 0.0091732 | 2.0154 | 167.6371 | 0.9945 | 2.0837 | 0.55447 | 0.68197

## re2019
Prepare for full recalculation of the 2019 measurement
On 4 December 2019 Minyu did comparison of twisted vs coax using the UB
|  | Coax | [7 empty] | For python processing
|  |  | L | R | [6 empty] | xdial | rdial
5Z |  | 10 | 988108 | 599631 | 599625 | 2660 | 3225 | 3198 | 0.3236% | 0 | 0 | 6
|  | 9 | 880699 | 558587 | 558581 | 2112 | 3018 | 2991 | 0.3396% | 1 | 103923 | 167631
|  | 8 | 785011 | 519627 | 519621 | 1678 | 2304 | 2277 | 0.2901% | 2 | 207807 | 243243
|  | 7 | 691283 | 480736 | 480730 | 1296 | 1803 | 1776 | 0.2569% | 3 | 304574 | 300796
|  | 6 | 593444 | 438983 | 438977 | 969 | 1102 | 1075 | 0.1811% | 4 | 399408 | 349552
|  | 5 | 495450 | 395440 | 395434 | 672 | 934 | 907 | 0.1831% | 5 | 495450 | 395440
|  | 4 | 399408 | 349552 | 349546 | 439 | 575 | 548 | 0.1372% | 6 | 593444 | 438983
|  | 3 | 304574 | 300796 | 300790 | 251 | 452 | 425 | 0.1395% | 7 | 691283 | 480736
|  | 2 | 207807 | 243243 | 243237 | 116 | 260 | 233 | 0.1121% | 8 | 785011 | 519627
|  | 1 | 103923 | 167631 | 167625 | 28 | 95 | 68 | 0.06543% | 9 | 880699 | 558587
|  | 0 | 0 | 6 | [5 empty] | 10 | 988108 | 599631
|  | Twisted | [8 empty] | xdial | rdial
|  | 10 | 985448 | 596406 | 596427 |  |  |  |  | 0 | 0 | -21
|  | 9 | 878587 | 555569 | 555590 |  |  |  |  | 1 | 103895 | 167536
|  | 8 | 783333 | 517323 | 517344 |  |  |  |  | 2 | 207691 | 242983
|  | 7 | 689987 | 478933 | 478954 |  |  |  |  | 3 | 304323 | 300344
|  | 6 | 592475 | 437881 | 437902 |  |  |  |  | 4 | 398969 | 348977
|  | 5 | 494778 | 394506 | 394527 |  |  |  |  | 5 | 494778 | 394506
|  | 4 | 398969 | 348977 | 348998 |  |  |  |  | 6 | 592475 | 437881
|  | 3 | 304323 | 300344 | 300365 |  |  |  |  | 7 | 689987 | 478933
|  | 2 | 207691 | 242983 | 243004 |  |  |  |  | 8 | 783333 | 517323
|  | 1 | 103895 | 167536 | 167557 |  |  |  |  | 9 | 878587 | 555569
|  | 0 | 0 | -21 | [5 empty] | 10 | 985448 | 596406
4Z |  | Coax
|  |  | L | R | [6 empty] | xdial | rdial
|  | 10 | 1011026 | 697686 | 697324 | 277 | 1131 | 831 | 0.08219% | 0 | 3 | 362
|  | 9 | 889871 | 647063 | 646701 | 212 | 960 | 660 | 0.07417% | 1 | 101490 | 194431
|  | 8 | 788799 | 601706 | 601344 | 166 | 816 | 516 | 0.06542% | 2 | 199878 | 279371
|  | 7 | 689418 | 556121 | 555759 | 126 | 796 | 496 | 0.07194% | 3 | 292647 | 343815
| dislexic swap? | 6 | 590951 | 509106 | 508744 | 89 | 641 | 341 | 0.0577% | 4 | 393472 | 404401
|  | 5 | 490506 | 458031 | 457669 | 58 | 556 | 256 | 0.05219% | 5 | 490506 | 458031
|  | 4 | 393472 | 404401 | 404039 | 38 | 368 | 68 | 0.01728% | 6 | 590951 | 509106
|  | 3 | 292647 | 343815 | 343453 | 17 | 322 | 22 | 0.007518% | 7 | 689418 | 556121
|  | 2 | 199878 | 279371 | 279009 | 6 | 296 | -4 | -0.002001% | 8 | 788799 | 601706
|  | 1 | 101490 | 194431 | 194069 | 1 | 296 | -4 | -0.003941% | 9 | 889871 | 647063
|  | 0 | 3 | 362 | [5 empty] | 10 | 1011026 | 697686
|  | Twisted | [8 empty] | xdial | rdial
|  | 10 | 1010749 | 696555 | 696493 |  |  |  |  | 0 | 0 | 62
|  | 9 | 889659 | 646103 | 646041 |  |  |  |  | 1 | 101490 | 194135
|  | 8 | 788633 | 600890 | 600828 |  |  |  |  | 2 | 199872 | 279075
|  | 7 | 689292 | 555325 | 555263 |  |  |  |  | 3 | 292630 | 343493
|  | 6 | 590862 | 508465 | 508403 |  |  |  |  | 4 | 393434 | 404033
|  | 5 | 490448 | 457475 | 457413 |  |  |  |  | 5 | 490448 | 457475
|  | 4 | 393434 | 404033 | 403971 |  |  |  |  | 6 | 590862 | 508465
|  | 3 | 292630 | 343493 | 343431 |  |  |  |  | 7 | 689292 | 555325
|  | 2 | 199872 | 279075 | 279013 |  |  |  |  | 8 | 788633 | 600890
|  | 1 | 101490 | 194135 | 194073 |  |  |  |  | 9 | 889659 | 646103
|  | 0 | 0 | 62 | 0 |  |  |  |  | 10 | 1010749 | 696555

## coax_v_twist
On 3 December 2019 the effect of twisted lead vs coaxial connection were investigated.
Measurements in AI VIII
LCR
Twisted lead variability
|  |  |  | normal/position ratio |  | ppm
| normal | 984.9926 |  | 1 |  | 0
| below | 984.9485 |  | 1 |  | 44.7739
| horizontal | 984.9939 |  | 1 |  | -1.3198
| above | 984.9565 |  | 1 |  | 36.6514
| change twist |  |  | low/high twist ratio
| below | 985.4645 |  | 0.99948 |  | -523.6109
| horizontal | 985.4827 |  | 0.9995 |  | -496.0006
| above | 985.4838 |  | 0.99946 |  | -535.0671
Above data suggests that dependence of angle of lead entry is 40 ppm
Change in twist has a much larger 500 ppm shift!
| Coax connection |  |  | horizontal/postion |  |  |  | twisted horizontal/coax |  | ppm
| below | 988.2161 |  | 1 |  | 4.4525 |  | 0.99723 |  | -2765.9942
| horizontal | 988.2205 |  | 1 |  | 0 |  | 0.99723 |  | -2770.4343
| above | 988.2231 |  | 1 |  | -2.631 |  | 0.99723 |  | -2773.058
The coax connection is not sensitive to lead position.
The coax connection gives a higher inductance!
The very large change seen here might have been caused by some other change?
|  |  |  | twisted/coax |  | ppm
UB | twisted | 590859 |  | 0.99985 |  | -152.2974
|  | 1010745 |  | 0.99972 |  | -275.9578
| coax | 590949 |  |  |  |  | coax reading is higher
1011024
Note that the zero dial setting was not included in the checks.
28.04
On 4 December 2019 Minyu did comparison of twisted vs coax using the UB | [9 empty] | 1.9604
Coax
| Raw dial readings |  |  | coax/twisted |  | ppm |  | coax-twisted |  |  | calc-meas
10 | 988108 |  |  | 1.0027 |  | 2699.2799 |  | 2660 |  | 2559.6348 | -100.3652
9 | 880699 |  |  | 1.0024 |  | 2403.8598 |  | 2112 |  | 2081.9727 | -30.0273
8 | 785011 |  |  | 1.0021 |  | 2142.1286 |  | 1678 |  | 1652.7061 | -25.2939
7 | 691283 |  |  | 1.0019 |  | 1878.2963 |  | 1296 |  | 1272.0618 | -23.9382
6 | 593444 |  |  | 1.0016 |  | 1635.512 |  | 969 |  | 940.2985 | -28.7015
5 | 495450 |  |  | 1.0014 |  | 1358.1849 |  | 672 |  | 657.7166 | -14.2834
4 | 399408 |  |  | 1.0011 |  | 1100.3361 |  | 439 |  | 424.6747 | -14.3253
3 | 304574 |  |  | 1.0008 |  | 824.7816 |  | 251 |  | 241.6165 | -9.3835
2 | 207807 |  |  | 1.0006 |  | 558.522 |  | 116 |  | 109.1232 | -6.8768
1 | 103923 |  |  | 1.0003 |  | 269.5029 |  | 28 |  | 28.04 | 0.04
0 | 0
[6 empty] | 10.0158 |  | 95
Twisted
10 | 985448
9 | 878587
8 | 783333
7 | 689987
6 | 592475
5 | 494778
4 | 398969
3 | 304323
2 | 207691
1 | 103895
0 | 0
This pattern of difference might be modelled as a mutual inductance per section.
Why does the ppm difference scale with dial setting rather than being constant?
Note that coil inductance is proportional to the square of the number of turns
| 100 | turns max | 2700 | delta at max turns
| 0.133 | effective fractional turn of lead |  |  |  | 0.131
| 1.5 | power |  |  |  | 1.5
Inductance | N |  |  | N+ | total L | calculated | measured | calc-meas | N^2 | (N+)^2 | ratio
100 | 100 |  | 1 | 100.133 | 100.2662 | 0.26618 | 0.266 | 0.00017689 | 10000 | 10026.6177 | 1.0027
90 | 94.8683 |  | 0.94868 | 94.9819 | 90.2156 | 0.21559 | 0.2112 | 0.004389 | 9000 | 9021.5589 | 1.0024
80 | 89.4427 |  | 0.89443 | 89.5379 | 80.1703 | 0.17033 | 0.1678 | 0.0025306 | 8000 | 8017.0331 | 1.0021
70 | 83.666 |  | 0.83666 | 83.7439 | 70.1304 | 0.1304 | 0.1296 | 0.00080067 | 7000 | 7013.0401 | 1.0019
60 | 77.4597 |  | 0.7746 | 77.5215 | 60.0958 | 0.095798 | 0.0969 | -0.0011018 | 6000 | 6009.5798 | 1.0016
50 | 70.7107 |  | 0.70711 | 70.7577 | 50.0665 | 0.066522 | 0.0672 | -0.00067789 | 5000 | 5006.6522 | 1.0013
40 | 63.2456 |  | 0.63246 | 63.2792 | 40.0426 | 0.042571 | 0.0439 | -0.0013287 | 4000 | 4004.2571 | 1.0011
30 | 54.7723 |  | 0.54772 | 54.7941 | 30.0239 | 0.023945 | 0.0251 | -0.0011552 | 3000 | 3002.3945 | 1.0008
20 | 44.7214 |  | 0.44721 | 44.7333 | 20.0106 | 0.010641 | 0.0116 | -0.00095858 | 2000 | 2001.0641 | 1.0005
10 | 31.6228 |  | 0.31623 | 31.627 | 10.0027 | 0.0026602 | 0.0028 | -0.00013982 | 1000 | 1000.266 | 1.0003
In essence fewer turns are needed for the final step than for the first step.
This would suggest that M had proportionally more impact on the lower values.
0.01
|  |  | 100 | 100.01 | 1.0002 | [7 empty] | plot against turns rather than inductance
|  |  | 50 | 50.01 | 1.0004
|  |  | 10 | 10.01 | 1.002
Inductance | N |  |  | N+ | L
100 | 100 |  | 1 | 100.133 | 100.2662
90 | 94.8683 |  | 0.94868 | 95.0013 | 90.2525
80 | 89.4427 |  | 0.89443 | 89.5757 | 80.2381
70 | 83.666 |  | 0.83666 | 83.799 | 70.2227
60 | 77.4597 |  | 0.7746 | 77.5927 | 60.2062
50 | 70.7107 |  | 0.70711 | 70.8437 | 50.1883
40 | 63.2456 |  | 0.63246 | 63.3786 | 40.1684
30 | 54.7723 |  | 0.54772 | 54.9053 | 30.1459
20 | 44.7214 |  | 0.44721 | 44.8544 | 20.1191
10 | 31.6228 |  | 0.31623 | 31.7558 | 10.0843
|  |  | 101 | 1001
|  |  | 100 | 1000
|  |  | 1.0201 | 1.002
Now understand capacitance to be the main cause of change when using the coax cable.
| Readings |  |  | Zero correcte
| C | G |  | C | G
Short coax | 3.442e-11 | 3.91e-10 |  | 3.4418e-11 | 3.9475e-10 |  | For 1 H unit assuming nominal L and 600 ohm
twisted leads | 9.6947e-13 | 6.417e-11 |  | 9.6746e-13 | 6.792e-11
LCR zero | 2.01e-15 | -3.75e-12 |  |  |  |  | cap induced L change |  |  | coax-twisted
[7 empty] | short coax | twisted | [5 empty] | error / %
|  |  | R | [6 empty] | calc | measured
|  |  | 168.4377 | 0.28003 | 168.0187 | 1 | 3.3459e-05 | 9.402e-07 |  | 3.2519e-05 | 3.1285e-05 | 0.96207 | 31.2855 | 0.003346%
|  |  | 244.4601 | 0.40642 | 243.8521 | 2 | 0.00013572 | 3.8125e-06 |  | 0.00013191 | 0.0001305 | 0.98931 | 130.5014 | 0.006786%
|  |  | 302.212 | 0.50243 | 301.4603 | 3 | 0.00030696 | 8.6198e-06 |  | 0.00029834 | 0.00028345 | 0.95009 | 283.4527 | 0.01023%
|  |  | 351.3596 | 0.58414 | 350.4856 | 4 | 0.00054723 | 1.5362e-05 |  | 0.00053187 | 0.00049125 | 0.92362 | 491.2475 | 0.01368%
|  |  | 397.4338 | 0.66074 | 396.4452 | 5 | 0.00085655 | 2.4037e-05 |  | 0.00083251 | 0.00075915 | 0.91188 | 759.1475 | 0.01713%
|  |  | 441.3742 | 0.73379 | 440.2763 | 6 | 0.001235 | 3.4644e-05 |  | 0.0012003 | 0.0010919 | 0.90964 | 1091.8631 | 0.02058%
|  |  | 483.0297 | 0.80305 | 481.8282 | 7 | 0.0016826 | 4.7186e-05 |  | 0.0016354 | 0.001485 | 0.908 | 1484.962 | 0.02404%
|  |  | 521.6111 | 0.86719 | 520.3137 | 8 | 0.0021996 | 6.1663e-05 |  | 0.0021379 | 0.0019185 | 0.89739 | 1918.53 | 0.02749%
|  |  | 560.1833 | 0.93132 | 558.7899 | 9 | 0.0027858 | 7.8072e-05 |  | 0.0027078 | 0.0024168 | 0.89256 | 2416.8473 | 0.03095%
|  |  | 601.4962 | 1 | 600 | 10 | 0.0034414 | 9.6411e-05 |  | 0.003345 | 0.0030448 | 0.91028 | 3044.8386 | 0.03441%
[9 empty] | 10 | 0.003345
[9 empty] | 9 | 0.0027078
[9 empty] | 8 | 0.0021379
[9 empty] | 7 | 0.0016354
[9 empty] | 6 | 0.0012003
[9 empty] | 5 | 0.00083251
[9 empty] | 4 | 0.00053187
[9 empty] | 3 | 0.00029834
[9 empty] | 2 | 0.00013191
[9 empty] | 1 | 3.2519e-05

## LHistory
History
[8 empty] | corrected for capacitance | [7 empty] | 40585
[21 empty] | as ratio | Typical U
|  | S20803 | S21193 | S21431 | S21637 | S21784 | S21960 | S22075 | [11 empty] | twist
| Nominal | 1992-09-03 | 1997-04-03 | 2001-09-14 | 2006-06-15 | 2011-02-11 | 2015-08-20 | 2019-12-04 | 2024-12-20 |  |  | 1992-09-03 | 1997-04-03 | 2001-09-14 | 2006-06-15 | 2011-02-11 | 2015-08-20 | 2019-12-04 | 2024-12-20 | 2025-11-30
| 100 | 101.1 | 101.0931 | 101.0775 | 101.0768 | 101.0934 | 101.1144 | 101.1231 | 101.1258 |  |  | 1 | 0.99993 | 0.99978 | 0.99977 | 0.99993 | 1.0001 | 1.0002 | 1.0003 |  | 6.9238e-05 | 0.0070
| 90 | 89 | 88.9737 | 88.9669 | 88.9616 | 88.9860 | 89.0063 | 89.0051 | 89.007 |  |  | 1 | 0.9997 | 0.99963 | 0.99957 | 0.99984 | 1.0001 | 1.0001 | 1.0001 |  | 7.3034e-05 | 0.0065
| 80 | 78.9 | 78.8686 | 78.8643 | 78.8583 | 78.8826 | 78.9009 | 78.8959 | 78.8976 |  |  | 1 | 0.9996 | 0.99955 | 0.99947 | 0.99978 | 1 | 0.99995 | 0.99997 |  | 7.4778e-05 | 0.0059
| 70 | 68.9 | 68.933 | 68.9293 | 68.9246 | 68.9465 | 68.9631 | 68.9557 | 68.9571 |  |  | 1 | 1.0005 | 1.0004 | 1.0004 | 1.0007 | 1.0009 | 1.0008 | 1.0008 |  | 7.5472e-05 | 0.0052
| 60 | 59.1 | 59.0888 | 59.0861 | 59.082 | 59.1017 | 59.1163 | 59.107 | 59.1088 |  |  | 1 | 0.99981 | 0.99976 | 0.9997 | 1 | 1.0003 | 1.0001 | 1.0001 |  | 7.6142e-05 | 0.0045
| 50 | 49.05 | 49.0463 | 49.0445 | 49.0405 | 49.0579 | 49.0702 | 49.0604 | 49.0614 |  |  | 1 | 0.99993 | 0.99989 | 0.99981 | 1.0002 | 1.0004 | 1.0002 | 1.0002 |  | 7.7472e-05 | 0.0038
| 40 | 39.35 | 39.3448 | 39.3434 | 39.3402 | 39.3546 | 39.3650 | 39.355 | 39.3543 |  |  | 1 | 0.99987 | 0.99983 | 0.99975 | 1.0001 | 1.0004 | 1.0001 | 1.0001 |  | 7.878e-05 | 0.0031
| 30 | 29.26 | 29.2635 | 29.2625 | 29.2602 | 29.2715 | 29.2794 | 29.2704 | 29.2713 |  |  | 1 | 1.0001 | 1.0001 | 1 | 1.0004 | 1.0007 | 1.0004 | 1.0004 |  | 7.8606e-05 | 0.0023
| 20 | 19.99 | 19.9873 | 19.9867 | 19.985 | 19.9931 | 19.9988 | 19.9916 | 19.9925 |  |  | 1 | 0.99987 | 0.99983 | 0.99975 | 1.0002 | 1.0004 | 1.0001 | 1.0001 |  | 8.004e-05 | 0.0016
| 10 | 10.15 | 10.1489 | 10.1486 | 10.1477 | 10.1520 | 10.1550 | 10.1508 | 10.1511 |  |  | 1 | 0.9999 | 0.99986 | 0.99977 | 1.0002 | 1.0005 | 1.0001 | 1.0001 |  | 8.867e-05 | 0.0009
| 1000 | 985 | 986.0623 | 985.2113 | 985.1001 | 984.939 | 985.244 | 984.9668 | 985.0298 | 0.003345 | 985.1441 | 1 | 1.0011 | 1.0002 | 1.0001 | 0.99994 | 1.0002 | 0.99997 | 1 | 1.0001 | 7.0051e-05 | 0.069
| 900 | 879 | 879.06 | 878.3959 | 878.2876 | 878.232 | 878.380 | 878.1728 | 878.2248 | 0.0027078 | 878.352 | 1 | 1.0001 | 0.99931 | 0.99919 | 0.99913 | 0.99929 | 0.99906 | 0.99912 | 0.99926 | 6.8259e-05 | 0.060
| 800 | 783 | 783.7116 | 783.1934 | 783.0962 | 783.086 | 783.228 | 783.035 | 783.0848 | 0.0021379 | 783.1533 | 1 | 1.0009 | 1.0002 | 1.0001 | 1.0001 | 1.0003 | 1 | 1.0001 | 1.0002 | 6.7688e-05 | 0.053
| 700 | 690 | 690.2875 | 689.8752 | 689.8051 | 689.813 | 689.929 | 689.7901 | 689.8379 | 0.0016354 | 689.8499 | 1 | 1.0004 | 0.99982 | 0.99972 | 0.99973 | 0.9999 | 0.9997 | 0.99977 | 0.99978 | 6.8116e-05 | 0.047
| 600 | 592 | 592.7167 | 592.4074 | 592.3583 | 592.360 | 592.471 | 592.366 | 592.3987 | 0.0012003 | 592.3817 | 1 | 1.0012 | 1.0007 | 1.0006 | 1.0006 | 1.0008 | 1.0006 | 1.0007 | 1.0006 | 6.7568e-05 | 0.040
| 500 | 494.8 | 494.9729 | 494.7435 | 494.7064 | 494.710 | 494.803 | 494.7197 | 494.7452 | 0.00083251 | 494.7115 | 1 | 1.0003 | 0.99989 | 0.99981 | 0.99982 | 1 | 0.99984 | 0.99989 | 0.99982 | 6.8715e-05 | 0.034
| 400 | 399 | 399.1135 | 398.9589 | 398.9299 | 398.931 | 399.008 | 398.9585 | 398.9739 | 0.00053187 | 398.9307 | 1 | 1.0003 | 0.9999 | 0.99982 | 0.99983 | 1 | 0.9999 | 0.99993 | 0.99983 | 6.7669e-05 | 0.027
| 300 | 304.4 | 304.4292 | 304.3255 | 304.3075 | 304.305 | 304.371 | 304.3385 | 304.3558 | 0.00029834 | 304.3051 | 1 | 1.0001 | 0.99976 | 0.9997 | 0.99969 | 0.9999 | 0.9998 | 0.99985 | 0.99969 | 6.8988e-05 | 0.021
| 200 | 207.7 | 207.7596 | 207.6978 | 207.6929 | 207.691 | 207.733 | 207.7179 | 207.7292 | 0.00013191 | 207.682 | 1 | 1.0003 | 0.99999 | 0.99997 | 0.99996 | 1.0002 | 1.0001 | 1.0001 | 0.99991 | 6.7405e-05 | 0.014
| 100 | 103.9 | 103.9278 | 103.9029 | 103.9039 | 103.900 | 103.923 | 103.9119 | 103.917 | 3.2519e-05 | 103.8955 | 1 | 1.0003 | 1 | 1 | 1 | 1.0002 | 1.0001 | 1.0002 | 0.99996 | 7.6997e-05 | 0.008
Note on D16: Christian: Actually UB read 985.254

## exampleU
Budget for L8 1591.5
ureal(60.08704790830853,0.1713140491986457,20.451673221652594)
uuttempL8 | 0.12049
temperature | 0.12043
Amp_phase_err | 0.013329
rtempcoL8 | 0.012019
Adialresureal(62.0,10.0,10.0) | 0.0011776
Adialresureal(600890.0,10.0,10.0) | 0.0011776
Adial_lin62 | 0.0010001
Adial_lin600890 | 0.0010001
Amperr | 0.00084121
R4Aacdc | 0.00015022
G1acdc | 0.00015022
G2acdc | 0.00015022
G1alpha | 6.0087e-05
G1stab | 6.0087e-05
R4Aalpha | 6.0087e-05
G2alpha | 6.0087e-05
R4Astab | 6.0087e-05
G2stab | 6.0087e-05
cal_temp | 4.2488e-05
G2err | 3.0043e-05
G1err | 3.0043e-05
R4Aerr | 3.0042e-05
uuttempL0 | 1.2432e-05
rtempcoL0 | 1.2401e-06
C1err | 0
C1alpha | 0
C1stab | 0
Bdial_lin0 | 0
Bdialresureal(0.0,1.0,10.0, label='Bdial_lin0') | 0
Bdial_lin788633 | 0
Bdialresureal(788633.0,1.0,10.0) | 0
ureal(0.0788795386314938,3.031875053076101e-06,27.783955120984952)
uuttempL8 | 1.9769e-06
temperature | 1.9211e-06
Amperr | 1.1043e-06
C1stab | 3.944e-07
R4Aacdc | 1.972e-07
G2acdc | 1.972e-07
xtempcoL8 | 1.972e-07
Bdialres0.0 | 1.1778e-07
Bdialres788633.0 | 1.1778e-07
Amp_phase_err | 1.0161e-07
Bdial_lin0 | 1.0002e-07
Bdial_lin788633 | 1.0002e-07
C1alpha | 7.888e-08
R4Aalpha | 7.888e-08
G2alpha | 7.888e-08
R4Astab | 7.888e-08
G2stab | 7.888e-08
C1err | 7.8868e-08
cal_temp | 5.5776e-08
G2err | 3.9439e-08
R4Aerr | 3.9438e-08
G1err | 0
G1acdc | 0
G1alpha | 0
G1stab | 0
Adial_lin62 | 0
Adialres62.0 | 0
uuttempL0 | 0
xtempcoL0 | 0
Adial_lin600890 | 0
Adialres600890.0 | 0

## Room Conditions
2017-12-04 16:30 | 20.9 | 53.6 |  | RH correction | -2.128 | from Y:\Quality\TECHPROC\E058-Lab Air Cond\loggercalibration_2017.xls
2017-12-04 16:35 | 20.9 | 53.8 | Temperature | RH
2017-12-04 16:40 | 20.9 | 54.8 | #REF! | 52.2805 | 50.1525
2017-12-04 16:45 | 20.9 | 54.4 | #REF! | 1.1136
2017-12-04 16:50 | 20.9 | 54.3
2017-12-04 16:55 | 20.8 | 54.2 | from GMH
2017-12-04 17:00 | 20.8 | 53.6
2017-12-04 17:05 | 20.9 | 53.4
2017-12-04 17:10 | 20.9 | 53.3
2017-12-04 17:15 | 20.8 | 53.7
2017-12-04 17:20 | 20.8 | 53
2017-12-04 17:25 | 20.8 | 52.6
2017-12-04 17:30 | 20.9 | 52.4
2017-12-04 17:35 | 20.9 | 52
2017-12-04 17:40 | 20.9 | 52
2017-12-04 17:45 | 20.9 | 52.1
2017-12-04 17:50 | 20.9 | 52
2017-12-04 17:55 | 20.9 | 51.8
2017-12-04 18:00 | 20.9 | 52.1
2017-12-04 18:05 | 20.9 | 52.1
2017-12-04 18:10 | 20.9 | 52
2017-12-04 18:15 | 20.9 | 52.3
2017-12-04 18:20 | 20.9 | 52.3
2017-12-04 18:25 | 20.9 | 52
2017-12-04 18:30 | 21 | 51.9
2017-12-04 18:35 | 21 | 52
2017-12-04 18:40 | 20.9 | 51.9
2017-12-04 18:45 | 20.9 | 51.9
2017-12-04 18:50 | 20.9 | 51.8
2017-12-04 18:55 | 20.9 | 51.8
2017-12-04 19:00 | 20.9 | 51.8
2017-12-04 19:05 | 20.9 | 51.8
2017-12-04 19:10 | 21 | 51.6
2017-12-04 19:15 | 21 | 51.7
2017-12-04 19:20 | 21 | 51.6
2017-12-04 19:25 | 21 | 51.7
2017-12-04 19:30 | 21 | 51.7
2017-12-04 19:35 | 21 | 51.7
2017-12-04 19:40 | 21 | 51.9
2017-12-04 19:45 | 21 | 51.9
2017-12-04 19:50 | 21 | 51.6
2017-12-04 19:55 | 21 | 51.7
2017-12-04 20:00 | 21 | 51.7
2017-12-04 20:05 | 21 | 51.7
2017-12-04 20:10 | 21 | 51.6
2017-12-04 20:15 | 21 | 51.9
2017-12-04 20:20 | 21 | 51.9
2017-12-04 20:25 | 21 | 51.9
2017-12-04 20:30 | 21 | 51.8
2017-12-04 20:35 | 21 | 51.7
2017-12-04 20:40 | 21.1 | 51.6
2017-12-04 20:45 | 21.1 | 51.8
2017-12-04 20:50 | 21 | 52
2017-12-04 20:55 | 21 | 52.2
2017-12-04 21:00 | 21 | 52.2
2017-12-04 21:05 | 21 | 52
2017-12-04 21:10 | 21 | 51.9
2017-12-04 21:15 | 21 | 52.1
2017-12-04 21:20 | 21.1 | 52
2017-12-04 21:25 | 21 | 52.2
2017-12-04 21:30 | 21 | 51.9
2017-12-04 21:35 | 21 | 52
2017-12-04 21:40 | 21 | 51.8
2017-12-04 21:45 | 21.1 | 51.5
2017-12-04 21:50 | 21 | 51.7
2017-12-04 21:55 | 21.1 | 51.6
2017-12-04 22:00 | 21.1 | 51.5
2017-12-04 22:05 | 21.1 | 51.3
2017-12-04 22:10 | 21.1 | 51.5
2017-12-04 22:15 | 21 | 51.7
2017-12-04 22:20 | 21.1 | 51.6
2017-12-04 22:25 | 21.1 | 52
2017-12-04 22:30 | 21.1 | 51.5
2017-12-04 22:35 | 21.1 | 51.4
2017-12-04 22:40 | 21.1 | 51.2
2017-12-04 22:45 | 21.1 | 51.2
2017-12-04 22:50 | 21.1 | 51.3
2017-12-04 22:55 | 21.1 | 51.4
2017-12-04 23:00 | 21.1 | 51.3
2017-12-04 23:05 | 21.1 | 51.3
2017-12-04 23:10 | 21.2 | 51.4
2017-12-04 23:15 | 21.2 | 51.3
2017-12-04 23:20 | 21.2 | 51.5
2017-12-04 23:25 | 21.1 | 51.7
2017-12-04 23:30 | 21.1 | 51.7
2017-12-04 23:35 | 21.1 | 51.3
2017-12-04 23:40 | 21.1 | 51.5
2017-12-04 23:45 | 21.1 | 51.4
2017-12-04 23:50 | 21.1 | 51.4
2017-12-04 23:55 | 21.1 | 51.3
2017-12-05 | 21.1 | 51.4
2017-12-05 00:05 | 21.1 | 51.3
2017-12-05 00:10 | 21.1 | 51
2017-12-05 00:15 | 21.1 | 51.3
2017-12-05 00:20 | 21.1 | 51.1
2017-12-05 00:25 | 21.2 | 51.2
2017-12-05 00:30 | 21.1 | 51.4
2017-12-05 00:35 | 21.2 | 51.3
2017-12-05 00:40 | 21.1 | 51.5
2017-12-05 00:45 | 21.2 | 51.3
2017-12-05 00:50 | 21.2 | 51.3
2017-12-05 00:55 | 21.1 | 51.3
2017-12-05 01:00 | 21.2 | 51.3
2017-12-05 01:05 | 21.1 | 51.4
2017-12-05 01:10 | 21.2 | 51.2
2017-12-05 01:15 | 21.2 | 51.1
2017-12-05 01:20 | 21.1 | 51.3
2017-12-05 01:25 | 21.2 | 51.9
2017-12-05 01:30 | 21.2 | 51.9
2017-12-05 01:35 | 21.1 | 52
2017-12-05 01:40 | 21.1 | 51.8
2017-12-05 01:45 | 21.2 | 51.8
2017-12-05 01:50 | 21.1 | 51.8
2017-12-05 01:55 | 21.1 | 51.9
2017-12-05 02:00 | 21.1 | 52.1
2017-12-05 02:05 | 21.1 | 51.9
2017-12-05 02:10 | 21.2 | 51.8
2017-12-05 02:15 | 21.2 | 51.7
2017-12-05 02:20 | 21.2 | 51.9
2017-12-05 02:25 | 21.2 | 51.9
2017-12-05 02:30 | 21.2 | 51.8
2017-12-05 02:35 | 21.1 | 51.6
2017-12-05 02:40 | 21.2 | 51.7
2017-12-05 02:45 | 21.2 | 51.6
2017-12-05 02:50 | 21.1 | 51.7
2017-12-05 02:55 | 21.1 | 51.6
2017-12-05 03:00 | 21.1 | 51.6
2017-12-05 03:05 | 21.1 | 51.6
2017-12-05 03:10 | 21.1 | 51.3
2017-12-05 03:15 | 21.2 | 51.1
2017-12-05 03:20 | 21.1 | 51.3
2017-12-05 03:25 | 21.1 | 51.6
2017-12-05 03:30 | 21.1 | 51.7
2017-12-05 03:35 | 21.1 | 51.6
2017-12-05 03:40 | 21.2 | 51.5
2017-12-05 03:45 | 21.1 | 51.3
2017-12-05 03:50 | 21.1 | 51.2
2017-12-05 03:55 | 21.1 | 51.2
2017-12-05 04:00 | 21.2 | 51
2017-12-05 04:05 | 21.2 | 51.1
2017-12-05 04:10 | 21.2 | 51.4
2017-12-05 04:15 | 21.2 | 51.6
2017-12-05 04:20 | 21.1 | 51.5
2017-12-05 04:25 | 21.1 | 51.8
2017-12-05 04:30 | 21.2 | 51.7
2017-12-05 04:35 | 21.1 | 52
2017-12-05 04:40 | 21.1 | 51.9
2017-12-05 04:45 | 21.2 | 52
2017-12-05 04:50 | 21.1 | 52
2017-12-05 04:55 | 21.2 | 51.8
2017-12-05 05:00 | 21.2 | 51.8
2017-12-05 05:05 | 21.1 | 51.9
2017-12-05 05:10 | 21.1 | 51.8
2017-12-05 05:15 | 21.1 | 51.6
2017-12-05 05:20 | 21.2 | 51.7
2017-12-05 05:25 | 21.2 | 51.8
2017-12-05 05:30 | 21.2 | 51.7
2017-12-05 05:35 | 21.1 | 51.8
2017-12-05 05:40 | 21.1 | 51.5
2017-12-05 05:45 | 21.1 | 51.4
2017-12-05 05:50 | 21.1 | 51.5
2017-12-05 05:55 | 21.1 | 51.4
2017-12-05 06:00 | 21.1 | 51.5
2017-12-05 06:05 | 21.1 | 51.4
2017-12-05 06:10 | 21.2 | 51.3
2017-12-05 06:15 | 21.2 | 51.2
2017-12-05 06:20 | 21.2 | 51.2
2017-12-05 06:25 | 21.2 | 51.2
2017-12-05 06:30 | 21.2 | 51.3
2017-12-05 06:35 | 21.1 | 51.3
2017-12-05 06:40 | 21.2 | 51.1
2017-12-05 06:45 | 21.2 | 51.4
2017-12-05 06:50 | 21.2 | 51.6
2017-12-05 06:55 | 21.1 | 51.4
2017-12-05 07:00 | 21.2 | 51.4
2017-12-05 07:05 | 21.1 | 51.3
2017-12-05 07:10 | 21.2 | 51.3
2017-12-05 07:15 | 21.2 | 51.3
2017-12-05 07:20 | 21.1 | 51.4
2017-12-05 07:25 | 21.1 | 51.4
2017-12-05 07:30 | 21.1 | 51.3
2017-12-05 07:35 | 21.1 | 51.3
2017-12-05 07:40 | 21.1 | 51.3
2017-12-05 07:45 | 21.1 | 51.4
2017-12-05 07:50 | 21.2 | 51.3
2017-12-05 07:55 | 21.1 | 51.3
2017-12-05 08:00 | 21.1 | 51.4
2017-12-05 08:05 | 21.1 | 51.4
2017-12-05 08:10 | 21.1 | 51.4
2017-12-05 08:15 | 21.2 | 51.5
2017-12-05 08:20 | 21.2 | 51.6
2017-12-05 08:25 | 21 | 51.8
2017-12-05 08:30 | 21 | 51.8
2017-12-05 08:35 | 21 | 51.8
2017-12-05 08:40 | 21.1 | 51.6
2017-12-05 08:45 | 21 | 51.8
2017-12-05 08:50 | 21 | 51.7
2017-12-05 08:55 | 21 | 51.9
2017-12-05 09:00 | 21.1 | 52.1
2017-12-05 09:05 | 21 | 52.2
2017-12-05 09:10 | 21 | 52.1
2017-12-05 09:15 | 21 | 52
2017-12-05 09:20 | 21 | 52
2017-12-05 09:25 | 21 | 52.1
2017-12-05 09:30 | 21 | 52.1
2017-12-05 09:35 | 21 | 52.3
2017-12-05 09:40 | 20.9 | 52
2017-12-05 09:45 | 21 | 52
2017-12-05 09:50 | 20.9 | 51.8
2017-12-05 09:55 | 21 | 51.6
2017-12-05 10:00 | 20.9 | 51.7
2017-12-05 10:05 | 21 | 51.8
2017-12-05 10:10 | 21 | 51.7
2017-12-05 10:15 | 20.9 | 51.9
2017-12-05 10:20 | 20.9 | 51.9
2017-12-05 10:25 | 20.9 | 51.7
2017-12-05 10:30 | 20.8 | 51.9
2017-12-05 10:35 | 20.9 | 51.9
2017-12-05 10:40 | 20.9 | 52.3
2017-12-05 10:45 | 20.9 | 52.2
2017-12-05 10:50 | 20.8 | 52.4
2017-12-05 10:55 | 20.8 | 52.8
2017-12-05 11:00 | 20.9 | 52.4
2017-12-05 11:05 | 20.8 | 52.5
2017-12-05 11:10 | 20.9 | 52.8
2017-12-05 11:15 | 20.8 | 52.6
2017-12-05 11:20 | 20.8 | 52.7
2017-12-05 11:25 | 20.8 | 52.2
2017-12-05 11:30 | 20.8 | 52.5
2017-12-05 11:35 | 20.8 | 52.4
2017-12-05 11:40 | 20.8 | 52.4
2017-12-05 11:45 | 20.8 | 52.9
2017-12-05 11:50 | 20.8 | 52.7
2017-12-05 11:55 | 20.8 | 52.7
2017-12-05 12:00 | 20.8 | 52.9
2017-12-05 12:05 | 20.9 | 52.9
2017-12-05 12:10 | 20.9 | 53.1
2017-12-05 12:15 | 20.9 | 52.8
2017-12-05 12:20 | 20.9 | 52.9
2017-12-05 12:25 | 20.9 | 53.5
2017-12-05 12:30 | 20.9 | 53.2
2017-12-05 12:35 | 20.9 | 53.2
2017-12-05 12:40 | 20.8 | 53.4
2017-12-05 12:45 | 20.8 | 53.2
2017-12-05 12:50 | 20.9 | 53.7
2017-12-05 12:55 | 20.9 | 53.7
2017-12-05 13:00 | 20.9 | 53.2
2017-12-05 13:05 | 21 | 53.4
2017-12-05 13:10 | 20.9 | 53.8
2017-12-05 13:15 | 20.8 | 54.1
2017-12-05 13:20 | 20.8 | 54
2017-12-05 13:25 | 20.8 | 54
2017-12-05 13:30 | 20.8 | 54.2
2017-12-05 13:35 | 20.8 | 54.2
2017-12-05 13:40 | 20.8 | 54.5
2017-12-05 13:45 | 20.8 | 54.5
2017-12-05 13:50 | 20.8 | 54.4
2017-12-05 13:55 | 20.8 | 54.6
2017-12-05 14:00 | 20.7 | 54.7
2017-12-05 14:05 | 20.8 | 54.9
2017-12-05 14:10 | 20.8 | 54.9
2017-12-05 14:15 | 20.7 | 55.1
2017-12-05 14:20 | 20.7 | 55.6
2017-12-05 14:25 | 20.7 | 55.2
2017-12-05 14:30 | 20.8 | 55.3
2017-12-05 14:35 | 20.8 | 54.6
2017-12-05 14:40 | 20.7 | 54.8
2017-12-05 14:45 | 20.7 | 54.9
2017-12-05 14:50 | 20.7 | 54.7
2017-12-05 14:55 | 20.7 | 54.8
2017-12-05 15:00 | 20.8 | 54.7
2017-12-05 15:05 | 20.8 | 54.4
2017-12-05 15:10 | 20.7 | 54.5
2017-12-05 15:15 | 20.8 | 54.6
2017-12-05 15:20 | 20.8 | 54.6
2017-12-05 15:25 | 20.8 | 54.2
2017-12-05 15:30 | 20.8 | 54.8
2017-12-05 15:35 | 20.8 | 54.2
2017-12-05 15:40 | 20.7 | 54
2017-12-05 15:45 | 20.7 | 54.8
2017-12-05 15:50 | 20.7 | 54.3
2017-12-05 15:55 | 20.7 | 53.8
2017-12-05 16:00 | 20.7 | 53.8
2017-12-05 16:05 | 20.6 | 54.6
2017-12-05 16:10 | 20.6 | 54
2017-12-05 16:15 | 20.7 | 54.2
2017-12-05 16:20 | 20.7 | 54.4
2017-12-05 16:25 | 20.6 | 54.3
2017-12-05 16:30 | 20.7 | 54.2
2017-12-05 16:35 | 20.7 | 54.1
2017-12-05 16:40 | 20.7 | 53.7
2017-12-05 16:45 | 20.7 | 53.9

## Conditions
| Temp |  | RH
All GMH5I readings | #REF! | #REF!
Room | #REF! | #REF! | #REF! | #REF!