## Commentary
Start with spreadsheet S21740 from the July 2009 calibration.
Remove all repeats from Results sheet
Temperature correction updated for GMH51
Measured frequency included, but not really necessary
Results transcribed from AI VII
Room data logger information for Primary Lab insered in 'Room data logger' sheet
Replace 'Universal Bridge' page with the same from S21942b, change zeros block and reference.
Insert uncertainty sheet from S21942b, with subsequent modifications but not to the UB uncertainty
The copper temperature coefficient is identical for all, but I have left the parameter table that would allow variation by component.
Add worksheet 'Histroical' to give a cleaner graphical view of previous calibrations, but some of the data is still picked up off the 'Previous History' worksheet.
Resolution term for L in uncertainty, cell B40, has been made consistent with the CMC formula.
On results page there is an analysis of HP4284 measurements designed to cross-check that the 600 microH result is reasonable (oscillation problem).
Restart using S21952, 2014, but update to interact with Python for analysis.
Enter lab book readings in 'pyUBreadings' worksheet and check the temperautre coefficients entered for each line.
Run sullivan.py to create a set of results in sullivan_Results.xlsx, being sure to use the latest ubdict_date.csv for the universal bridge calibration.
Paste these results into the uncoloured part of the pyUBresults worksheet.
Final report appears in the pyReport worksheet.
May2025
All worksheet in this job marked as this cell
Job | MSLJob3176.
S_Number | S22167
Historical Analysing in “Historical!”

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | sullivan |  | 4: Client; 5: Babcock; 10: Enter, say, 5 ppm as 5e-6
Date from | 2025-05-15 |  | 4: Job number; 5: MSLJob3176; 10: 1 sigma uncertainties from specifications
to | 2026-05-16 |  | 4: S Number; 5: S22167
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
coax zero | 1591.5 | 3Z | -1 | -10 | 1592.54 | 20.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A1 | 1591.5 | 3Z | 99364 | 571520 | 1592.34 | 20.88 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A2 | 1591.5 | 3Z | 197974 | 841874 | 1592.19 | 20.87 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A3 | 1591.5 | 3Z | 296819 | 1070422 | 1593.41 | 20.65 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A4 | 1591.5 | 3Z | 394747 | 1269967 | 1592.77 | 20.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A5 | 1591.5 | 3Z | 493101 | 1447708 | 1592.47 | 20.82 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A6 | 1591.5 | 3Z | 591332 | 1617418 | 1591.8 | 20.66 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A7 | 1591.5 | 3Z | 688992 | 1782538 | 1591.74 | 20.62 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A8 | 1591.5 | 3Z | 786206 | 1934775 | 1591.67 | 20.68 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A9 | 1591.5 | 3Z | 883126 | 2081205 | 1591.51 | 20.72 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A10 | 1591.5 | 3Z | 980165 | 2229664 | 1592.88 | 20.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
coax zero | 1591.5 | 2Z | -3 | -12 | 1593.08 | 20.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B1 | 1591.5 | 2Z | 99109 | 1168727 | 1593.85 | 20.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B2 | 1591.5 | 2Z | 198937 | 1743376 | 1593.64 | 20.72 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B3 | 1591.5 | 2Z | 298138 | 2213587 | 1593.08 | 20.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B4 | 1591.5 | 2Z | 396394 | 2803515 | 1593.56 | 20.68 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B5 | 1591.5 | 2Z | 494345 | 3190209 | 1593.49 | 20.84 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B6 | 1591.5 | 2Z | 593381 | 3553499 | 1593.17 | 20.85 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B7 | 1591.5 | 2Z | 691176 | 3910839 | 1593.32 | 20.86 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B8 | 1591.5 | 2Z | 788907 | 4213640 | 1593.89 | 20.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B9 | 1591.5 | 2Z | 887110 | 4532510 | 1593.65 | 20.75 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B10 | 1591.5 | 2Z | 984585 | 4826330 | 1595.54 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
20.7064
0.21849

## freq_coeff
Some checks were made on the frequency sensitivity of the impedance measurements.
This follows from observing that the Muirhead decade inductors showed a high, but manageable, variation of resistance with frequency. (Up to 70 ppm/Hz for resistance and up to 2.5 ppm/Hz for inductance.)
Conclusion is that the frequency dependence is not significant for the Sullivan inductors | [16 empty] | A10 R |  | A10 L
Assume effect is largest on the higher value inductor | [15 empty] | 1432.4 | 22.2254 | 5.3747e-05 | 0.0098071 | 1.088e-08
A manual LCR check was made on 13/6/25 in AI VIII | [15 empty] | 1512.4 | 22.2357 | 0.00013119 | 0.009806 | 8.2022e-09
[16 empty] | 1591.5 | 22.2456 | 8.6568e-05 | 0.0098049 | 1.0338e-08
| f | L | R | [12 empty] | 1670.8 | 22.2551 | 0.00011834 | 0.0098039 | 7.5229e-09
| kHz | mH | Ω | [12 empty] | 1749.6 | 22.2646 | 0.00041727 | 0.0098028 | 3.797e-08
| 1.592 | 9.8046 | 22.2332 | [13 empty] | B6 R |  | B6 L
| 1.67 | 9.8036 | 22.2426 | [12 empty] | 1432.4 | 3.5418 | 1.4197e-05 | 0.00059357 | 1.5776e-09
| 1.592 | 9.8046 | 22.2345 | rel delta f | rel delta L | rel delta R | [9 empty] | 1512.4 | 3.5423 | 1.2811e-05 | 0.00059354 | 2.0471e-09
|  |  |  | % | ppm | ppm | [9 empty] | 1591.5 | 3.5427 | 2.0027e-05 | 0.0005935 | 1.5738e-09
average | 1.592 | 9.8046 | 22.2339 | 4.8995 | -104.1344 | 393.9935 | [9 empty] | 1670.8 | 3.5431 | 1.9316e-05 | 0.00059347 | 1.4722e-09
[5 empty] | -1.3082 | 4.9497 | in ppm/Hz | [8 empty] | 1749.6 | 3.5435 | 7.1406e-05 | 0.00059344 | 6.6733e-09
quick assumption is that the deltas will be at least ten times smaller for the actual frequency deviation. | [14 empty] | Muirhead used ppm/Hz
[17 empty] | R |  | L
| target f | 1591.5494 | [14 empty] | ppm/Hz |  | ppm/Hz
|  | 1592.54 |  | 0.06224% | [11 empty] | A10 | 5.5337 |  | -1.383
|  | 1592.34 |  | 0.04967% | [11 empty] | B6 | 1.5254 |  | -0.68239
|  | 1592.19 |  | 0.04025%
|  | 1593.41 |  | 0.1169%
|  | 1592.77 |  | 0.07669%
|  | 1592.47 |  | 0.05784%
|  | 1591.8 |  | 0.01574%
|  | 1591.74 |  | 0.01197%
|  | 1591.67 |  | 0.007576%
|  | 1591.51 |  | -0.002478%
|  | 1592.88 |  | 0.0836%
|  | 1593.85 |  | 0.1445%
|  | 1593.64 |  | 0.1314%
|  | 1593.08 |  | 0.09617%
|  | 1593.56 |  | 0.1263%
|  | 1593.49 |  | 0.1219%
|  | 1593.17 |  | 0.1018%
|  | 1593.32 |  | 0.1112%
|  | 1593.89 |  | 0.1471%
|  | 1593.65 |  | 0.132%
| mean | 1592.8485 |  | 0.08162%
| sd | 0.77511
| rel sd | 0.04866%
From E4980_02_2025-06-18.csv gathered on lab machine
Items A10 and B6 were measured
[7 empty] | lcr name: | KeySight E4980A
[7 empty] | Name and description of Device under test: | Sullivan A10 Coaxial GMH3I_2043
[7 empty] | GPIB source: | <bound method MessageBasedResource.read of <'GPIBInstrument'('GPIB0::17::INSTR')>>
[7 empty] | 2025-06-18 | 16:02
[7 empty] | select the freq range 90% - 110% | [1432.4, 1512.4, 1591.5, 1670.8, 1749.6]
A10 | [6 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2254
B6 | [6 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2253
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2254
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | 1432.4 | 0.0098071 | 22.2255
[7 empty] | Freq | 1432.4 | Rs | 22.2254 | 5.3747e-05 | Ls | 0.0098071 | 1.088e-08
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1432;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1512.4 | 0.009806 | 22.2356
[7 empty] | 1512.4 | 0.009806 | 22.2356
[7 empty] | 1512.4 | 0.009806 | 22.2356
[7 empty] | 1512.4 | 0.009806 | 22.2357
[7 empty] | 1512.4 | 0.009806 | 22.2358
[7 empty] | 1512.4 | 0.009806 | 22.2355
[7 empty] | 1512.4 | 0.009806 | 22.2358
[7 empty] | 1512.4 | 0.009806 | 22.2357
[7 empty] | 1512.4 | 0.009806 | 22.2358
[7 empty] | 1512.4 | 0.009806 | 22.2358
[7 empty] | 1512.4 | 0.009806 | 22.236
[7 empty] | 1512.4 | 0.009806 | 22.2357
[7 empty] | 1512.4 | 0.009806 | 22.2359
[7 empty] | 1512.4 | 0.009806 | 22.2357
[7 empty] | 1512.4 | 0.009806 | 22.2359
[7 empty] | 1512.4 | 0.009806 | 22.2358
[7 empty] | Freq | 1512.4 | Rs | 22.2357 | 0.00013119 | Ls | 0.009806 | 8.2022e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1512;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | 1591.5 | 0.0098049 | 22.2454
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2458
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2457
[7 empty] | 1591.5 | 0.0098049 | 22.2456
[7 empty] | Freq | 1591.5 | Rs | 22.2456 | 8.6568e-05 | Ls | 0.0098049 | 1.0338e-08
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1592;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1670.8 | 0.0098039 | 22.255
[7 empty] | 1670.8 | 0.0098039 | 22.2549
[7 empty] | 1670.8 | 0.0098039 | 22.2553
[7 empty] | 1670.8 | 0.0098039 | 22.255
[7 empty] | 1670.8 | 0.0098039 | 22.255
[7 empty] | 1670.8 | 0.0098039 | 22.2549
[7 empty] | 1670.8 | 0.0098039 | 22.255
[7 empty] | 1670.8 | 0.0098039 | 22.2552
[7 empty] | 1670.8 | 0.0098039 | 22.255
[7 empty] | 1670.8 | 0.0098039 | 22.2551
[7 empty] | 1670.8 | 0.0098039 | 22.2551
[7 empty] | 1670.8 | 0.0098039 | 22.2551
[7 empty] | 1670.8 | 0.0098039 | 22.2553
[7 empty] | 1670.8 | 0.0098039 | 22.2551
[7 empty] | 1670.8 | 0.0098039 | 22.2553
[7 empty] | 1670.8 | 0.0098039 | 22.2552
[7 empty] | Freq | 1670.8 | Rs | 22.2551 | 0.00011834 | Ls | 0.0098039 | 7.5229e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1671;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1749.6 | 0.0098028 | 22.2644
[7 empty] | 1749.6 | 0.0098028 | 22.2646
[7 empty] | 1749.6 | 0.0098028 | 22.2644
[7 empty] | 1749.6 | 0.0098028 | 22.2643
[7 empty] | 1749.6 | 0.0098029 | 22.2649
[7 empty] | 1749.6 | 0.0098028 | 22.2657
[7 empty] | 1749.6 | 0.0098028 | 22.2643
[7 empty] | 1749.6 | 0.0098028 | 22.2644
[7 empty] | 1749.6 | 0.0098028 | 22.2646
[7 empty] | 1749.6 | 0.0098029 | 22.2646
[7 empty] | 1749.6 | 0.0098028 | 22.2644
[7 empty] | 1749.6 | 0.0098028 | 22.2646
[7 empty] | 1749.6 | 0.0098028 | 22.2648
[7 empty] | 1749.6 | 0.0098027 | 22.2646
[7 empty] | 1749.6 | 0.0098029 | 22.2637
[7 empty] | 1749.6 | 0.0098028 | 22.265
[7 empty] | Freq | 1749.6 | Rs | 22.2646 | 0.00041727 | Ls | 0.0098028 | 3.797e-08
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1750;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1750;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 100;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 2025-06-18 | 16:08
---------run finished---------
[7 empty] | lcr name: | KeySight E4980A
[7 empty] | Name and description of Device under test: | Sullivan B6 Coaxial GMH3I_2066
[7 empty] | GPIB source: | <bound method MessageBasedResource.read of <'GPIBInstrument'('GPIB0::17::INSTR')>>
[7 empty] | 2025-06-18 | 16:14
[7 empty] | select the freq range 90% - 110% | [1432.4, 1512.4, 1591.5, 1670.8, 1749.6]
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5419
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059356 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059356 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5419
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | 1432.4 | 0.00059357 | 3.5418
[7 empty] | Freq | 1432.4 | Rs | 3.5418 | 1.4197e-05 | Ls | 0.00059357 | 1.5776e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1432;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059353 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059353 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059353 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | 1512.4 | 0.00059353 | 3.5423
[7 empty] | 1512.4 | 0.00059354 | 3.5423
[7 empty] | Freq | 1512.4 | Rs | 3.5423 | 1.2811e-05 | Ls | 0.00059354 | 2.0471e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1512;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.0005935 | 3.5427
[7 empty] | 1591.5 | 0.00059351 | 3.5427
[7 empty] | Freq | 1591.5 | Rs | 3.5427 | 2.0027e-05 | Ls | 0.0005935 | 1.5738e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1592;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059348 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5432
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | 1670.8 | 0.00059347 | 3.5432
[7 empty] | 1670.8 | 0.00059347 | 3.5432
[7 empty] | 1670.8 | 0.00059347 | 3.5431
[7 empty] | Freq | 1670.8 | Rs | 3.5431 | 1.9316e-05 | Ls | 0.00059347 | 1.4722e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1671;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 1749.6 | 0.00059344 | 3.5435
[7 empty] | 1749.6 | 0.00059344 | 3.5435
[7 empty] | 1749.6 | 0.00059344 | 3.5435
[7 empty] | 1749.6 | 0.00059343 | 3.5437
[7 empty] | 1749.6 | 0.00059343 | 3.5436
[7 empty] | 1749.6 | 0.00059344 | 3.5435
[7 empty] | 1749.6 | 0.00059344 | 3.5436
[7 empty] | 1749.6 | 0.00059343 | 3.5435
[7 empty] | 1749.6 | 0.00059345 | 3.5436
[7 empty] | 1749.6 | 0.00059343 | 3.5434
[7 empty] | 1749.6 | 0.00059343 | 3.5436
[7 empty] | 1749.6 | 0.00059344 | 3.5436
[7 empty] | 1749.6 | 0.00059345 | 3.5435
[7 empty] | 1749.6 | 0.00059345 | 3.5436
[7 empty] | 1749.6 | 0.00059344 | 3.5435
[7 empty] | 1749.6 | 0.00059344 | 3.5436
[7 empty] | Freq | 1749.6 | Rs | 3.5435 | 7.1406e-05 | Ls | 0.00059344 | 6.6733e-09
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1750;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
:SOUR:DCS:STAT OFF;:BIAS:STAT OFF;:COMP:BIN:COUN:CLE;:LIST:CLE:ALL;:COMP:BIN:CLE;:AMPL:ALC OFF;:APER LONG,20;:BIAS:POL:AUTO OFF;:BIAS:RANG:AUTO ON;:BIAS:VOLT:LEV 0;:COMP:ABIN OFF;:COMP:BIN:COUN:STAT OFF;:COMP:MODE PTOL;:COMP:SLIM -9.9e+037,9.9e+037;:COMP:STAT OFF;:COMP:SWAP OFF;:COMP:TOL:NOM 0;:DISP:ENAB ON;:DISP:LINE "";:DISP:WIND:TEXT1:DATA:FMSD:DATA 1e-009;:DISP:WIND:TEXT2:DATA:FMSD:DATA 1e-009;:FORM:ASC:LONG ON;:FORM:BORD NORM;:FORM:DATA ASC,64;:FREQ:CW 1750;:FUNC:DEV1:MODE OFF;:FUNC:DEV1:REF:VAL 0;:FUNC:DEV2:MODE OFF;:FUNC:DEV2:REF:VAL 0;:FUNC:IMP:TYPE LSRS;:FUNC:SMON:IDC:STAT OFF;:FUNC:SMON:VDC:STAT OFF;:INIT:CONT OFF;:LIST:MODE SEQ;:LIST:STIM:TYPE FREQ,NONE;:SOUR:DCS:VOLT:LEV 0;:TRIG:DEL 0;:TRIG:SOUR INT;:TRIG:TDEL 0;:VOLT:LEV 0.1;:DISP:WIND:TEXT1:DATA:FMSD:STAT OFF;:DISP:WIND:TEXT2:DATA:FMSD:STAT OFF;:FUNC:DCR:RANG:VAL 100;:FUNC:IMP:RANG:VAL 10;:OUTP:DC:ISOL:LEV:VAL 0.02;:OUTP:DC:ISOL:STAT OFF;:FUNC:DCR:RANG:AUTO ON;:FUNC:IMP:RANG:AUTO ON;:OUTP:DC:ISOL:LEV:AUTO ON;:DISP:PAGE MEASurement;
[7 empty] | 2025-06-18 | 16:20
---------run finished---------

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
coax zero | 1591.5 | 3Z | -1 | -10 | 1592.54 | 20.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2559e-08 | 0.00012559 | 0 | 0 | coax zero1591.5 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
A1 | 1591.5 | 3Z | 99364 | 571520 | 1592.34 | 20.88 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00099386 | 6.3258e-08 | 1.9889 | 5.7358 | 0.033902 | 2.0825 | 1.9596e-08 | 0.00017904 | 3.2281 | 189.3508 | A11591.5 | 1 | 5.7358 | 0.033902 | 2.0825 | 1000000 | 993.865 | 0.063258 | 1.9889 | 1591.5
A2 | 1591.5 | 3Z | 197974 | 841874 | 1592.19 | 20.87 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0019802 | 9.8608e-08 | 2.0002 | 8.4486 | 0.04986 | 2.0836 | 3.1578e-08 | 0.00026245 | 3.1227 | 189.9775 | A21591.5 | 1 | 8.4486 | 0.04986 | 2.0836 | 1000000 | 1980.1779 | 0.098608 | 2.0002 | 1591.5
A3 | 1591.5 | 3Z | 296819 | 1070422 | 1593.41 | 20.65 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0029688 | 1.3834e-07 | 2.0101 | 10.7328 | 0.063326 | 2.0844 | 4.4717e-08 | 0.00035546 | 3.0937 | 178.1516 | A31591.5 | 1 | 10.7328 | 0.063326 | 2.0844 | 1000000 | 2968.8285 | 0.13834 | 2.0101 | 1591.5
A4 | 1591.5 | 3Z | 394747 | 1269967 | 1592.77 | 20.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0039483 | 1.791e-07 | 2.0156 | 12.7391 | 0.075136 | 2.0844 | 5.8089e-08 | 0.00045085 | 3.0831 | 166.6532 | A41591.5 | 1 | 12.7391 | 0.075136 | 2.0844 | 1000000 | 3948.3246 | 0.1791 | 2.0156 | 1591.5
A5 | 1591.5 | 3Z | 493101 | 1447708 | 1592.47 | 20.82 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0049321 | 2.2055e-07 | 2.0189 | 14.5255 | 0.085658 | 2.0843 | 7.1659e-08 | 0.00054782 | 3.0779 | 156.3624 | A51591.5 | 1 | 14.5255 | 0.085658 | 2.0843 | 1000000 | 4932.0809 | 0.22055 | 2.0189 | 1591.5
A6 | 1591.5 | 3Z | 591332 | 1617418 | 1591.8 | 20.66 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0059146 | 2.6221e-07 | 2.0211 | 16.2179 | 0.095651 | 2.0847 | 8.5298e-08 | 0.00064566 | 3.0741 | 148.1434 | A61591.5 | 1 | 16.2179 | 0.095651 | 2.0847 | 1000000 | 5914.5833 | 0.26221 | 2.0211 | 1591.5
A7 | 1591.5 | 3Z | 688992 | 1782538 | 1591.74 | 20.62 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0068914 | 3.0386e-07 | 2.0225 | 17.8708 | 0.1054 | 2.0848 | 9.8912e-08 | 0.00074362 | 3.072 | 141.74 | A71591.5 | 1 | 17.8708 | 0.1054 | 2.0848 | 1000000 | 6891.3848 | 0.30386 | 2.0225 | 1591.5
A8 | 1591.5 | 3Z | 786206 | 1934775 | 1591.67 | 20.68 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0078637 | 3.4542e-07 | 2.0236 | 19.4016 | 0.11442 | 2.0847 | 1.1248e-07 | 0.00084115 | 3.0709 | 136.0237 | A81591.5 | 1 | 19.4016 | 0.11442 | 2.0847 | 1000000 | 7863.7395 | 0.34542 | 2.0236 | 1591.5
A9 | 1591.5 | 3Z | 883126 | 2081205 | 1591.51 | 20.72 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0088332 | 3.8692e-07 | 2.0243 | 20.8733 | 0.12309 | 2.0846 | 1.2603e-07 | 0.00093857 | 3.07 | 131.1424 | A91591.5 | 1 | 20.8733 | 0.12309 | 2.0846 | 1000000 | 8833.1524 | 0.38692 | 2.0243 | 1591.5
A10 | 1591.5 | 3Z | 980165 | 2229664 | 1592.88 | 20.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0098038 | 4.2854e-07 | 2.0249 | 22.3632 | 0.13187 | 2.0846 | 1.3962e-07 | 0.0010365 | 3.0694 | 127.2261 | A101591.5 | 1 | 22.3632 | 0.13187 | 2.0846 | 1000000 | 9803.7512 | 0.42854 | 2.0249 | 1591.5
coax zero | 1591.5 | 2Z | -3 | -12 | 1593.08 | 20.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2554e-09 | 1.2554e-05 | 0 | 0 | coax zero1591.5 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
B1 | 1591.5 | 2Z | 99109 | 1168727 | 1593.85 | 20.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 9.9133e-05 | 7.2195e-09 | 1.9977 | 1.172 | 0.0069133 | 2.0845 | 2.2061e-09 | 2.288e-05 | 3.2726 | 302.1581 | B11591.5 | 1 | 1.172 | 0.0069133 | 2.0845 | 1000000 | 99.1331 | 0.0072195 | 1.9977 | 1591.5
B2 | 1591.5 | 2Z | 198937 | 1743376 | 1593.64 | 20.72 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00019898 | 1.1174e-08 | 2.0028 | 1.7485 | 0.010309 | 2.0847 | 3.5178e-09 | 3.39e-05 | 3.1765 | 304.1118 | B21591.5 | 1 | 1.7485 | 0.010309 | 2.0847 | 1000000 | 198.9824 | 0.011174 | 2.0028 | 1591.5
B3 | 1591.5 | 2Z | 298138 | 2213587 | 1593.08 | 20.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0002982 | 1.5343e-08 | 2.0065 | 2.2202 | 0.013089 | 2.0848 | 4.8893e-09 | 4.4803e-05 | 3.138 | 292.1394 | B31591.5 | 1 | 2.2202 | 0.013089 | 2.0848 | 1000000 | 298.2046 | 0.015343 | 2.0065 | 1591.5
B4 | 1591.5 | 2Z | 396394 | 2803515 | 1593.56 | 20.68 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00039648 | 1.9829e-08 | 2.0094 | 2.8113 | 0.016574 | 2.0849 | 6.3439e-09 | 5.7185e-05 | 3.1257 | 289.8262 | B41591.5 | 1 | 2.8113 | 0.016574 | 2.0849 | 1000000 | 396.4812 | 0.019829 | 2.0094 | 1591.5
B5 | 1591.5 | 2Z | 494345 | 3190209 | 1593.49 | 20.84 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00049445 | 2.4048e-08 | 2.0109 | 3.2011 | 0.018868 | 2.0846 | 7.7247e-09 | 6.7803e-05 | 3.1131 | 278.2729 | B51591.5 | 1 | 3.2011 | 0.018868 | 2.0846 | 1000000 | 494.4546 | 0.024048 | 2.0109 | 1591.5
B6 | 1591.5 | 2Z | 593381 | 3553499 | 1593.17 | 20.85 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00059351 | 2.8324e-08 | 2.0122 | 3.5658 | 0.021017 | 2.0846 | 9.1238e-09 | 7.8455e-05 | 3.1044 | 267.882 | B61591.5 | 1 | 3.5658 | 0.021017 | 2.0846 | 1000000 | 593.512 | 0.028324 | 2.0122 | 1591.5
B7 | 1591.5 | 2Z | 691176 | 3910839 | 1593.32 | 20.86 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00069133 | 3.2581e-08 | 2.0133 | 3.9245 | 0.023131 | 2.0846 | 1.0514e-08 | 8.9078e-05 | 3.0988 | 259.6702 | B71591.5 | 1 | 3.9245 | 0.023131 | 2.0846 | 1000000 | 691.3283 | 0.032581 | 2.0133 | 1591.5
B8 | 1591.5 | 2Z | 788907 | 4213640 | 1593.89 | 20.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00078908 | 3.6779e-08 | 2.0144 | 4.2255 | 0.024909 | 2.0849 | 1.189e-08 | 9.9318e-05 | 3.0933 | 250.8053 | B81591.5 | 1 | 4.2255 | 0.024909 | 2.0849 | 1000000 | 789.0777 | 0.036779 | 2.0144 | 1591.5
B9 | 1591.5 | 2Z | 887110 | 4532510 | 1593.65 | 20.75 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0008873 | 4.1047e-08 | 2.0152 | 4.5464 | 0.026799 | 2.0848 | 1.3283e-08 | 0.00010981 | 3.0903 | 244.0554 | B91591.5 | 1 | 4.5464 | 0.026799 | 2.0848 | 1000000 | 887.3026 | 0.041047 | 2.0152 | 1591.5
B10 | 1591.5 | 2Z | 984585 | 4826330 | 1595.54 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00098478 | 4.5247e-08 | 2.0161 | 4.8226 | 0.028508 | 2.0855 | 1.4661e-08 | 0.00012009 | 3.0862 | 237.396 | B101591.5 | 1 | 4.8226 | 0.028508 | 2.0855 | 1000000 | 984.7795 | 0.045247 | 2.0161 | 1591.5
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35

## pyReport
| Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
A1 | 1 | 993.865 | 0.063258 | 1.9889 | 5.7358 | 0.033902 | 2.0825
A2 | 2 | 1980.1779 | 0.098608 | 2.0002 | 8.4486 | 0.04986 | 2.0836
A3 | 3 | 2968.8285 | 0.13834 | 2.0101 | 10.7328 | 0.063326 | 2.0844
A4 | 4 | 3948.3246 | 0.1791 | 2.0156 | 12.7391 | 0.075136 | 2.0844
A5 | 5 | 4932.0809 | 0.22055 | 2.0189 | 14.5255 | 0.085658 | 2.0843
A6 | 6 | 5914.5833 | 0.26221 | 2.0211 | 16.2179 | 0.095651 | 2.0847
A7 | 7 | 6891.3848 | 0.30386 | 2.0225 | 17.8708 | 0.1054 | 2.0848
A8 | 8 | 7863.7395 | 0.34542 | 2.0236 | 19.4016 | 0.11442 | 2.0847
A9 | 9 | 8833.1524 | 0.38692 | 2.0243 | 20.8733 | 0.12309 | 2.0846
A10 | 10 | 9803.7512 | 0.42854 | 2.0249 | 22.3632 | 0.13187 | 2.0846
| Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
B1 | 1 | 99.1331 | 0.0072195 | 1.9977 | 1.172 | 0.0069133 | 2.0845
B2 | 2 | 198.9824 | 0.011174 | 2.0028 | 1.7485 | 0.010309 | 2.0847
B3 | 3 | 298.2046 | 0.015343 | 2.0065 | 2.2202 | 0.013089 | 2.0848
B4 | 4 | 396.4812 | 0.019829 | 2.0094 | 2.8113 | 0.016574 | 2.0849
B5 | 5 | 494.4546 | 0.024048 | 2.0109 | 3.2011 | 0.018868 | 2.0846
B6 | 6 | 593.512 | 0.028324 | 2.0122 | 3.5658 | 0.021017 | 2.0846
B7 | 7 | 691.3283 | 0.032581 | 2.0133 | 3.9245 | 0.023131 | 2.0846
B8 | 8 | 789.0777 | 0.036779 | 2.0144 | 4.2255 | 0.024909 | 2.0849
B9 | 9 | 887.3026 | 0.041047 | 2.0152 | 4.5464 | 0.026799 | 2.0848
B10 | 10 | 984.7795 | 0.045247 | 2.0161 | 4.8226 | 0.028508 | 2.0855
B10 | 10 | 984.7795 | 0.045247 | 2.0161 | 4.8226 | 0.028508 | 2.0855

## environment
FCU10 Impedance
Date UTC | Temperature VSL | Humidity VSL | NZST_Time
2025-05-14 20:00:09 | 20.42 | 48.65 | 45792.3334 | 2025-05-15 08:00
2025-05-14 20:02:08 | 20.41 | 48.63 | 45792.3348 | 2025-05-15 08:02
2025-05-14 20:04:09 | 20.41 | 48.74 | 45792.3362 | 2025-05-15 08:04
2025-05-14 20:06:09 | 20.41 | 48.68 | 45792.3376 | 2025-05-15 08:06
2025-05-14 20:08:09 | 20.4 | 48.83 | 45792.339 | 2025-05-15 08:08
2025-05-14 20:10:08 | 20.37 | 48.81 | 45792.3404 | 2025-05-15 08:10
2025-05-14 20:12:09 | 20.39 | 48.88 | 45792.3418 | 2025-05-15 08:12
2025-05-14 20:14:08 | 20.38 | 48.95 | 45792.3431 | 2025-05-15 08:14
2025-05-14 20:16:08 | 20.39 | 48.95 | 45792.3445 | 2025-05-15 08:16
2025-05-14 20:18:08 | 20.37 | 49.02 | 45792.3459 | 2025-05-15 08:18
2025-05-14 20:20:09 | 20.4 | 48.95 | 45792.3473 | 2025-05-15 08:20
2025-05-14 20:22:08 | 20.4 | 49.01 | 45792.3487 | 2025-05-15 08:22
2025-05-14 20:24:09 | 20.41 | 48.99 | 45792.3501 | 2025-05-15 08:24
2025-05-14 20:26:09 | 20.42 | 49.06 | 45792.3515 | 2025-05-15 08:26
2025-05-14 20:28:09 | 20.43 | 48.98 | 45792.3529 | 2025-05-15 08:28
2025-05-14 20:30:08 | 20.4 | 49.09 | 45792.3543 | 2025-05-15 08:30
2025-05-14 20:32:09 | 20.44 | 48.92 | 45792.3557 | 2025-05-15 08:32
2025-05-14 20:34:08 | 20.43 | 48.9 | 45792.357 | 2025-05-15 08:34
2025-05-14 20:36:08 | 20.42 | 48.95 | 45792.3584 | 2025-05-15 08:36
2025-05-14 20:38:09 | 20.42 | 48.93 | 45792.3598 | 2025-05-15 08:38
2025-05-14 20:40:09 | 20.4 | 49.09 | 45792.3612 | 2025-05-15 08:40
2025-05-14 20:42:08 | 20.4 | 49.08 | 45792.3626 | 2025-05-15 08:42
2025-05-14 20:44:09 | 20.41 | 49.15 | 45792.364 | 2025-05-15 08:44
2025-05-14 20:46:08 | 20.42 | 49.22 | 45792.3654 | 2025-05-15 08:46
2025-05-14 20:48:09 | 20.41 | 49.23 | 45792.3668 | 2025-05-15 08:48
2025-05-14 20:50:08 | 20.4 | 49.32 | 45792.3681 | 2025-05-15 08:50
2025-05-14 20:52:09 | 20.4 | 49.18 | 45792.3695 | 2025-05-15 08:52
2025-05-14 20:54:08 | 20.4 | 49.2 | 45792.3709 | 2025-05-15 08:54
2025-05-14 20:56:09 | 20.38 | 49.26 | 45792.3723 | 2025-05-15 08:56
2025-05-14 20:58:08 | 20.39 | 49.31 | 45792.3737 | 2025-05-15 08:58
2025-05-14 21:00:09 | 20.38 | 49.4 | 45792.3751 | 2025-05-15 09:00
2025-05-14 21:02:09 | 20.39 | 49.23 | 45792.3765 | 2025-05-15 09:02
2025-05-14 21:04:09 | 20.4 | 49.24 | 45792.3779 | 2025-05-15 09:04
2025-05-14 21:06:08 | 20.39 | 49.33 | 45792.3793 | 2025-05-15 09:06
2025-05-14 21:08:09 | 20.38 | 49.37 | 45792.3807 | 2025-05-15 09:08
2025-05-14 21:10:09 | 20.41 | 49.43 | 45792.382 | 2025-05-15 09:10
2025-05-14 21:12:08 | 20.42 | 49.37 | 45792.3834 | 2025-05-15 09:12
2025-05-14 21:14:08 | 20.43 | 49.43 | 45792.3848 | 2025-05-15 09:14
2025-05-14 21:16:09 | 20.44 | 49.47 | 45792.3862 | 2025-05-15 09:16
2025-05-14 21:18:09 | 20.44 | 49.29 | 45792.3876 | 2025-05-15 09:18
2025-05-14 21:20:09 | 20.43 | 49.39 | 45792.389 | 2025-05-15 09:20
2025-05-14 21:22:08 | 20.43 | 49.38 | 45792.3904 | 2025-05-15 09:22
2025-05-14 21:24:09 | 20.44 | 49.52 | 45792.3918 | 2025-05-15 09:24
2025-05-14 21:26:08 | 20.43 | 49.58 | 45792.3931 | 2025-05-15 09:26
2025-05-14 21:28:08 | 20.44 | 49.45 | 45792.3945 | 2025-05-15 09:28
2025-05-14 21:30:08 | 20.44 | 49.5 | 45792.3959 | 2025-05-15 09:30
2025-05-14 21:32:09 | 20.44 | 49.61 | 45792.3973 | 2025-05-15 09:32
2025-05-14 21:34:08 | 20.43 | 49.56 | 45792.3987 | 2025-05-15 09:34
2025-05-14 21:36:08 | 20.43 | 49.66 | 45792.4001 | 2025-05-15 09:36
2025-05-14 21:38:09 | 20.43 | 49.62 | 45792.4015 | 2025-05-15 09:38
2025-05-14 21:40:09 | 20.42 | 49.55 | 45792.4029 | 2025-05-15 09:40
2025-05-14 21:42:08 | 20.42 | 49.16 | 45792.4043 | 2025-05-15 09:42
2025-05-14 21:44:09 | 20.42 | 48.89 | 45792.4057 | 2025-05-15 09:44
2025-05-14 21:46:08 | 20.43 | 48.8 | 45792.407 | 2025-05-15 09:46
2025-05-14 21:48:08 | 20.44 | 49.12 | 45792.4084 | 2025-05-15 09:48
2025-05-14 21:50:08 | 20.45 | 49.3 | 45792.4098 | 2025-05-15 09:50
2025-05-14 21:52:09 | 20.46 | 49.36 | 45792.4112 | 2025-05-15 09:52
2025-05-14 21:54:08 | 20.47 | 49.45 | 45792.4126 | 2025-05-15 09:54
2025-05-14 21:56:08 | 20.47 | 49.38 | 45792.414 | 2025-05-15 09:56
2025-05-14 21:58:08 | 20.46 | 49.51 | 45792.4154 | 2025-05-15 09:58
2025-05-14 22:00:09 | 20.47 | 49.39 | 45792.4168 | 2025-05-15 10:00
2025-05-14 22:02:08 | 20.46 | 49.32 | 45792.4181 | 2025-05-15 10:02
2025-05-14 22:04:08 | 20.46 | 49.35 | 45792.4195 | 2025-05-15 10:04
2025-05-14 22:06:08 | 20.46 | 49.32 | 45792.4209 | 2025-05-15 10:06
2025-05-14 22:08:09 | 20.47 | 49.35 | 45792.4223 | 2025-05-15 10:08
2025-05-14 22:10:08 | 20.48 | 49.41 | 45792.4237 | 2025-05-15 10:10
2025-05-14 22:12:08 | 20.47 | 49.31 | 45792.4251 | 2025-05-15 10:12
2025-05-14 22:14:08 | 20.48 | 49.4 | 45792.4265 | 2025-05-15 10:14
2025-05-14 22:16:09 | 20.46 | 49.34 | 45792.4279 | 2025-05-15 10:16
2025-05-14 22:18:08 | 20.47 | 49.35 | 45792.4293 | 2025-05-15 10:18
2025-05-14 22:20:09 | 20.46 | 49.45 | 45792.4307 | 2025-05-15 10:20
2025-05-14 22:22:08 | 20.46 | 49.35 | 45792.432 | 2025-05-15 10:22
2025-05-14 22:24:09 | 20.47 | 49.32 | 45792.4334 | 2025-05-15 10:24
2025-05-14 22:26:09 | 20.48 | 49.42 | 45792.4348 | 2025-05-15 10:26
2025-05-14 22:28:09 | 20.46 | 49.41 | 45792.4362 | 2025-05-15 10:28
2025-05-14 22:30:08 | 20.45 | 49.51 | 45792.4376 | 2025-05-15 10:30
2025-05-14 22:32:09 | 20.44 | 49.19 | 45792.439 | 2025-05-15 10:32
2025-05-14 22:34:08 | 20.44 | 48.86 | 45792.4404 | 2025-05-15 10:34
2025-05-14 22:36:08 | 20.43 | 48.63 | 45792.4418 | 2025-05-15 10:36
2025-05-14 22:38:08 | 20.43 | 48.93 | 45792.4431 | 2025-05-15 10:38
2025-05-14 22:40:09 | 20.44 | 49.04 | 45792.4445 | 2025-05-15 10:40
2025-05-14 22:42:09 | 20.43 | 49 | 45792.4459 | 2025-05-15 10:42
2025-05-14 22:44:09 | 20.43 | 49.02 | 45792.4473 | 2025-05-15 10:44
2025-05-14 22:46:08 | 20.43 | 49.17 | 45792.4487 | 2025-05-15 10:46
2025-05-14 22:48:09 | 20.43 | 49.24 | 45792.4501 | 2025-05-15 10:48
2025-05-14 22:50:08 | 20.42 | 49.39 | 45792.4515 | 2025-05-15 10:50
2025-05-14 22:52:08 | 20.42 | 49.44 | 45792.4529 | 2025-05-15 10:52
2025-05-14 22:54:08 | 20.43 | 49.43 | 45792.4543 | 2025-05-15 10:54
2025-05-14 22:56:09 | 20.42 | 49.43 | 45792.4557 | 2025-05-15 10:56
2025-05-14 22:58:08 | 20.43 | 49.38 | 45792.457 | 2025-05-15 10:58
2025-05-14 23:00:08 | 20.42 | 49.35 | 45792.4584 | 2025-05-15 11:00
2025-05-14 23:02:08 | 20.41 | 49.33 | 45792.4598 | 2025-05-15 11:02
2025-05-14 23:04:09 | 20.42 | 49.18 | 45792.4612 | 2025-05-15 11:04
2025-05-14 23:06:08 | 20.44 | 49.19 | 45792.4626 | 2025-05-15 11:06
2025-05-14 23:08:08 | 20.43 | 49.29 | 45792.464 | 2025-05-15 11:08
2025-05-14 23:10:08 | 20.44 | 49.14 | 45792.4654 | 2025-05-15 11:10
2025-05-14 23:12:09 | 20.43 | 49.15 | 45792.4668 | 2025-05-15 11:12
2025-05-14 23:14:09 | 20.44 | 49.22 | 45792.4682 | 2025-05-15 11:14
2025-05-14 23:16:09 | 20.44 | 49.18 | 45792.4695 | 2025-05-15 11:16
2025-05-14 23:18:08 | 20.43 | 49.33 | 45792.4709 | 2025-05-15 11:18
2025-05-14 23:20:09 | 20.42 | 49.41 | 45792.4723 | 2025-05-15 11:20
2025-05-14 23:22:08 | 20.44 | 49.35 | 45792.4737 | 2025-05-15 11:22
2025-05-14 23:24:08 | 20.45 | 49.4 | 45792.4751 | 2025-05-15 11:24
2025-05-14 23:26:08 | 20.41 | 49.44 | 45792.4765 | 2025-05-15 11:26
2025-05-14 23:28:09 | 20.42 | 49.22 | 45792.4779 | 2025-05-15 11:28
2025-05-14 23:30:08 | 20.42 | 49.22 | 45792.4793 | 2025-05-15 11:30
2025-05-14 23:32:08 | 20.43 | 49.28 | 45792.4806 | 2025-05-15 11:32
2025-05-14 23:34:08 | 20.42 | 49.23 | 45792.482 | 2025-05-15 11:34
2025-05-14 23:36:09 | 20.44 | 49.21 | 45792.4834 | 2025-05-15 11:36
2025-05-14 23:38:09 | 20.43 | 49.12 | 45792.4848 | 2025-05-15 11:38
2025-05-14 23:40:09 | 20.43 | 48.98 | 45792.4862 | 2025-05-15 11:40
2025-05-14 23:42:08 | 20.43 | 48.98 | 45792.4876 | 2025-05-15 11:42
2025-05-14 23:44:09 | 20.41 | 48.95 | 45792.489 | 2025-05-15 11:44
2025-05-14 23:46:08 | 20.44 | 48.84 | 45792.4904 | 2025-05-15 11:46
2025-05-14 23:48:08 | 20.43 | 48.83 | 45792.4918 | 2025-05-15 11:48
2025-05-14 23:50:08 | 20.42 | 48.89 | 45792.4931 | 2025-05-15 11:50
2025-05-14 23:52:09 | 20.44 | 48.72 | 45792.4945 | 2025-05-15 11:52
2025-05-14 23:54:08 | 20.44 | 48.75 | 45792.4959 | 2025-05-15 11:54
2025-05-14 23:56:09 | 20.43 | 48.79 | 45792.4973 | 2025-05-15 11:56
2025-05-14 23:58:08 | 20.44 | 48.83 | 45792.4987 | 2025-05-15 11:58
2025-05-15 00:00:09 | 20.42 | 48.71 | 45792.5001 | 2025-05-15 12:00
2025-05-15 00:02:08 | 20.42 | 48.78 | 45792.5015 | 2025-05-15 12:02
2025-05-15 00:04:08 | 20.43 | 48.76 | 45792.5029 | 2025-05-15 12:04
2025-05-15 00:06:08 | 20.42 | 48.52 | 45792.5043 | 2025-05-15 12:06
2025-05-15 00:08:09 | 20.44 | 48.41 | 45792.5057 | 2025-05-15 12:08
2025-05-15 00:10:08 | 20.44 | 48.36 | 45792.507 | 2025-05-15 12:10
2025-05-15 00:12:09 | 20.44 | 48.37 | 45792.5084 | 2025-05-15 12:12
2025-05-15 00:14:08 | 20.44 | 48.4 | 45792.5098 | 2025-05-15 12:14
2025-05-15 00:16:09 | 20.43 | 48.29 | 45792.5112 | 2025-05-15 12:16
2025-05-15 00:18:08 | 20.45 | 48.28 | 45792.5126 | 2025-05-15 12:18
2025-05-15 00:20:09 | 20.44 | 48.33 | 45792.514 | 2025-05-15 12:20
2025-05-15 00:22:08 | 20.45 | 48.53 | 45792.5154 | 2025-05-15 12:22
2025-05-15 00:24:09 | 20.44 | 48.6 | 45792.5168 | 2025-05-15 12:24
2025-05-15 00:26:09 | 20.43 | 48.53 | 45792.5182 | 2025-05-15 12:26
2025-05-15 00:28:09 | 20.41 | 48.64 | 45792.5195 | 2025-05-15 12:28
2025-05-15 00:30:08 | 20.43 | 48.7 | 45792.5209 | 2025-05-15 12:30
2025-05-15 00:32:09 | 20.42 | 48.64 | 45792.5223 | 2025-05-15 12:32
2025-05-15 00:34:09 | 20.41 | 48.76 | 45792.5237 | 2025-05-15 12:34
2025-05-15 00:36:09 | 20.42 | 48.83 | 45792.5251 | 2025-05-15 12:36
2025-05-15 00:38:08 | 20.42 | 48.76 | 45792.5265 | 2025-05-15 12:38
2025-05-15 00:40:09 | 20.43 | 48.84 | 45792.5279 | 2025-05-15 12:40
2025-05-15 00:42:08 | 20.43 | 48.87 | 45792.5293 | 2025-05-15 12:42
2025-05-15 00:44:09 | 20.43 | 48.81 | 45792.5307 | 2025-05-15 12:44
2025-05-15 00:46:08 | 20.42 | 48.84 | 45792.532 | 2025-05-15 12:46
2025-05-15 00:48:09 | 20.43 | 48.85 | 45792.5334 | 2025-05-15 12:48
2025-05-15 00:50:09 | 20.44 | 48.81 | 45792.5348 | 2025-05-15 12:50
2025-05-15 00:52:09 | 20.42 | 48.86 | 45792.5362 | 2025-05-15 12:52
2025-05-15 00:54:08 | 20.43 | 48.69 | 45792.5376 | 2025-05-15 12:54
2025-05-15 00:56:09 | 20.44 | 48.61 | 45792.539 | 2025-05-15 12:56
2025-05-15 00:58:09 | 20.44 | 48.69 | 45792.5404 | 2025-05-15 12:58
2025-05-15 01:00:09 | 20.44 | 48.63 | 45792.5418 | 2025-05-15 13:00
2025-05-15 01:02:08 | 20.42 | 48.76 | 45792.5431 | 2025-05-15 13:02
2025-05-15 01:04:09 | 20.41 | 48.86 | 45792.5445 | 2025-05-15 13:04
2025-05-15 01:06:09 | 20.42 | 48.69 | 45792.5459 | 2025-05-15 13:06
2025-05-15 01:08:09 | 20.43 | 48.66 | 45792.5473 | 2025-05-15 13:08
2025-05-15 01:10:08 | 20.42 | 48.74 | 45792.5487 | 2025-05-15 13:10
2025-05-15 01:12:09 | 20.43 | 48.79 | 45792.5501 | 2025-05-15 13:12
2025-05-15 01:14:09 | 20.42 | 48.81 | 45792.5515 | 2025-05-15 13:14
2025-05-15 01:16:09 | 20.44 | 48.92 | 45792.5529 | 2025-05-15 13:16
2025-05-15 01:18:08 | 20.44 | 48.95 | 45792.5543 | 2025-05-15 13:18
2025-05-15 01:20:09 | 20.45 | 48.86 | 45792.5557 | 2025-05-15 13:20
2025-05-15 01:22:09 | 20.44 | 48.92 | 45792.557 | 2025-05-15 13:22
2025-05-15 01:24:09 | 20.45 | 48.88 | 45792.5584 | 2025-05-15 13:24
2025-05-15 01:26:08 | 20.46 | 48.74 | 45792.5598 | 2025-05-15 13:26
2025-05-15 01:28:08 | 20.44 | 48.77 | 45792.5612 | 2025-05-15 13:28
2025-05-15 01:30:08 | 20.43 | 48.86 | 45792.5626 | 2025-05-15 13:30
2025-05-15 01:32:09 | 20.43 | 48.84 | 45792.564 | 2025-05-15 13:32
2025-05-15 01:34:08 | 20.42 | 48.89 | 45792.5654 | 2025-05-15 13:34
2025-05-15 01:36:08 | 20.42 | 48.99 | 45792.5668 | 2025-05-15 13:36
2025-05-15 01:38:08 | 20.41 | 49.02 | 45792.5681 | 2025-05-15 13:38
2025-05-15 01:40:09 | 20.44 | 48.94 | 45792.5695 | 2025-05-15 13:40
2025-05-15 01:42:08 | 20.44 | 49.01 | 45792.5709 | 2025-05-15 13:42
2025-05-15 01:44:08 | 20.44 | 48.99 | 45792.5723 | 2025-05-15 13:44
2025-05-15 01:46:08 | 20.43 | 49 | 45792.5737 | 2025-05-15 13:46
2025-05-15 01:48:09 | 20.41 | 49.09 | 45792.5751 | 2025-05-15 13:48
2025-05-15 01:50:08 | 20.44 | 49.04 | 45792.5765 | 2025-05-15 13:50
2025-05-15 01:52:09 | 20.42 | 49.13 | 45792.5779 | 2025-05-15 13:52
2025-05-15 01:54:08 | 20.41 | 49.24 | 45792.5793 | 2025-05-15 13:54
2025-05-15 01:56:09 | 20.43 | 49.18 | 45792.5807 | 2025-05-15 13:56
2025-05-15 01:58:08 | 20.44 | 49.3 | 45792.582 | 2025-05-15 13:58
2025-05-15 02:00:08 | 20.43 | 49.34 | 45792.5834 | 2025-05-15 14:00
2025-05-15 02:02:08 | 20.45 | 49.23 | 45792.5848 | 2025-05-15 14:02
2025-05-15 02:04:09 | 20.45 | 49.26 | 45792.5862 | 2025-05-15 14:04
2025-05-15 02:06:09 | 20.44 | 49.2 | 45792.5876 | 2025-05-15 14:06
2025-05-15 02:08:09 | 20.45 | 49.23 | 45792.589 | 2025-05-15 14:08
2025-05-15 02:10:08 | 20.45 | 49.23 | 45792.5904 | 2025-05-15 14:10
2025-05-15 02:12:09 | 20.44 | 49.04 | 45792.5918 | 2025-05-15 14:12
2025-05-15 02:14:09 | 20.43 | 49.03 | 45792.5932 | 2025-05-15 14:14
2025-05-15 02:16:09 | 20.42 | 49.03 | 45792.5945 | 2025-05-15 14:16
2025-05-15 02:18:08 | 20.44 | 48.96 | 45792.5959 | 2025-05-15 14:18
2025-05-15 02:20:09 | 20.45 | 49.03 | 45792.5973 | 2025-05-15 14:20
2025-05-15 02:22:08 | 20.43 | 48.87 | 45792.5987 | 2025-05-15 14:22
2025-05-15 02:24:09 | 20.43 | 48.86 | 45792.6001 | 2025-05-15 14:24
2025-05-15 02:26:08 | 20.44 | 48.84 | 45792.6015 | 2025-05-15 14:26
2025-05-15 02:28:09 | 20.42 | 48.97 | 45792.6029 | 2025-05-15 14:28
2025-05-15 02:30:08 | 20.42 | 49.1 | 45792.6043 | 2025-05-15 14:30
2025-05-15 02:32:09 | 20.42 | 49.04 | 45792.6057 | 2025-05-15 14:32
2025-05-15 02:34:08 | 20.43 | 49.12 | 45792.607 | 2025-05-15 14:34
2025-05-15 02:36:09 | 20.41 | 49.15 | 45792.6084 | 2025-05-15 14:36
2025-05-15 02:38:09 | 20.42 | 49.07 | 45792.6098 | 2025-05-15 14:38
2025-05-15 02:40:09 | 20.44 | 49.06 | 45792.6112 | 2025-05-15 14:40
2025-05-15 02:42:08 | 20.45 | 49.09 | 45792.6126 | 2025-05-15 14:42
2025-05-15 02:44:09 | 20.43 | 49.02 | 45792.614 | 2025-05-15 14:44
2025-05-15 02:46:09 | 20.43 | 49.04 | 45792.6154 | 2025-05-15 14:46
2025-05-15 02:48:09 | 20.43 | 49.1 | 45792.6168 | 2025-05-15 14:48
2025-05-15 02:50:08 | 20.44 | 48.96 | 45792.6181 | 2025-05-15 14:50
2025-05-15 02:52:08 | 20.44 | 49.05 | 45792.6195 | 2025-05-15 14:52
2025-05-15 02:54:08 | 20.45 | 49.08 | 45792.6209 | 2025-05-15 14:54
2025-05-15 02:56:09 | 20.44 | 49.04 | 45792.6223 | 2025-05-15 14:56
2025-05-15 02:58:09 | 20.43 | 49.1 | 45792.6237 | 2025-05-15 14:58
2025-05-15 03:00:09 | 20.45 | 49.04 | 45792.6251 | 2025-05-15 15:00
2025-05-15 03:02:08 | 20.44 | 49.09 | 45792.6265 | 2025-05-15 15:02
2025-05-15 03:04:09 | 20.43 | 49.18 | 45792.6279 | 2025-05-15 15:04
2025-05-15 03:06:09 | 20.42 | 49.14 | 45792.6293 | 2025-05-15 15:06
2025-05-15 03:08:09 | 20.44 | 49.22 | 45792.6307 | 2025-05-15 15:08
2025-05-15 03:10:08 | 20.42 | 49.15 | 45792.632 | 2025-05-15 15:10
2025-05-15 03:12:09 | 20.42 | 49.23 | 45792.6334 | 2025-05-15 15:12
2025-05-15 03:14:08 | 20.44 | 49.27 | 45792.6348 | 2025-05-15 15:14
2025-05-15 03:16:09 | 20.44 | 49.2 | 45792.6362 | 2025-05-15 15:16
2025-05-15 03:18:08 | 20.44 | 49.27 | 45792.6376 | 2025-05-15 15:18
2025-05-15 03:20:08 | 20.44 | 49.2 | 45792.639 | 2025-05-15 15:20
2025-05-15 03:22:08 | 20.44 | 49.27 | 45792.6404 | 2025-05-15 15:22
2025-05-15 03:24:09 | 20.45 | 49.34 | 45792.6418 | 2025-05-15 15:24
2025-05-15 03:26:09 | 20.44 | 49.29 | 45792.6432 | 2025-05-15 15:26
2025-05-15 03:28:09 | 20.44 | 49.38 | 45792.6445 | 2025-05-15 15:28
2025-05-15 03:30:08 | 20.43 | 49.24 | 45792.6459 | 2025-05-15 15:30
2025-05-15 03:32:09 | 20.41 | 49.23 | 45792.6473 | 2025-05-15 15:32
2025-05-15 03:34:08 | 20.42 | 49.28 | 45792.6487 | 2025-05-15 15:34
2025-05-15 03:36:09 | 20.42 | 49.2 | 45792.6501 | 2025-05-15 15:36
2025-05-15 03:38:09 | 20.42 | 49.28 | 45792.6515 | 2025-05-15 15:38
2025-05-15 03:40:09 | 20.42 | 49.27 | 45792.6529 | 2025-05-15 15:40
2025-05-15 03:42:08 | 20.42 | 48.81 | 45792.6543 | 2025-05-15 15:42
2025-05-15 03:44:09 | 20.42 | 48.46 | 45792.6557 | 2025-05-15 15:44
2025-05-15 03:46:08 | 20.41 | 48.42 | 45792.657 | 2025-05-15 15:46
2025-05-15 03:48:09 | 20.42 | 48.63 | 45792.6584 | 2025-05-15 15:48
2025-05-15 03:50:08 | 20.42 | 48.71 | 45792.6598 | 2025-05-15 15:50
2025-05-15 03:52:08 | 20.42 | 48.82 | 45792.6612 | 2025-05-15 15:52
2025-05-15 03:54:08 | 20.42 | 48.92 | 45792.6626 | 2025-05-15 15:54
2025-05-15 03:56:09 | 20.41 | 48.91 | 45792.664 | 2025-05-15 15:56
2025-05-15 03:58:08 | 20.4 | 49.02 | 45792.6654 | 2025-05-15 15:58
2025-05-15 04:00:09 | 20.42 | 48.98 | 45792.6668 | 2025-05-15 16:00
2025-05-15 04:02:08 | 20.43 | 49.04 | 45792.6681 | 2025-05-15 16:02
2025-05-15 04:04:09 | 20.43 | 49.03 | 45792.6695 | 2025-05-15 16:04
2025-05-15 04:06:08 | 20.44 | 49.06 | 45792.6709 | 2025-05-15 16:06
2025-05-15 04:08:09 | 20.42 | 49.17 | 45792.6723 | 2025-05-15 16:08
2025-05-15 04:10:08 | 20.43 | 49.13 | 45792.6737 | 2025-05-15 16:10
2025-05-15 04:12:09 | 20.41 | 49.22 | 45792.6751 | 2025-05-15 16:12
2025-05-15 04:14:08 | 20.41 | 49.11 | 45792.6765 | 2025-05-15 16:14
2025-05-15 04:16:09 | 20.4 | 49.1 | 45792.6779 | 2025-05-15 16:16
2025-05-15 04:18:08 | 20.42 | 49.17 | 45792.6793 | 2025-05-15 16:18
2025-05-15 04:20:09 | 20.44 | 49.1 | 45792.6807 | 2025-05-15 16:20
2025-05-15 04:22:08 | 20.45 | 49.18 | 45792.682 | 2025-05-15 16:22
2025-05-15 04:24:09 | 20.45 | 49.22 | 45792.6834 | 2025-05-15 16:24
2025-05-15 04:26:08 | 20.44 | 49.22 | 45792.6848 | 2025-05-15 16:26
2025-05-15 04:28:08 | 20.43 | 49.32 | 45792.6862 | 2025-05-15 16:28
2025-05-15 04:30:08 | 20.42 | 49.25 | 45792.6876 | 2025-05-15 16:30
2025-05-15 04:32:09 | 20.43 | 49.31 | 45792.689 | 2025-05-15 16:32
2025-05-15 04:34:08 | 20.43 | 49.36 | 45792.6904 | 2025-05-15 16:34
2025-05-15 04:36:08 | 20.43 | 49.34 | 45792.6918 | 2025-05-15 16:36
2025-05-15 04:38:09 | 20.42 | 49.45 | 45792.6932 | 2025-05-15 16:38
2025-05-15 04:40:09 | 20.43 | 49.06 | 45792.6945 | 2025-05-15 16:40
2025-05-15 04:42:08 | 20.44 | 48.73 | 45792.6959 | 2025-05-15 16:42
2025-05-15 04:44:09 | 20.43 | 48.49 | 45792.6973 | 2025-05-15 16:44
2025-05-15 04:46:09 | 20.43 | 48.67 | 45792.6987 | 2025-05-15 16:46
2025-05-15 04:48:09 | 20.42 | 48.82 | 45792.7001 | 2025-05-15 16:48
2025-05-15 04:50:08 | 20.42 | 48.87 | 45792.7015 | 2025-05-15 16:50
2025-05-15 04:52:09 | 20.43 | 48.95 | 45792.7029 | 2025-05-15 16:52
2025-05-15 04:54:09 | 20.43 | 48.92 | 45792.7043 | 2025-05-15 16:54
2025-05-15 04:56:09 | 20.44 | 49.01 | 45792.7057 | 2025-05-15 16:56
2025-05-15 04:58:08 | 20.42 | 48.98 | 45792.707 | 2025-05-15 16:58
2025-05-15 05:00:08 | 20.42 | 49.07 | 45792.7084 | 2025-05-15 17:00
2025-05-15 05:02:08 | 20.44 | 48.99 | 45792.7098 | 2025-05-15 17:02
2025-05-15 05:04:09 | 20.42 | 49.11 | 45792.7112 | 2025-05-15 17:04
2025-05-15 05:06:08 | 20.43 | 49.09 | 45792.7126 | 2025-05-15 17:06
2025-05-15 05:08:09 | 20.44 | 49.18 | 45792.714 | 2025-05-15 17:08
2025-05-15 05:10:09 | 20.43 | 49.2 | 45792.7154 | 2025-05-15 17:10
2025-05-15 05:12:09 | 20.43 | 49.17 | 45792.7168 | 2025-05-15 17:12
2025-05-15 05:14:08 | 20.42 | 49.27 | 45792.7181 | 2025-05-15 17:14
2025-05-15 05:16:08 | 20.42 | 49.22 | 45792.7195 | 2025-05-15 17:16
2025-05-15 05:18:09 | 20.43 | 49.28 | 45792.7209 | 2025-05-15 17:18
2025-05-15 05:20:09 | 20.42 | 49.19 | 45792.7223 | 2025-05-15 17:20
2025-05-15 05:22:08 | 20.43 | 49.13 | 45792.7237 | 2025-05-15 17:22
2025-05-15 05:24:09 | 20.43 | 49.18 | 45792.7251 | 2025-05-15 17:24
2025-05-15 05:26:09 | 20.42 | 49.13 | 45792.7265 | 2025-05-15 17:26
2025-05-15 05:28:09 | 20.42 | 49.23 | 45792.7279 | 2025-05-15 17:28
2025-05-15 05:30:08 | 20.41 | 49.21 | 45792.7293 | 2025-05-15 17:30
2025-05-15 05:32:08 | 20.41 | 49.26 | 45792.7306 | 2025-05-15 17:32
2025-05-15 05:34:08 | 20.42 | 48.81 | 45792.732 | 2025-05-15 17:34
2025-05-15 05:36:09 | 20.42 | 48.49 | 45792.7334 | 2025-05-15 17:36
2025-05-15 05:38:08 | 20.42 | 48.38 | 45792.7348 | 2025-05-15 17:38
2025-05-15 05:40:09 | 20.41 | 48.62 | 45792.7362 | 2025-05-15 17:40
2025-05-15 05:42:08 | 20.42 | 48.65 | 45792.7376 | 2025-05-15 17:42
2025-05-15 05:44:09 | 20.43 | 48.66 | 45792.739 | 2025-05-15 17:44
2025-05-15 05:46:08 | 20.41 | 48.75 | 45792.7404 | 2025-05-15 17:46
2025-05-15 05:48:08 | 20.43 | 48.77 | 45792.7418 | 2025-05-15 17:48
2025-05-15 05:50:08 | 20.44 | 48.85 | 45792.7431 | 2025-05-15 17:50
2025-05-15 05:52:09 | 20.43 | 48.8 | 45792.7445 | 2025-05-15 17:52
2025-05-15 05:54:08 | 20.43 | 48.88 | 45792.7459 | 2025-05-15 17:54
2025-05-15 05:56:09 | 20.41 | 48.85 | 45792.7473 | 2025-05-15 17:56
2025-05-15 05:58:08 | 20.41 | 48.93 | 45792.7487 | 2025-05-15 17:58
2025-05-15 06:00:09 | 20.42 | 48.91 | 45792.7501 | 2025-05-15 18:00
2025-05-15 06:02:08 | 20.42 | 48.97 | 45792.7515 | 2025-05-15 18:02
2025-05-15 06:04:09 | 20.43 | 48.99 | 45792.7529 | 2025-05-15 18:04
2025-05-15 06:06:08 | 20.42 | 48.97 | 45792.7543 | 2025-05-15 18:06
2025-05-15 06:08:08 | 20.4 | 49.07 | 45792.7556 | 2025-05-15 18:08
2025-05-15 06:10:08 | 20.4 | 48.89 | 45792.757 | 2025-05-15 18:10
2025-05-15 06:12:09 | 20.39 | 48.87 | 45792.7584 | 2025-05-15 18:12
2025-05-15 06:14:08 | 20.41 | 48.96 | 45792.7598 | 2025-05-15 18:14
2025-05-15 06:16:08 | 20.4 | 48.91 | 45792.7612 | 2025-05-15 18:16
2025-05-15 06:18:08 | 20.41 | 48.95 | 45792.7626 | 2025-05-15 18:18
2025-05-15 06:20:09 | 20.4 | 48.9 | 45792.764 | 2025-05-15 18:20
2025-05-15 06:22:08 | 20.41 | 48.93 | 45792.7654 | 2025-05-15 18:22
2025-05-15 06:24:08 | 20.42 | 48.95 | 45792.7668 | 2025-05-15 18:24
2025-05-15 06:26:08 | 20.42 | 48.75 | 45792.7681 | 2025-05-15 18:26
2025-05-15 06:28:09 | 20.41 | 48.71 | 45792.7695 | 2025-05-15 18:28
2025-05-15 06:30:09 | 20.4 | 48.74 | 45792.7709 | 2025-05-15 18:30
2025-05-15 06:32:09 | 20.39 | 48.67 | 45792.7723 | 2025-05-15 18:32
2025-05-15 06:34:08 | 20.4 | 48.75 | 45792.7737 | 2025-05-15 18:34
2025-05-15 06:36:09 | 20.4 | 48.8 | 45792.7751 | 2025-05-15 18:36
2025-05-15 06:38:09 | 20.39 | 48.76 | 45792.7765 | 2025-05-15 18:38
2025-05-15 06:40:09 | 20.38 | 48.87 | 45792.7779 | 2025-05-15 18:40
2025-05-15 06:42:08 | 20.38 | 48.8 | 45792.7793 | 2025-05-15 18:42
2025-05-15 06:44:09 | 20.39 | 48.93 | 45792.7807 | 2025-05-15 18:44
2025-05-15 06:46:09 | 20.4 | 49.04 | 45792.782 | 2025-05-15 18:46
2025-05-15 06:48:09 | 20.41 | 48.97 | 45792.7834 | 2025-05-15 18:48
2025-05-15 06:50:08 | 20.41 | 49.1 | 45792.7848 | 2025-05-15 18:50
2025-05-15 06:52:09 | 20.43 | 49.28 | 45792.7862 | 2025-05-15 18:52
2025-05-15 06:54:08 | 20.44 | 49.36 | 45792.7876 | 2025-05-15 18:54
2025-05-15 06:56:09 | 20.42 | 49.68 | 45792.789 | 2025-05-15 18:56
2025-05-15 06:58:08 | 20.43 | 49.78 | 45792.7904 | 2025-05-15 18:58
2025-05-15 07:00:09 | 20.43 | 49.96 | 45792.7918 | 2025-05-15 19:00
2025-05-15 07:02:09 | 20.42 | 49.66 | 45792.7932 | 2025-05-15 19:02
2025-05-15 07:04:08 | 20.43 | 49.52 | 45792.7945 | 2025-05-15 19:04
2025-05-15 07:06:08 | 20.42 | 49.44 | 45792.7959 | 2025-05-15 19:06
2025-05-15 07:08:09 | 20.42 | 49.39 | 45792.7973 | 2025-05-15 19:08
2025-05-15 07:10:09 | 20.41 | 49.34 | 45792.7987 | 2025-05-15 19:10
2025-05-15 07:12:09 | 20.41 | 49.31 | 45792.8001 | 2025-05-15 19:12
2025-05-15 07:14:08 | 20.42 | 49.21 | 45792.8015 | 2025-05-15 19:14
2025-05-15 07:16:09 | 20.41 | 49.12 | 45792.8029 | 2025-05-15 19:16
2025-05-15 07:18:09 | 20.38 | 49.08 | 45792.8043 | 2025-05-15 19:18
2025-05-15 07:20:09 | 20.41 | 49.07 | 45792.8057 | 2025-05-15 19:20
2025-05-15 07:22:08 | 20.41 | 49.17 | 45792.807 | 2025-05-15 19:22
2025-05-15 07:24:09 | 20.41 | 49.73 | 45792.8084 | 2025-05-15 19:24
2025-05-15 07:26:09 | 20.42 | 50.05 | 45792.8098 | 2025-05-15 19:26
2025-05-15 07:28:09 | 20.43 | 50.06 | 45792.8112 | 2025-05-15 19:28
2025-05-15 07:30:08 | 20.42 | 50.03 | 45792.8126 | 2025-05-15 19:30
2025-05-15 07:32:09 | 20.42 | 50.01 | 45792.814 | 2025-05-15 19:32
2025-05-15 07:34:08 | 20.4 | 50.08 | 45792.8154 | 2025-05-15 19:34
2025-05-15 07:36:09 | 20.41 | 50.12 | 45792.8168 | 2025-05-15 19:36
2025-05-15 07:38:09 | 20.43 | 50.15 | 45792.8182 | 2025-05-15 19:38
2025-05-15 07:40:09 | 20.43 | 50.18 | 45792.8195 | 2025-05-15 19:40
2025-05-15 07:42:08 | 20.42 | 50.19 | 45792.8209 | 2025-05-15 19:42
2025-05-15 07:44:09 | 20.42 | 50.19 | 45792.8223 | 2025-05-15 19:44
2025-05-15 07:46:08 | 20.42 | 50.23 | 45792.8237 | 2025-05-15 19:46
2025-05-15 07:48:09 | 20.41 | 50.26 | 45792.8251 | 2025-05-15 19:48
2025-05-15 07:50:09 | 20.43 | 50.29 | 45792.8265 | 2025-05-15 19:50
2025-05-15 07:52:09 | 20.41 | 50.3 | 45792.8279 | 2025-05-15 19:52
2025-05-15 07:54:08 | 20.41 | 50.35 | 45792.8293 | 2025-05-15 19:54
2025-05-15 07:56:09 | 20.41 | 50.33 | 45792.8307 | 2025-05-15 19:56
2025-05-15 07:58:08 | 20.42 | 50.37 | 45792.832 | 2025-05-15 19:58
2025-05-15 08:00:09 | 20.41 | 50.38 | 45792.8334 | 2025-05-15 20:00
2025-05-15 08:02:08 | 20.42 | 50.37 | 45792.8348 | 2025-05-15 20:02
2025-05-15 08:04:09 | 20.41 | 50.33 | 45792.8362 | 2025-05-15 20:04
2025-05-15 08:06:09 | 20.42 | 50.32 | 45792.8376 | 2025-05-15 20:06
2025-05-15 08:08:09 | 20.42 | 50.28 | 45792.839 | 2025-05-15 20:08
2025-05-15 08:10:08 | 20.43 | 50.27 | 45792.8404 | 2025-05-15 20:10
2025-05-15 08:12:09 | 20.4 | 50.26 | 45792.8418 | 2025-05-15 20:12
2025-05-15 08:14:09 | 20.4 | 50.24 | 45792.8432 | 2025-05-15 20:14
2025-05-15 08:16:09 | 20.41 | 50.22 | 45792.8445 | 2025-05-15 20:16
2025-05-15 08:18:08 | 20.41 | 50.18 | 45792.8459 | 2025-05-15 20:18
2025-05-15 08:20:09 | 20.42 | 50.12 | 45792.8473 | 2025-05-15 20:20
2025-05-15 08:22:08 | 20.4 | 50.07 | 45792.8487 | 2025-05-15 20:22
2025-05-15 08:24:09 | 20.41 | 50.04 | 45792.8501 | 2025-05-15 20:24
2025-05-15 08:26:09 | 20.42 | 49.96 | 45792.8515 | 2025-05-15 20:26
2025-05-15 08:28:09 | 20.4 | 49.94 | 45792.8529 | 2025-05-15 20:28
2025-05-15 08:30:08 | 20.42 | 49.9 | 45792.8543 | 2025-05-15 20:30
2025-05-15 08:32:08 | 20.42 | 49.89 | 45792.8556 | 2025-05-15 20:32
2025-05-15 08:34:09 | 20.42 | 49.88 | 45792.857 | 2025-05-15 20:34
2025-05-15 08:36:09 | 20.39 | 49.89 | 45792.8584 | 2025-05-15 20:36
2025-05-15 08:38:08 | 20.38 | 49.95 | 45792.8598 | 2025-05-15 20:38
2025-05-15 08:40:09 | 20.41 | 49.98 | 45792.8612 | 2025-05-15 20:40
2025-05-15 08:42:08 | 20.4 | 49.99 | 45792.8626 | 2025-05-15 20:42
2025-05-15 08:44:09 | 20.4 | 49.98 | 45792.864 | 2025-05-15 20:44
2025-05-15 08:46:08 | 20.4 | 49.98 | 45792.8654 | 2025-05-15 20:46
2025-05-15 08:48:10 | 20.41 | 49.99 | 45792.8668 | 2025-05-15 20:48
2025-05-15 08:50:08 | 20.4 | 49.99 | 45792.8681 | 2025-05-15 20:50
2025-05-15 08:52:09 | 20.41 | 50.02 | 45792.8695 | 2025-05-15 20:52
2025-05-15 08:54:08 | 20.41 | 50.07 | 45792.8709 | 2025-05-15 20:54
2025-05-15 08:56:10 | 20.42 | 50.14 | 45792.8723 | 2025-05-15 20:56
2025-05-15 08:58:08 | 20.42 | 50.15 | 45792.8737 | 2025-05-15 20:58
2025-05-15 09:00:09 | 20.42 | 50.22 | 45792.8751 | 2025-05-15 21:00
2025-05-15 09:02:09 | 20.41 | 50.3 | 45792.8765 | 2025-05-15 21:02
2025-05-15 09:04:09 | 20.43 | 50.33 | 45792.8779 | 2025-05-15 21:04
2025-05-15 09:06:08 | 20.42 | 50.37 | 45792.8793 | 2025-05-15 21:06
2025-05-15 09:08:09 | 20.42 | 50.43 | 45792.8807 | 2025-05-15 21:08
2025-05-15 09:10:08 | 20.42 | 50.46 | 45792.882 | 2025-05-15 21:10
2025-05-15 09:12:08 | 20.44 | 50.49 | 45792.8834 | 2025-05-15 21:12
2025-05-15 09:14:08 | 20.42 | 50.52 | 45792.8848 | 2025-05-15 21:14
2025-05-15 09:16:09 | 20.42 | 50.55 | 45792.8862 | 2025-05-15 21:16
2025-05-15 09:18:08 | 20.42 | 50.56 | 45792.8876 | 2025-05-15 21:18
2025-05-15 09:20:09 | 20.41 | 50.59 | 45792.889 | 2025-05-15 21:20
2025-05-15 09:22:08 | 20.42 | 50.58 | 45792.8904 | 2025-05-15 21:22
2025-05-15 09:24:09 | 20.42 | 50.64 | 45792.8918 | 2025-05-15 21:24
2025-05-15 09:26:08 | 20.44 | 50.71 | 45792.8931 | 2025-05-15 21:26
2025-05-15 09:28:09 | 20.43 | 50.8 | 45792.8945 | 2025-05-15 21:28
2025-05-15 09:30:09 | 20.42 | 50.89 | 45792.8959 | 2025-05-15 21:30
2025-05-15 09:32:09 | 20.41 | 51.01 | 45792.8973 | 2025-05-15 21:32
2025-05-15 09:34:08 | 20.42 | 51.12 | 45792.8987 | 2025-05-15 21:34
2025-05-15 09:36:09 | 20.4 | 51.26 | 45792.9001 | 2025-05-15 21:36
2025-05-15 09:38:09 | 20.41 | 51.33 | 45792.9015 | 2025-05-15 21:38
2025-05-15 09:40:09 | 20.41 | 51.43 | 45792.9029 | 2025-05-15 21:40
2025-05-15 09:42:08 | 20.41 | 51.51 | 45792.9043 | 2025-05-15 21:42
2025-05-15 09:44:09 | 20.41 | 51.61 | 45792.9057 | 2025-05-15 21:44
2025-05-15 09:46:09 | 20.42 | 51.72 | 45792.907 | 2025-05-15 21:46
2025-05-15 09:48:09 | 20.42 | 51.86 | 45792.9084 | 2025-05-15 21:48
2025-05-15 09:50:08 | 20.42 | 51.98 | 45792.9098 | 2025-05-15 21:50
2025-05-15 09:52:09 | 20.43 | 52.05 | 45792.9112 | 2025-05-15 21:52
2025-05-15 09:54:08 | 20.43 | 52.21 | 45792.9126 | 2025-05-15 21:54
2025-05-15 09:56:08 | 20.42 | 52.33 | 45792.914 | 2025-05-15 21:56
2025-05-15 09:58:08 | 20.41 | 52.4 | 45792.9154 | 2025-05-15 21:58
2025-05-15 10:00:09 | 20.42 | 52.49 | 45792.9168 | 2025-05-15 22:00
2025-05-15 10:02:08 | 20.43 | 52.55 | 45792.9181 | 2025-05-15 22:02
2025-05-15 10:04:09 | 20.42 | 52.62 | 45792.9195 | 2025-05-15 22:04
2025-05-15 10:06:09 | 20.42 | 52.65 | 45792.9209 | 2025-05-15 22:06
2025-05-15 10:08:09 | 20.42 | 52.68 | 45792.9223 | 2025-05-15 22:08
2025-05-15 10:10:08 | 20.44 | 52.63 | 45792.9237 | 2025-05-15 22:10
2025-05-15 10:12:09 | 20.42 | 52.7 | 45792.9251 | 2025-05-15 22:12
2025-05-15 10:14:09 | 20.43 | 52.72 | 45792.9265 | 2025-05-15 22:14
2025-05-15 10:16:09 | 20.42 | 52.82 | 45792.9279 | 2025-05-15 22:16
2025-05-15 10:18:08 | 20.41 | 52.86 | 45792.9293 | 2025-05-15 22:18
2025-05-15 10:20:09 | 20.4 | 52.95 | 45792.9307 | 2025-05-15 22:20
2025-05-15 10:22:08 | 20.43 | 52.95 | 45792.932 | 2025-05-15 22:22
2025-05-15 10:24:09 | 20.43 | 53.01 | 45792.9334 | 2025-05-15 22:24
2025-05-15 10:26:08 | 20.43 | 53.07 | 45792.9348 | 2025-05-15 22:26
2025-05-15 10:28:09 | 20.42 | 53.11 | 45792.9362 | 2025-05-15 22:28
2025-05-15 10:30:08 | 20.43 | 53.09 | 45792.9376 | 2025-05-15 22:30
2025-05-15 10:32:08 | 20.44 | 53.09 | 45792.939 | 2025-05-15 22:32
2025-05-15 10:34:09 | 20.43 | 53.11 | 45792.9404 | 2025-05-15 22:34
2025-05-15 10:36:09 | 20.42 | 53.11 | 45792.9418 | 2025-05-15 22:36
2025-05-15 10:38:08 | 20.43 | 53.07 | 45792.9431 | 2025-05-15 22:38
2025-05-15 10:40:09 | 20.42 | 53.05 | 45792.9445 | 2025-05-15 22:40
2025-05-15 10:42:08 | 20.42 | 53.08 | 45792.9459 | 2025-05-15 22:42
2025-05-15 10:44:09 | 20.43 | 53.06 | 45792.9473 | 2025-05-15 22:44
2025-05-15 10:46:08 | 20.45 | 53.04 | 45792.9487 | 2025-05-15 22:46
2025-05-15 10:48:09 | 20.43 | 53.04 | 45792.9501 | 2025-05-15 22:48
2025-05-15 10:50:09 | 20.43 | 52.99 | 45792.9515 | 2025-05-15 22:50
2025-05-15 10:52:09 | 20.43 | 52.94 | 45792.9529 | 2025-05-15 22:52
2025-05-15 10:54:08 | 20.45 | 52.92 | 45792.9543 | 2025-05-15 22:54
2025-05-15 10:56:09 | 20.44 | 52.92 | 45792.9557 | 2025-05-15 22:56
2025-05-15 10:58:08 | 20.44 | 52.9 | 45792.957 | 2025-05-15 22:58
2025-05-15 11:00:09 | 20.43 | 52.91 | 45792.9584 | 2025-05-15 23:00
2025-05-15 11:02:08 | 20.42 | 52.93 | 45792.9598 | 2025-05-15 23:02
2025-05-15 11:04:08 | 20.44 | 52.94 | 45792.9612 | 2025-05-15 23:04
2025-05-15 11:06:09 | 20.41 | 52.94 | 45792.9626 | 2025-05-15 23:06
2025-05-15 11:08:09 | 20.42 | 52.92 | 45792.964 | 2025-05-15 23:08
2025-05-15 11:10:08 | 20.42 | 52.91 | 45792.9654 | 2025-05-15 23:10
2025-05-15 11:12:08 | 20.42 | 52.91 | 45792.9668 | 2025-05-15 23:12
2025-05-15 11:14:08 | 20.43 | 52.89 | 45792.9681 | 2025-05-15 23:14
2025-05-15 11:16:10 | 20.43 | 52.86 | 45792.9696 | 2025-05-15 23:16
2025-05-15 11:18:09 | 20.43 | 52.82 | 45792.9709 | 2025-05-15 23:18
2025-05-15 11:20:09 | 20.44 | 52.73 | 45792.9723 | 2025-05-15 23:20
2025-05-15 11:22:08 | 20.44 | 52.66 | 45792.9737 | 2025-05-15 23:22
2025-05-15 11:24:08 | 20.4 | 52.65 | 45792.9751 | 2025-05-15 23:24
2025-05-15 11:26:08 | 20.41 | 52.68 | 45792.9765 | 2025-05-15 23:26
2025-05-15 11:28:09 | 20.41 | 52.66 | 45792.9779 | 2025-05-15 23:28
2025-05-15 11:30:08 | 20.41 | 52.59 | 45792.9793 | 2025-05-15 23:30
2025-05-15 11:32:09 | 20.42 | 52.59 | 45792.9807 | 2025-05-15 23:32
2025-05-15 11:34:08 | 20.42 | 52.56 | 45792.982 | 2025-05-15 23:34
2025-05-15 11:36:09 | 20.43 | 52.49 | 45792.9834 | 2025-05-15 23:36
2025-05-15 11:38:09 | 20.42 | 52.42 | 45792.9848 | 2025-05-15 23:38
2025-05-15 11:40:09 | 20.39 | 52.39 | 45792.9862 | 2025-05-15 23:40
2025-05-15 11:42:08 | 20.39 | 52.43 | 45792.9876 | 2025-05-15 23:42
2025-05-15 11:44:08 | 20.4 | 52.38 | 45792.989 | 2025-05-15 23:44
2025-05-15 11:46:08 | 20.4 | 52.36 | 45792.9904 | 2025-05-15 23:46
2025-05-15 11:48:09 | 20.42 | 52.31 | 45792.9918 | 2025-05-15 23:48
2025-05-15 11:50:08 | 20.42 | 52.21 | 45792.9931 | 2025-05-15 23:50
2025-05-15 11:52:09 | 20.41 | 52.12 | 45792.9945 | 2025-05-15 23:52
2025-05-15 11:54:08 | 20.42 | 52.13 | 45792.9959 | 2025-05-15 23:54
2025-05-15 11:56:09 | 20.4 | 52.16 | 45792.9973 | 2025-05-15 23:56
2025-05-15 11:58:08 | 20.41 | 52.13 | 45792.9987 | 2025-05-15 23:58
2025-05-15 12:00:09 | 20.44 | 52.05 | 45793.0001 | 2025-05-16
2025-05-15 12:02:09 | 20.41 | 51.95 | 45793.0015 | 2025-05-16 00:02
2025-05-15 12:04:09 | 20.4 | 51.97 | 45793.0029 | 2025-05-16 00:04
2025-05-15 12:06:08 | 20.41 | 52.01 | 45793.0043 | 2025-05-16 00:06
2025-05-15 12:08:09 | 20.41 | 52 | 45793.0057 | 2025-05-16 00:08
2025-05-15 12:10:09 | 20.41 | 51.98 | 45793.007 | 2025-05-16 00:10
2025-05-15 12:12:08 | 20.42 | 51.88 | 45793.0084 | 2025-05-16 00:12
2025-05-15 12:14:08 | 20.41 | 51.81 | 45793.0098 | 2025-05-16 00:14
2025-05-15 12:16:09 | 20.41 | 51.8 | 45793.0112 | 2025-05-16 00:16
2025-05-15 12:18:08 | 20.4 | 51.82 | 45793.0126 | 2025-05-16 00:18
2025-05-15 12:20:08 | 20.41 | 51.8 | 45793.014 | 2025-05-16 00:20
2025-05-15 12:22:08 | 20.41 | 51.71 | 45793.0154 | 2025-05-16 00:22
2025-05-15 12:24:09 | 20.41 | 51.65 | 45793.0168 | 2025-05-16 00:24
2025-05-15 12:26:08 | 20.4 | 51.67 | 45793.0181 | 2025-05-16 00:26
2025-05-15 12:28:09 | 20.4 | 51.67 | 45793.0195 | 2025-05-16 00:28
2025-05-15 12:30:09 | 20.41 | 51.67 | 45793.0209 | 2025-05-16 00:30
2025-05-15 12:32:09 | 20.43 | 51.64 | 45793.0223 | 2025-05-16 00:32
2025-05-15 12:34:08 | 20.41 | 51.57 | 45793.0237 | 2025-05-16 00:34
2025-05-15 12:36:09 | 20.4 | 51.52 | 45793.0251 | 2025-05-16 00:36
2025-05-15 12:38:09 | 20.39 | 51.54 | 45793.0265 | 2025-05-16 00:38
2025-05-15 12:40:09 | 20.41 | 51.53 | 45793.0279 | 2025-05-16 00:40
2025-05-15 12:42:08 | 20.4 | 51.53 | 45793.0293 | 2025-05-16 00:42
2025-05-15 12:44:09 | 20.43 | 51.51 | 45793.0307 | 2025-05-16 00:44
2025-05-15 12:46:09 | 20.43 | 51.41 | 45793.032 | 2025-05-16 00:46
2025-05-15 12:48:09 | 20.43 | 51.35 | 45793.0334 | 2025-05-16 00:48
2025-05-15 12:50:08 | 20.41 | 51.38 | 45793.0348 | 2025-05-16 00:50
2025-05-15 12:52:09 | 20.41 | 51.4 | 45793.0362 | 2025-05-16 00:52
2025-05-15 12:54:09 | 20.42 | 51.38 | 45793.0376 | 2025-05-16 00:54
2025-05-15 12:56:09 | 20.41 | 51.36 | 45793.039 | 2025-05-16 00:56
2025-05-15 12:58:08 | 20.42 | 51.28 | 45793.0404 | 2025-05-16 00:58
2025-05-15 13:00:09 | 20.41 | 51.22 | 45793.0418 | 2025-05-16 01:00
2025-05-15 13:02:08 | 20.4 | 51.26 | 45793.0431 | 2025-05-16 01:02
2025-05-15 13:04:08 | 20.41 | 51.27 | 45793.0445 | 2025-05-16 01:04
2025-05-15 13:06:08 | 20.42 | 51.21 | 45793.0459 | 2025-05-16 01:06
2025-05-15 13:08:09 | 20.43 | 51.14 | 45793.0473 | 2025-05-16 01:08
2025-05-15 13:10:08 | 20.42 | 51.07 | 45793.0487 | 2025-05-16 01:10
2025-05-15 13:12:09 | 20.4 | 51.03 | 45793.0501 | 2025-05-16 01:12
2025-05-15 13:14:09 | 20.4 | 51.05 | 45793.0515 | 2025-05-16 01:14
2025-05-15 13:16:09 | 20.39 | 51.07 | 45793.0529 | 2025-05-16 01:16
2025-05-15 13:18:08 | 20.41 | 51.01 | 45793.0543 | 2025-05-16 01:18
2025-05-15 13:20:09 | 20.41 | 50.94 | 45793.0557 | 2025-05-16 01:20
2025-05-15 13:22:09 | 20.4 | 50.86 | 45793.057 | 2025-05-16 01:22
2025-05-15 13:24:09 | 20.39 | 50.82 | 45793.0584 | 2025-05-16 01:24
2025-05-15 13:26:08 | 20.4 | 50.81 | 45793.0598 | 2025-05-16 01:26
2025-05-15 13:28:09 | 20.39 | 50.84 | 45793.0612 | 2025-05-16 01:28
2025-05-15 13:30:08 | 20.4 | 50.84 | 45793.0626 | 2025-05-16 01:30
2025-05-15 13:32:08 | 20.42 | 50.77 | 45793.064 | 2025-05-16 01:32
2025-05-15 13:34:08 | 20.42 | 50.66 | 45793.0654 | 2025-05-16 01:34
2025-05-15 13:36:09 | 20.42 | 50.62 | 45793.0668 | 2025-05-16 01:36
2025-05-15 13:38:08 | 20.42 | 50.63 | 45793.0681 | 2025-05-16 01:38
2025-05-15 13:40:08 | 20.39 | 50.65 | 45793.0695 | 2025-05-16 01:40
2025-05-15 13:42:08 | 20.42 | 50.65 | 45793.0709 | 2025-05-16 01:42
2025-05-15 13:44:09 | 20.41 | 50.63 | 45793.0723 | 2025-05-16 01:44
2025-05-15 13:46:08 | 20.41 | 50.59 | 45793.0737 | 2025-05-16 01:46
2025-05-15 13:48:08 | 20.41 | 50.56 | 45793.0751 | 2025-05-16 01:48
2025-05-15 13:50:08 | 20.39 | 50.57 | 45793.0765 | 2025-05-16 01:50
2025-05-15 13:52:09 | 20.38 | 50.64 | 45793.0779 | 2025-05-16 01:52
2025-05-15 13:54:08 | 20.39 | 50.66 | 45793.0793 | 2025-05-16 01:54
2025-05-15 13:56:09 | 20.39 | 50.63 | 45793.0807 | 2025-05-16 01:56
2025-05-15 13:58:08 | 20.41 | 50.52 | 45793.082 | 2025-05-16 01:58
2025-05-15 14:00:09 | 20.4 | 50.47 | 45793.0834 | 2025-05-16 02:00
2025-05-15 14:02:08 | 20.38 | 50.52 | 45793.0848 | 2025-05-16 02:02
2025-05-15 14:04:09 | 20.37 | 50.55 | 45793.0862 | 2025-05-16 02:04
2025-05-15 14:06:08 | 20.41 | 50.58 | 45793.0876 | 2025-05-16 02:06
2025-05-15 14:08:09 | 20.41 | 50.54 | 45793.089 | 2025-05-16 02:08
2025-05-15 14:10:08 | 20.42 | 50.48 | 45793.0904 | 2025-05-16 02:10
2025-05-15 14:12:08 | 20.4 | 50.41 | 45793.0918 | 2025-05-16 02:12
2025-05-15 14:14:08 | 20.38 | 50.41 | 45793.0931 | 2025-05-16 02:14
2025-05-15 14:16:09 | 20.37 | 50.47 | 45793.0945 | 2025-05-16 02:16
2025-05-15 14:18:08 | 20.39 | 50.51 | 45793.0959 | 2025-05-16 02:18
2025-05-15 14:20:09 | 20.4 | 50.49 | 45793.0973 | 2025-05-16 02:20
2025-05-15 14:22:08 | 20.39 | 50.44 | 45793.0987 | 2025-05-16 02:22
2025-05-15 14:24:09 | 20.41 | 50.4 | 45793.1001 | 2025-05-16 02:24
2025-05-15 14:26:08 | 20.39 | 50.38 | 45793.1015 | 2025-05-16 02:26
2025-05-15 14:28:08 | 20.38 | 50.44 | 45793.1029 | 2025-05-16 02:28
2025-05-15 14:30:08 | 20.38 | 50.52 | 45793.1043 | 2025-05-16 02:30
2025-05-15 14:32:09 | 20.4 | 50.51 | 45793.1057 | 2025-05-16 02:32
2025-05-15 14:34:08 | 20.42 | 50.5 | 45793.107 | 2025-05-16 02:34
2025-05-15 14:36:09 | 20.4 | 50.45 | 45793.1084 | 2025-05-16 02:36
2025-05-15 14:38:09 | 20.4 | 50.49 | 45793.1098 | 2025-05-16 02:38
2025-05-15 14:40:09 | 20.39 | 50.55 | 45793.1112 | 2025-05-16 02:40
2025-05-15 14:42:08 | 20.41 | 50.58 | 45793.1126 | 2025-05-16 02:42
2025-05-15 14:44:09 | 20.41 | 50.62 | 45793.114 | 2025-05-16 02:44
2025-05-15 14:46:09 | 20.41 | 50.58 | 45793.1154 | 2025-05-16 02:46
2025-05-15 14:48:09 | 20.42 | 50.56 | 45793.1168 | 2025-05-16 02:48
2025-05-15 14:50:08 | 20.41 | 50.62 | 45793.1181 | 2025-05-16 02:50
2025-05-15 14:52:09 | 20.4 | 50.71 | 45793.1195 | 2025-05-16 02:52
2025-05-15 14:54:09 | 20.4 | 50.74 | 45793.1209 | 2025-05-16 02:54
2025-05-15 14:56:09 | 20.4 | 50.74 | 45793.1223 | 2025-05-16 02:56
2025-05-15 14:58:08 | 20.42 | 50.74 | 45793.1237 | 2025-05-16 02:58
2025-05-15 15:00:09 | 20.4 | 50.76 | 45793.1251 | 2025-05-16 03:00
2025-05-15 15:02:09 | 20.4 | 50.81 | 45793.1265 | 2025-05-16 03:02
2025-05-15 15:04:09 | 20.39 | 50.88 | 45793.1279 | 2025-05-16 03:04
2025-05-15 15:06:08 | 20.4 | 50.95 | 45793.1293 | 2025-05-16 03:06
2025-05-15 15:08:09 | 20.4 | 50.9 | 45793.1307 | 2025-05-16 03:08
2025-05-15 15:10:08 | 20.41 | 50.85 | 45793.132 | 2025-05-16 03:10
2025-05-15 15:12:09 | 20.42 | 50.87 | 45793.1334 | 2025-05-16 03:12
2025-05-15 15:14:08 | 20.38 | 50.97 | 45793.1348 | 2025-05-16 03:14
2025-05-15 15:16:09 | 20.4 | 51 | 45793.1362 | 2025-05-16 03:16
2025-05-15 15:18:09 | 20.41 | 50.99 | 45793.1376 | 2025-05-16 03:18
2025-05-15 15:20:09 | 20.42 | 50.93 | 45793.139 | 2025-05-16 03:20
2025-05-15 15:22:08 | 20.43 | 50.88 | 45793.1404 | 2025-05-16 03:22
2025-05-15 15:24:09 | 20.4 | 50.91 | 45793.1418 | 2025-05-16 03:24
2025-05-15 15:26:08 | 20.4 | 50.96 | 45793.1431 | 2025-05-16 03:26
2025-05-15 15:28:08 | 20.41 | 50.95 | 45793.1445 | 2025-05-16 03:28
2025-05-15 15:30:08 | 20.4 | 50.89 | 45793.1459 | 2025-05-16 03:30
2025-05-15 15:32:09 | 20.41 | 50.82 | 45793.1473 | 2025-05-16 03:32
2025-05-15 15:34:08 | 20.41 | 50.78 | 45793.1487 | 2025-05-16 03:34
2025-05-15 15:36:09 | 20.41 | 50.82 | 45793.1501 | 2025-05-16 03:36
2025-05-15 15:38:09 | 20.4 | 50.83 | 45793.1515 | 2025-05-16 03:38
2025-05-15 15:40:09 | 20.4 | 50.82 | 45793.1529 | 2025-05-16 03:40
2025-05-15 15:42:08 | 20.41 | 50.8 | 45793.1543 | 2025-05-16 03:42
2025-05-15 15:44:09 | 20.42 | 50.71 | 45793.1557 | 2025-05-16 03:44
2025-05-15 15:46:08 | 20.4 | 50.67 | 45793.157 | 2025-05-16 03:46
2025-05-15 15:48:08 | 20.39 | 50.73 | 45793.1584 | 2025-05-16 03:48
2025-05-15 15:50:08 | 20.41 | 50.77 | 45793.1598 | 2025-05-16 03:50
2025-05-15 15:52:09 | 20.42 | 50.73 | 45793.1612 | 2025-05-16 03:52
2025-05-15 15:54:08 | 20.41 | 50.71 | 45793.1626 | 2025-05-16 03:54
2025-05-15 15:56:09 | 20.42 | 50.63 | 45793.164 | 2025-05-16 03:56
2025-05-15 15:58:08 | 20.41 | 50.61 | 45793.1654 | 2025-05-16 03:58
2025-05-15 16:00:09 | 20.39 | 50.66 | 45793.1668 | 2025-05-16 04:00
2025-05-15 16:02:08 | 20.38 | 50.71 | 45793.1681 | 2025-05-16 04:02
2025-05-15 16:04:09 | 20.42 | 50.7 | 45793.1695 | 2025-05-16 04:04
2025-05-15 16:06:09 | 20.42 | 50.65 | 45793.1709 | 2025-05-16 04:06
2025-05-15 16:08:09 | 20.43 | 50.59 | 45793.1723 | 2025-05-16 04:08
2025-05-15 16:10:08 | 20.4 | 50.63 | 45793.1737 | 2025-05-16 04:10
2025-05-15 16:12:09 | 20.42 | 50.71 | 45793.1751 | 2025-05-16 04:12
2025-05-15 16:14:08 | 20.42 | 50.7 | 45793.1765 | 2025-05-16 04:14
2025-05-15 16:16:08 | 20.43 | 50.66 | 45793.1779 | 2025-05-16 04:16
2025-05-15 16:18:08 | 20.42 | 50.59 | 45793.1793 | 2025-05-16 04:18
2025-05-15 16:20:09 | 20.42 | 50.58 | 45793.1807 | 2025-05-16 04:20
2025-05-15 16:22:08 | 20.4 | 50.66 | 45793.182 | 2025-05-16 04:22
2025-05-15 16:24:09 | 20.39 | 50.72 | 45793.1834 | 2025-05-16 04:24
2025-05-15 16:26:08 | 20.41 | 50.71 | 45793.1848 | 2025-05-16 04:26
2025-05-15 16:28:08 | 20.43 | 50.66 | 45793.1862 | 2025-05-16 04:28
2025-05-15 16:30:08 | 20.42 | 50.59 | 45793.1876 | 2025-05-16 04:30
2025-05-15 16:32:09 | 20.39 | 50.57 | 45793.189 | 2025-05-16 04:32
2025-05-15 16:34:08 | 20.38 | 50.67 | 45793.1904 | 2025-05-16 04:34
2025-05-15 16:36:09 | 20.37 | 50.7 | 45793.1918 | 2025-05-16 04:36
2025-05-15 16:38:08 | 20.4 | 50.7 | 45793.1931 | 2025-05-16 04:38
2025-05-15 16:40:08 | 20.41 | 50.63 | 45793.1945 | 2025-05-16 04:40
2025-05-15 16:42:08 | 20.4 | 50.6 | 45793.1959 | 2025-05-16 04:42
2025-05-15 16:44:09 | 20.39 | 50.62 | 45793.1973 | 2025-05-16 04:44
2025-05-15 16:46:08 | 20.4 | 50.67 | 45793.1987 | 2025-05-16 04:46
2025-05-15 16:48:08 | 20.4 | 50.67 | 45793.2001 | 2025-05-16 04:48
2025-05-15 16:50:08 | 20.42 | 50.64 | 45793.2015 | 2025-05-16 04:50
2025-05-15 16:52:09 | 20.39 | 50.62 | 45793.2029 | 2025-05-16 04:52
2025-05-15 16:54:08 | 20.41 | 50.56 | 45793.2043 | 2025-05-16 04:54
2025-05-15 16:56:09 | 20.41 | 50.58 | 45793.2057 | 2025-05-16 04:56
2025-05-15 16:58:08 | 20.4 | 50.62 | 45793.207 | 2025-05-16 04:58
2025-05-15 17:00:09 | 20.4 | 50.65 | 45793.2084 | 2025-05-16 05:00
2025-05-15 17:02:08 | 20.4 | 50.6 | 45793.2098 | 2025-05-16 05:02
2025-05-15 17:04:09 | 20.4 | 50.53 | 45793.2112 | 2025-05-16 05:04
2025-05-15 17:06:09 | 20.4 | 50.55 | 45793.2126 | 2025-05-16 05:06
2025-05-15 17:08:09 | 20.39 | 50.61 | 45793.214 | 2025-05-16 05:08
2025-05-15 17:10:08 | 20.39 | 50.6 | 45793.2154 | 2025-05-16 05:10
2025-05-15 17:12:09 | 20.41 | 50.58 | 45793.2168 | 2025-05-16 05:12
2025-05-15 17:14:09 | 20.4 | 50.49 | 45793.2182 | 2025-05-16 05:14
2025-05-15 17:16:09 | 20.4 | 50.47 | 45793.2195 | 2025-05-16 05:16
2025-05-15 17:18:08 | 20.4 | 50.52 | 45793.2209 | 2025-05-16 05:18
2025-05-15 17:20:09 | 20.4 | 50.52 | 45793.2223 | 2025-05-16 05:20
2025-05-15 17:22:08 | 20.41 | 50.49 | 45793.2237 | 2025-05-16 05:22
2025-05-15 17:24:09 | 20.42 | 50.45 | 45793.2251 | 2025-05-16 05:24
2025-05-15 17:26:08 | 20.41 | 50.37 | 45793.2265 | 2025-05-16 05:26
2025-05-15 17:28:09 | 20.4 | 50.41 | 45793.2279 | 2025-05-16 05:28
2025-05-15 17:30:08 | 20.4 | 50.47 | 45793.2293 | 2025-05-16 05:30
2025-05-15 17:32:08 | 20.4 | 50.49 | 45793.2306 | 2025-05-16 05:32
2025-05-15 17:34:08 | 20.4 | 50.49 | 45793.232 | 2025-05-16 05:34
2025-05-15 17:36:09 | 20.41 | 50.46 | 45793.2334 | 2025-05-16 05:36
2025-05-15 17:38:08 | 20.4 | 50.39 | 45793.2348 | 2025-05-16 05:38
2025-05-15 17:40:09 | 20.39 | 50.39 | 45793.2362 | 2025-05-16 05:40
2025-05-15 17:42:09 | 20.38 | 50.43 | 45793.2376 | 2025-05-16 05:42
2025-05-15 17:44:09 | 20.39 | 50.43 | 45793.239 | 2025-05-16 05:44
2025-05-15 17:46:08 | 20.42 | 50.39 | 45793.2404 | 2025-05-16 05:46
2025-05-15 17:48:09 | 20.41 | 50.38 | 45793.2418 | 2025-05-16 05:48
2025-05-15 17:50:08 | 20.4 | 50.31 | 45793.2431 | 2025-05-16 05:50
2025-05-15 17:52:08 | 20.39 | 50.3 | 45793.2445 | 2025-05-16 05:52
2025-05-15 17:54:08 | 20.38 | 50.36 | 45793.2459 | 2025-05-16 05:54
2025-05-15 17:56:09 | 20.4 | 50.37 | 45793.2473 | 2025-05-16 05:56
2025-05-15 17:58:08 | 20.4 | 50.37 | 45793.2487 | 2025-05-16 05:58
2025-05-15 18:00:09 | 20.42 | 50.35 | 45793.2501 | 2025-05-16 06:00
2025-05-15 18:02:09 | 20.42 | 50.29 | 45793.2515 | 2025-05-16 06:02
2025-05-15 18:04:09 | 20.4 | 50.25 | 45793.2529 | 2025-05-16 06:04
2025-05-15 18:06:08 | 20.39 | 50.32 | 45793.2543 | 2025-05-16 06:06
2025-05-15 18:08:09 | 20.4 | 50.36 | 45793.2557 | 2025-05-16 06:08
2025-05-15 18:10:08 | 20.41 | 50.35 | 45793.257 | 2025-05-16 06:10
2025-05-15 18:12:08 | 20.39 | 50.34 | 45793.2584 | 2025-05-16 06:12
2025-05-15 18:14:08 | 20.4 | 50.33 | 45793.2598 | 2025-05-16 06:14
2025-05-15 18:16:09 | 20.4 | 50.3 | 45793.2612 | 2025-05-16 06:16
2025-05-15 18:18:08 | 20.38 | 50.32 | 45793.2626 | 2025-05-16 06:18
2025-05-15 18:20:09 | 20.38 | 50.39 | 45793.264 | 2025-05-16 06:20
2025-05-15 18:22:08 | 20.38 | 50.4 | 45793.2654 | 2025-05-16 06:22
2025-05-15 18:24:09 | 20.41 | 50.4 | 45793.2668 | 2025-05-16 06:24
2025-05-15 18:26:08 | 20.42 | 50.31 | 45793.2681 | 2025-05-16 06:26
2025-05-15 18:28:09 | 20.4 | 50.29 | 45793.2695 | 2025-05-16 06:28
2025-05-15 18:30:08 | 20.39 | 50.33 | 45793.2709 | 2025-05-16 06:30
2025-05-15 18:32:08 | 20.39 | 50.36 | 45793.2723 | 2025-05-16 06:32
2025-05-15 18:34:08 | 20.41 | 50.38 | 45793.2737 | 2025-05-16 06:34
2025-05-15 18:36:09 | 20.39 | 50.38 | 45793.2751 | 2025-05-16 06:36
2025-05-15 18:38:09 | 20.41 | 50.36 | 45793.2765 | 2025-05-16 06:38
2025-05-15 18:40:09 | 20.41 | 50.3 | 45793.2779 | 2025-05-16 06:40
2025-05-15 18:42:08 | 20.4 | 50.32 | 45793.2793 | 2025-05-16 06:42
2025-05-15 18:44:09 | 20.4 | 50.37 | 45793.2807 | 2025-05-16 06:44
2025-05-15 18:46:08 | 20.41 | 50.38 | 45793.282 | 2025-05-16 06:46
2025-05-15 18:48:08 | 20.4 | 50.38 | 45793.2834 | 2025-05-16 06:48
2025-05-15 18:50:08 | 20.42 | 50.36 | 45793.2848 | 2025-05-16 06:50
2025-05-15 18:52:09 | 20.41 | 50.29 | 45793.2862 | 2025-05-16 06:52
2025-05-15 18:54:08 | 20.39 | 50.3 | 45793.2876 | 2025-05-16 06:54
2025-05-15 18:56:09 | 20.38 | 50.33 | 45793.289 | 2025-05-16 06:56
2025-05-15 18:58:08 | 20.39 | 50.36 | 45793.2904 | 2025-05-16 06:58
2025-05-15 19:00:09 | 20.4 | 50.36 | 45793.2918 | 2025-05-16 07:00
2025-05-15 19:02:08 | 20.41 | 50.36 | 45793.2931 | 2025-05-16 07:02
2025-05-15 19:04:09 | 20.41 | 50.37 | 45793.2945 | 2025-05-16 07:04
2025-05-15 19:06:08 | 20.43 | 50.38 | 45793.2959 | 2025-05-16 07:06
2025-05-15 19:08:09 | 20.42 | 50.39 | 45793.2973 | 2025-05-16 07:08
2025-05-15 19:10:08 | 20.42 | 50.4 | 45793.2987 | 2025-05-16 07:10
2025-05-15 19:12:09 | 20.4 | 50.35 | 45793.3001 | 2025-05-16 07:12
2025-05-15 19:14:08 | 20.4 | 50.37 | 45793.3015 | 2025-05-16 07:14
2025-05-15 19:16:09 | 20.38 | 50.4 | 45793.3029 | 2025-05-16 07:16
2025-05-15 19:18:09 | 20.38 | 50.44 | 45793.3043 | 2025-05-16 07:18
2025-05-15 19:20:09 | 20.39 | 50.42 | 45793.3057 | 2025-05-16 07:20
2025-05-15 19:22:08 | 20.41 | 50.43 | 45793.307 | 2025-05-16 07:22
2025-05-15 19:24:09 | 20.4 | 50.39 | 45793.3084 | 2025-05-16 07:24
2025-05-15 19:26:09 | 20.4 | 50.38 | 45793.3098 | 2025-05-16 07:26
2025-05-15 19:28:09 | 20.39 | 50.44 | 45793.3112 | 2025-05-16 07:28
2025-05-15 19:30:08 | 20.4 | 50.45 | 45793.3126 | 2025-05-16 07:30
2025-05-15 19:32:09 | 20.4 | 50.44 | 45793.314 | 2025-05-16 07:32
2025-05-15 19:34:09 | 20.42 | 50.43 | 45793.3154 | 2025-05-16 07:34
2025-05-15 19:36:08 | 20.4 | 50.38 | 45793.3168 | 2025-05-16 07:36
2025-05-15 19:38:08 | 20.38 | 50.41 | 45793.3181 | 2025-05-16 07:38
2025-05-15 19:40:09 | 20.38 | 50.47 | 45793.3195 | 2025-05-16 07:40
2025-05-15 19:42:08 | 20.4 | 50.45 | 45793.3209 | 2025-05-16 07:42
2025-05-15 19:44:08 | 20.41 | 50.43 | 45793.3223 | 2025-05-16 07:44
2025-05-15 19:46:08 | 20.41 | 50.37 | 45793.3237 | 2025-05-16 07:46
2025-05-15 19:48:09 | 20.38 | 50.38 | 45793.3251 | 2025-05-16 07:48
2025-05-15 19:50:08 | 20.38 | 50.44 | 45793.3265 | 2025-05-16 07:50
2025-05-15 19:52:09 | 20.38 | 50.46 | 45793.3279 | 2025-05-16 07:52
2025-05-15 19:54:08 | 20.39 | 50.43 | 45793.3293 | 2025-05-16 07:54
2025-05-15 19:56:09 | 20.39 | 50.43 | 45793.3307 | 2025-05-16 07:56
2025-05-15 19:58:08 | 20.41 | 50.35 | 45793.332 | 2025-05-16 07:58
2025-05-15 20:00:09 | 20.38 | 50.33 | 45793.3334 | 2025-05-16 08:00
2025-05-15 20:02:08 | 20.38 | 50.36 | 45793.3348 | 2025-05-16 08:02
2025-05-15 20:04:08 | 20.38 | 50.41 | 45793.3362 | 2025-05-16 08:04
2025-05-15 20:06:08 | 20.41 | 50.38 | 45793.3376 | 2025-05-16 08:06
2025-05-15 20:08:09 | 20.4 | 50.39 | 45793.339 | 2025-05-16 08:08
2025-05-15 20:10:08 | 20.42 | 50.35 | 45793.3404 | 2025-05-16 08:10
2025-05-15 20:12:08 | 20.42 | 50.32 | 45793.3418 | 2025-05-16 08:12
2025-05-15 20:14:09 | 20.42 | 50.3 | 45793.3432 | 2025-05-16 08:14
2025-05-15 20:16:09 | 20.41 | 50.39 | 45793.3445 | 2025-05-16 08:16
2025-05-15 20:18:08 | 20.42 | 50.41 | 45793.3459 | 2025-05-16 08:18
2025-05-15 20:20:09 | 20.41 | 50.43 | 45793.3473 | 2025-05-16 08:20
2025-05-15 20:22:08 | 20.43 | 50.42 | 45793.3487 | 2025-05-16 08:22
2025-05-15 20:24:09 | 20.41 | 50.42 | 45793.3501 | 2025-05-16 08:24
2025-05-15 20:26:08 | 20.41 | 50.48 | 45793.3515 | 2025-05-16 08:26
2025-05-15 20:28:09 | 20.42 | 50.49 | 45793.3529 | 2025-05-16 08:28
2025-05-15 20:30:08 | 20.41 | 50.56 | 45793.3543 | 2025-05-16 08:30
2025-05-15 20:32:08 | 20.4 | 50.56 | 45793.3556 | 2025-05-16 08:32
2025-05-15 20:34:09 | 20.4 | 50.53 | 45793.357 | 2025-05-16 08:34
2025-05-15 20:36:09 | 20.41 | 50.52 | 45793.3584 | 2025-05-16 08:36
2025-05-15 20:38:08 | 20.39 | 50.6 | 45793.3598 | 2025-05-16 08:38
2025-05-15 20:40:09 | 20.4 | 50.61 | 45793.3612 | 2025-05-16 08:40
2025-05-15 20:42:08 | 20.42 | 50.61 | 45793.3626 | 2025-05-16 08:42
2025-05-15 20:44:09 | 20.39 | 50.61 | 45793.364 | 2025-05-16 08:44
2025-05-15 20:46:08 | 20.4 | 50.56 | 45793.3654 | 2025-05-16 08:46
2025-05-15 20:48:09 | 20.41 | 50.56 | 45793.3668 | 2025-05-16 08:48
2025-05-15 20:50:08 | 20.39 | 50.6 | 45793.3681 | 2025-05-16 08:50
2025-05-15 20:52:08 | 20.39 | 50.61 | 45793.3695 | 2025-05-16 08:52
2025-05-15 20:54:08 | 20.4 | 50.61 | 45793.3709 | 2025-05-16 08:54
2025-05-15 20:56:09 | 20.43 | 50.59 | 45793.3723 | 2025-05-16 08:56
2025-05-15 20:58:08 | 20.41 | 50.54 | 45793.3737 | 2025-05-16 08:58
2025-05-15 21:00:09 | 20.4 | 50.53 | 45793.3751 | 2025-05-16 09:00
2025-05-15 21:02:09 | 20.39 | 50.58 | 45793.3765 | 2025-05-16 09:02
2025-05-15 21:04:09 | 20.4 | 50.63 | 45793.3779 | 2025-05-16 09:04
2025-05-15 21:06:08 | 20.41 | 50.63 | 45793.3793 | 2025-05-16 09:06
2025-05-15 21:08:09 | 20.41 | 50.66 | 45793.3807 | 2025-05-16 09:08
2025-05-15 21:10:08 | 20.42 | 50.62 | 45793.382 | 2025-05-16 09:10
2025-05-15 21:12:08 | 20.41 | 50.68 | 45793.3834 | 2025-05-16 09:12
2025-05-15 21:14:08 | 20.41 | 50.81 | 45793.3848 | 2025-05-16 09:14
2025-05-15 21:16:09 | 20.42 | 50.89 | 45793.3862 | 2025-05-16 09:16
2025-05-15 21:18:08 | 20.43 | 50.97 | 45793.3876 | 2025-05-16 09:18
2025-05-15 21:20:09 | 20.41 | 51.03 | 45793.389 | 2025-05-16 09:20
2025-05-15 21:22:08 | 20.4 | 51.12 | 45793.3904 | 2025-05-16 09:22
2025-05-15 21:24:09 | 20.4 | 51.26 | 45793.3918 | 2025-05-16 09:24
2025-05-15 21:26:08 | 20.4 | 51.41 | 45793.3931 | 2025-05-16 09:26
2025-05-15 21:28:09 | 20.42 | 51.5 | 45793.3945 | 2025-05-16 09:28
2025-05-15 21:30:09 | 20.42 | 51.52 | 45793.3959 | 2025-05-16 09:30
2025-05-15 21:32:09 | 20.41 | 51.55 | 45793.3973 | 2025-05-16 09:32
2025-05-15 21:34:08 | 20.4 | 51.66 | 45793.3987 | 2025-05-16 09:34
2025-05-15 21:36:09 | 20.39 | 51.79 | 45793.4001 | 2025-05-16 09:36
2025-05-15 21:38:08 | 20.4 | 51.83 | 45793.4015 | 2025-05-16 09:38
2025-05-15 21:40:08 | 20.42 | 51.88 | 45793.4029 | 2025-05-16 09:40
2025-05-15 21:42:08 | 20.42 | 51.83 | 45793.4043 | 2025-05-16 09:42
2025-05-15 21:44:09 | 20.41 | 51.82 | 45793.4057 | 2025-05-16 09:44
2025-05-15 21:46:08 | 20.39 | 51.87 | 45793.407 | 2025-05-16 09:46
2025-05-15 21:48:08 | 20.38 | 52.02 | 45793.4084 | 2025-05-16 09:48
2025-05-15 21:50:08 | 20.39 | 52.06 | 45793.4098 | 2025-05-16 09:50
2025-05-15 21:52:09 | 20.42 | 52.09 | 45793.4112 | 2025-05-16 09:52
2025-05-15 21:54:08 | 20.4 | 52.03 | 45793.4126 | 2025-05-16 09:54
2025-05-15 21:56:09 | 20.41 | 52.06 | 45793.414 | 2025-05-16 09:56
2025-05-15 21:58:08 | 20.39 | 52.15 | 45793.4154 | 2025-05-16 09:58
2025-05-15 22:00:09 | 20.4 | 52.18 | 45793.4168 | 2025-05-16 10:00
2025-05-15 22:02:08 | 20.39 | 52.18 | 45793.4181 | 2025-05-16 10:02
2025-05-15 22:04:08 | 20.39 | 52.18 | 45793.4195 | 2025-05-16 10:04
2025-05-15 22:06:08 | 20.4 | 52.09 | 45793.4209 | 2025-05-16 10:06
2025-05-15 22:08:09 | 20.4 | 52.11 | 45793.4223 | 2025-05-16 10:08
2025-05-15 22:10:08 | 20.39 | 52.2 | 45793.4237 | 2025-05-16 10:10
2025-05-15 22:12:09 | 20.41 | 52.29 | 45793.4251 | 2025-05-16 10:12
2025-05-15 22:14:08 | 20.42 | 52.27 | 45793.4265 | 2025-05-16 10:14
2025-05-15 22:16:09 | 20.41 | 52.19 | 45793.4279 | 2025-05-16 10:16
2025-05-15 22:18:08 | 20.43 | 52.19 | 45793.4293 | 2025-05-16 10:18
2025-05-15 22:20:08 | 20.41 | 52.16 | 45793.4306 | 2025-05-16 10:20
2025-05-15 22:22:08 | 20.42 | 52.1 | 45793.432 | 2025-05-16 10:22
2025-05-15 22:24:09 | 20.42 | 52.11 | 45793.4334 | 2025-05-16 10:24
2025-05-15 22:26:08 | 20.42 | 52.19 | 45793.4348 | 2025-05-16 10:26
2025-05-15 22:28:09 | 20.43 | 52.26 | 45793.4362 | 2025-05-16 10:28
2025-05-15 22:30:09 | 20.46 | 52.21 | 45793.4376 | 2025-05-16 10:30
2025-05-15 22:32:09 | 20.47 | 52.14 | 45793.439 | 2025-05-16 10:32
2025-05-15 22:34:08 | 20.48 | 52.11 | 45793.4404 | 2025-05-16 10:34
2025-05-15 22:36:09 | 20.46 | 52.14 | 45793.4418 | 2025-05-16 10:36
2025-05-15 22:38:09 | 20.45 | 52.28 | 45793.4432 | 2025-05-16 10:38
2025-05-15 22:40:09 | 20.43 | 52.39 | 45793.4445 | 2025-05-16 10:40
2025-05-15 22:42:08 | 20.45 | 52.48 | 45793.4459 | 2025-05-16 10:42
2025-05-15 22:44:09 | 20.43 | 52.43 | 45793.4473 | 2025-05-16 10:44
2025-05-15 22:46:08 | 20.41 | 52.44 | 45793.4487 | 2025-05-16 10:46
2025-05-15 22:48:08 | 20.4 | 52.37 | 45793.4501 | 2025-05-16 10:48
2025-05-15 22:50:08 | 20.37 | 52.51 | 45793.4515 | 2025-05-16 10:50
2025-05-15 22:52:09 | 20.38 | 52.58 | 45793.4529 | 2025-05-16 10:52
2025-05-15 22:54:09 | 20.4 | 52.61 | 45793.4543 | 2025-05-16 10:54
2025-05-15 22:56:09 | 20.4 | 52.55 | 45793.4557 | 2025-05-16 10:56
2025-05-15 22:58:08 | 20.42 | 52.59 | 45793.457 | 2025-05-16 10:58
2025-05-15 23:00:09 | 20.44 | 52.56 | 45793.4584 | 2025-05-16 11:00
2025-05-15 23:02:08 | 20.4 | 52.47 | 45793.4598 | 2025-05-16 11:02
2025-05-15 23:04:08 | 20.38 | 52.51 | 45793.4612 | 2025-05-16 11:04
2025-05-15 23:06:08 | 20.4 | 52.59 | 45793.4626 | 2025-05-16 11:06
2025-05-15 23:08:09 | 20.41 | 52.6 | 45793.464 | 2025-05-16 11:08
2025-05-15 23:10:08 | 20.4 | 52.57 | 45793.4654 | 2025-05-16 11:10
2025-05-15 23:12:08 | 20.4 | 52.5 | 45793.4668 | 2025-05-16 11:12
2025-05-15 23:14:08 | 20.37 | 52.5 | 45793.4681 | 2025-05-16 11:14
2025-05-15 23:16:09 | 20.37 | 52.57 | 45793.4695 | 2025-05-16 11:16
2025-05-15 23:18:08 | 20.38 | 52.64 | 45793.4709 | 2025-05-16 11:18
2025-05-15 23:20:08 | 20.39 | 52.66 | 45793.4723 | 2025-05-16 11:20
2025-05-15 23:22:08 | 20.41 | 52.71 | 45793.4737 | 2025-05-16 11:22
2025-05-15 23:24:09 | 20.41 | 52.63 | 45793.4751 | 2025-05-16 11:24
2025-05-15 23:26:08 | 20.42 | 52.58 | 45793.4765 | 2025-05-16 11:26
2025-05-15 23:28:09 | 20.41 | 52.49 | 45793.4779 | 2025-05-16 11:28
2025-05-15 23:30:08 | 20.37 | 52.47 | 45793.4793 | 2025-05-16 11:30
2025-05-15 23:32:08 | 20.37 | 52.49 | 45793.4806 | 2025-05-16 11:32
2025-05-15 23:34:08 | 20.39 | 52.49 | 45793.482 | 2025-05-16 11:34
2025-05-15 23:36:09 | 20.39 | 52.48 | 45793.4834 | 2025-05-16 11:36
2025-05-15 23:38:08 | 20.4 | 52.46 | 45793.4848 | 2025-05-16 11:38
2025-05-15 23:40:09 | 20.4 | 52.4 | 45793.4862 | 2025-05-16 11:40
2025-05-15 23:42:09 | 20.4 | 52.35 | 45793.4876 | 2025-05-16 11:42
2025-05-15 23:44:09 | 20.39 | 52.31 | 45793.489 | 2025-05-16 11:44
2025-05-15 23:46:08 | 20.36 | 52.43 | 45793.4904 | 2025-05-16 11:46
2025-05-15 23:48:09 | 20.39 | 52.49 | 45793.4918 | 2025-05-16 11:48
2025-05-15 23:50:08 | 20.41 | 52.53 | 45793.4931 | 2025-05-16 11:50
2025-05-15 23:52:09 | 20.41 | 52.54 | 45793.4945 | 2025-05-16 11:52
2025-05-15 23:54:08 | 20.44 | 52.52 | 45793.4959 | 2025-05-16 11:54
2025-05-15 23:56:09 | 20.42 | 52.45 | 45793.4973 | 2025-05-16 11:56
2025-05-15 23:58:08 | 20.39 | 52.48 | 45793.4987 | 2025-05-16 11:58
2025-05-16 00:00:09 | 20.38 | 52.48 | 45793.5001 | 2025-05-16 12:00
2025-05-16 00:02:08 | 20.38 | 52.59 | 45793.5015 | 2025-05-16 12:02
2025-05-16 00:04:08 | 20.4 | 52.64 | 45793.5029 | 2025-05-16 12:04
2025-05-16 00:06:08 | 20.38 | 52.73 | 45793.5043 | 2025-05-16 12:06
2025-05-16 00:08:09 | 20.42 | 52.71 | 45793.5057 | 2025-05-16 12:08
2025-05-16 00:10:08 | 20.4 | 52.71 | 45793.507 | 2025-05-16 12:10
2025-05-16 00:12:09 | 20.38 | 52.72 | 45793.5084 | 2025-05-16 12:12
2025-05-16 00:14:09 | 20.36 | 52.71 | 45793.5098 | 2025-05-16 12:14
2025-05-16 00:16:09 | 20.37 | 52.72 | 45793.5112 | 2025-05-16 12:16
2025-05-16 00:18:08 | 20.37 | 52.82 | 45793.5126 | 2025-05-16 12:18
2025-05-16 00:20:09 | 20.39 | 52.93 | 45793.514 | 2025-05-16 12:20
2025-05-16 00:22:09 | 20.38 | 52.97 | 45793.5154 | 2025-05-16 12:22
2025-05-16 00:24:09 | 20.4 | 52.92 | 45793.5168 | 2025-05-16 12:24
2025-05-16 00:26:08 | 20.4 | 52.86 | 45793.5181 | 2025-05-16 12:26
2025-05-16 00:28:09 | 20.39 | 52.85 | 45793.5195 | 2025-05-16 12:28
2025-05-16 00:30:08 | 20.36 | 52.81 | 45793.5209 | 2025-05-16 12:30
2025-05-16 00:32:08 | 20.35 | 52.84 | 45793.5223 | 2025-05-16 12:32
2025-05-16 00:34:08 | 20.35 | 52.94 | 45793.5237 | 2025-05-16 12:34
2025-05-16 00:36:09 | 20.35 | 53 | 45793.5251 | 2025-05-16 12:36
2025-05-16 00:38:08 | 20.37 | 53.05 | 45793.5265 | 2025-05-16 12:38
2025-05-16 00:40:08 | 20.4 | 52.98 | 45793.5279 | 2025-05-16 12:40
2025-05-16 00:42:08 | 20.38 | 52.91 | 45793.5293 | 2025-05-16 12:42
2025-05-16 00:44:09 | 20.36 | 53.04 | 45793.5307 | 2025-05-16 12:44
2025-05-16 00:46:08 | 20.37 | 53.1 | 45793.532 | 2025-05-16 12:46
2025-05-16 00:48:09 | 20.38 | 53.07 | 45793.5334 | 2025-05-16 12:48
2025-05-16 00:50:09 | 20.37 | 53.04 | 45793.5348 | 2025-05-16 12:50
2025-05-16 00:52:09 | 20.37 | 53.01 | 45793.5362 | 2025-05-16 12:52
2025-05-16 00:54:08 | 20.36 | 52.98 | 45793.5376 | 2025-05-16 12:54
2025-05-16 00:56:09 | 20.37 | 53.05 | 45793.539 | 2025-05-16 12:56
2025-05-16 00:58:08 | 20.38 | 53.12 | 45793.5404 | 2025-05-16 12:58
2025-05-16 01:00:08 | 20.38 | 53.17 | 45793.5418 | 2025-05-16 13:00
2025-05-16 01:02:08 | 20.39 | 53.19 | 45793.5431 | 2025-05-16 13:02
2025-05-16 01:04:09 | 20.39 | 53.15 | 45793.5445 | 2025-05-16 13:04
2025-05-16 01:06:08 | 20.38 | 53.16 | 45793.5459 | 2025-05-16 13:06
2025-05-16 01:08:09 | 20.38 | 53.14 | 45793.5473 | 2025-05-16 13:08
2025-05-16 01:10:08 | 20.36 | 53.12 | 45793.5487 | 2025-05-16 13:10
2025-05-16 01:12:09 | 20.36 | 53.23 | 45793.5501 | 2025-05-16 13:12
2025-05-16 01:14:09 | 20.37 | 53.29 | 45793.5515 | 2025-05-16 13:14
2025-05-16 01:16:09 | 20.38 | 53.36 | 45793.5529 | 2025-05-16 13:16
2025-05-16 01:18:08 | 20.38 | 53.39 | 45793.5543 | 2025-05-16 13:18
2025-05-16 01:20:09 | 20.4 | 53.34 | 45793.5557 | 2025-05-16 13:20
2025-05-16 01:22:08 | 20.38 | 53.32 | 45793.557 | 2025-05-16 13:22
2025-05-16 01:24:08 | 20.38 | 53.38 | 45793.5584 | 2025-05-16 13:24
2025-05-16 01:26:08 | 20.4 | 53.33 | 45793.5598 | 2025-05-16 13:26
2025-05-16 01:28:09 | 20.38 | 53.46 | 45793.5612 | 2025-05-16 13:28
2025-05-16 01:30:08 | 20.41 | 53.6 | 45793.5626 | 2025-05-16 13:30
2025-05-16 01:32:09 | 20.45 | 53.6 | 45793.564 | 2025-05-16 13:32
2025-05-16 01:34:09 | 20.47 | 53.55 | 45793.5654 | 2025-05-16 13:34
2025-05-16 01:36:09 | 20.45 | 53.5 | 45793.5668 | 2025-05-16 13:36
2025-05-16 01:38:08 | 20.44 | 53.49 | 45793.5681 | 2025-05-16 13:38
2025-05-16 01:40:09 | 20.45 | 53.5 | 45793.5695 | 2025-05-16 13:40
2025-05-16 01:42:08 | 20.4 | 53.44 | 45793.5709 | 2025-05-16 13:42
2025-05-16 01:44:08 | 20.37 | 53.57 | 45793.5723 | 2025-05-16 13:44
2025-05-16 01:46:08 | 20.39 | 53.67 | 45793.5737 | 2025-05-16 13:46
2025-05-16 01:48:09 | 20.4 | 53.67 | 45793.5751 | 2025-05-16 13:48
2025-05-16 01:50:08 | 20.4 | 53.6 | 45793.5765 | 2025-05-16 13:50
2025-05-16 01:52:08 | 20.39 | 53.57 | 45793.5779 | 2025-05-16 13:52
2025-05-16 01:54:08 | 20.37 | 53.57 | 45793.5793 | 2025-05-16 13:54
2025-05-16 01:56:09 | 20.41 | 53.56 | 45793.5807 | 2025-05-16 13:56
2025-05-16 01:58:08 | 20.39 | 53.65 | 45793.582 | 2025-05-16 13:58
2025-05-16 02:00:08 | 20.39 | 53.75 | 45793.5834 | 2025-05-16 14:00
2025-05-16 02:02:08 | 20.42 | 53.82 | 45793.5848 | 2025-05-16 14:02
2025-05-16 02:04:09 | 20.43 | 53.77 | 45793.5862 | 2025-05-16 14:04
2025-05-16 02:06:09 | 20.42 | 53.77 | 45793.5876 | 2025-05-16 14:06
2025-05-16 02:08:09 | 20.4 | 53.76 | 45793.589 | 2025-05-16 14:08
2025-05-16 02:10:08 | 20.38 | 53.74 | 45793.5904 | 2025-05-16 14:10
2025-05-16 02:12:09 | 20.35 | 53.8 | 45793.5918 | 2025-05-16 14:12
2025-05-16 02:14:08 | 20.35 | 53.94 | 45793.5931 | 2025-05-16 14:14
2025-05-16 02:16:08 | 20.37 | 54.03 | 45793.5945 | 2025-05-16 14:16
2025-05-16 02:18:08 | 20.4 | 54.08 | 45793.5959 | 2025-05-16 14:18
2025-05-16 02:20:09 | 20.41 | 54.04 | 45793.5973 | 2025-05-16 14:20
2025-05-16 02:22:08 | 20.39 | 54.05 | 45793.5987 | 2025-05-16 14:22
2025-05-16 02:24:09 | 20.38 | 54.06 | 45793.6001 | 2025-05-16 14:24
2025-05-16 02:26:09 | 20.37 | 54.02 | 45793.6015 | 2025-05-16 14:26
2025-05-16 02:28:09 | 20.36 | 53.99 | 45793.6029 | 2025-05-16 14:28
2025-05-16 02:30:08 | 20.35 | 54.07 | 45793.6043 | 2025-05-16 14:30
2025-05-16 02:32:09 | 20.35 | 54.2 | 45793.6057 | 2025-05-16 14:32
2025-05-16 02:34:08 | 20.34 | 54.26 | 45793.607 | 2025-05-16 14:34
2025-05-16 02:36:08 | 20.37 | 54.3 | 45793.6084 | 2025-05-16 14:36
2025-05-16 02:38:08 | 20.38 | 54.27 | 45793.6098 | 2025-05-16 14:38
2025-05-16 02:40:09 | 20.38 | 54.22 | 45793.6112 | 2025-05-16 14:40
2025-05-16 02:42:08 | 20.38 | 54.15 | 45793.6126 | 2025-05-16 14:42
2025-05-16 02:44:08 | 20.37 | 54.1 | 45793.614 | 2025-05-16 14:44
2025-05-16 02:46:09 | 20.34 | 54.05 | 45793.6154 | 2025-05-16 14:46
2025-05-16 02:48:09 | 20.34 | 53.98 | 45793.6168 | 2025-05-16 14:48
2025-05-16 02:50:08 | 20.35 | 54.06 | 45793.6181 | 2025-05-16 14:50
2025-05-16 02:52:09 | 20.38 | 54.22 | 45793.6195 | 2025-05-16 14:52
2025-05-16 02:54:09 | 20.4 | 54.1 | 45793.6209 | 2025-05-16 14:54
2025-05-16 02:56:09 | 20.4 | 54.05 | 45793.6223 | 2025-05-16 14:56
2025-05-16 02:58:08 | 20.43 | 53.94 | 45793.6237 | 2025-05-16 14:58
2025-05-16 03:00:09 | 20.46 | 53.9 | 45793.6251 | 2025-05-16 15:00
2025-05-16 03:02:09 | 20.48 | 53.86 | 45793.6265 | 2025-05-16 15:02
2025-05-16 03:04:09 | 20.47 | 53.78 | 45793.6279 | 2025-05-16 15:04
2025-05-16 03:06:08 | 20.49 | 53.65 | 45793.6293 | 2025-05-16 15:06
2025-05-16 03:08:09 | 20.44 | 53.7 | 45793.6307 | 2025-05-16 15:08
2025-05-16 03:10:08 | 20.4 | 53.76 | 45793.632 | 2025-05-16 15:10
2025-05-16 03:12:09 | 20.44 | 53.73 | 45793.6334 | 2025-05-16 15:12
2025-05-16 03:14:08 | 20.44 | 53.8 | 45793.6348 | 2025-05-16 15:14
2025-05-16 03:16:09 | 20.42 | 53.9 | 45793.6362 | 2025-05-16 15:16
2025-05-16 03:18:08 | 20.42 | 53.95 | 45793.6376 | 2025-05-16 15:18
2025-05-16 03:20:08 | 20.42 | 53.88 | 45793.639 | 2025-05-16 15:20
2025-05-16 03:22:08 | 20.4 | 53.88 | 45793.6404 | 2025-05-16 15:22
2025-05-16 03:24:09 | 20.38 | 53.87 | 45793.6418 | 2025-05-16 15:24
2025-05-16 03:26:08 | 20.35 | 53.83 | 45793.6431 | 2025-05-16 15:26
2025-05-16 03:28:09 | 20.36 | 53.92 | 45793.6445 | 2025-05-16 15:28
2025-05-16 03:30:08 | 20.35 | 53.97 | 45793.6459 | 2025-05-16 15:30
2025-05-16 03:32:09 | 20.38 | 53.99 | 45793.6473 | 2025-05-16 15:32
2025-05-16 03:34:08 | 20.39 | 53.95 | 45793.6487 | 2025-05-16 15:34
2025-05-16 03:36:09 | 20.39 | 53.87 | 45793.6501 | 2025-05-16 15:36
2025-05-16 03:38:08 | 20.37 | 53.89 | 45793.6515 | 2025-05-16 15:38
2025-05-16 03:40:08 | 20.37 | 53.99 | 45793.6529 | 2025-05-16 15:40
2025-05-16 03:42:08 | 20.38 | 53.99 | 45793.6543 | 2025-05-16 15:42
2025-05-16 03:44:09 | 20.37 | 53.91 | 45793.6557 | 2025-05-16 15:44
2025-05-16 03:46:08 | 20.37 | 53.87 | 45793.657 | 2025-05-16 15:46
2025-05-16 03:48:08 | 20.36 | 53.83 | 45793.6584 | 2025-05-16 15:48
2025-05-16 03:50:08 | 20.38 | 53.86 | 45793.6598 | 2025-05-16 15:50
2025-05-16 03:52:09 | 20.37 | 53.94 | 45793.6612 | 2025-05-16 15:52
2025-05-16 03:54:08 | 20.41 | 53.92 | 45793.6626 | 2025-05-16 15:54
2025-05-16 03:56:09 | 20.46 | 53.9 | 45793.664 | 2025-05-16 15:56
2025-05-16 03:58:09 | 20.5 | 53.75 | 45793.6654 | 2025-05-16 15:58
2025-05-16 04:00:09 | 20.52 | 53.67 | 45793.6668 | 2025-05-16 16:00
2025-05-16 04:02:08 | 20.51 | 53.63 | 45793.6681 | 2025-05-16 16:02
2025-05-16 04:04:09 | 20.5 | 53.59 | 45793.6695 | 2025-05-16 16:04
2025-05-16 04:06:09 | 20.51 | 53.49 | 45793.6709 | 2025-05-16 16:06
2025-05-16 04:08:09 | 20.42 | 53.58 | 45793.6723 | 2025-05-16 16:08
2025-05-16 04:10:08 | 20.43 | 53.72 | 45793.6737 | 2025-05-16 16:10
2025-05-16 04:12:09 | 20.45 | 53.77 | 45793.6751 | 2025-05-16 16:12
2025-05-16 04:14:08 | 20.45 | 53.74 | 45793.6765 | 2025-05-16 16:14
2025-05-16 04:16:08 | 20.41 | 53.7 | 45793.6779 | 2025-05-16 16:16
2025-05-16 04:18:08 | 20.39 | 53.69 | 45793.6793 | 2025-05-16 16:18
2025-05-16 04:20:09 | 20.38 | 53.7 | 45793.6807 | 2025-05-16 16:20
2025-05-16 04:22:08 | 20.38 | 53.8 | 45793.682 | 2025-05-16 16:22
2025-05-16 04:24:09 | 20.39 | 53.87 | 45793.6834 | 2025-05-16 16:24
2025-05-16 04:26:08 | 20.41 | 53.92 | 45793.6848 | 2025-05-16 16:26
2025-05-16 04:28:09 | 20.41 | 53.9 | 45793.6862 | 2025-05-16 16:28
2025-05-16 04:30:08 | 20.39 | 53.92 | 45793.6876 | 2025-05-16 16:30
2025-05-16 04:32:09 | 20.39 | 53.87 | 45793.689 | 2025-05-16 16:32
2025-05-16 04:34:08 | 20.38 | 53.84 | 45793.6904 | 2025-05-16 16:34
2025-05-16 04:36:08 | 20.37 | 53.87 | 45793.6918 | 2025-05-16 16:36
2025-05-16 04:38:08 | 20.36 | 54 | 45793.6931 | 2025-05-16 16:38
2025-05-16 04:40:09 | 20.37 | 54.04 | 45793.6945 | 2025-05-16 16:40
2025-05-16 04:42:08 | 20.39 | 54.07 | 45793.6959 | 2025-05-16 16:42
2025-05-16 04:44:08 | 20.4 | 54.04 | 45793.6973 | 2025-05-16 16:44
2025-05-16 04:46:09 | 20.37 | 53.98 | 45793.6987 | 2025-05-16 16:46
2025-05-16 04:48:09 | 20.39 | 53.94 | 45793.7001 | 2025-05-16 16:48
2025-05-16 04:50:08 | 20.36 | 53.89 | 45793.7015 | 2025-05-16 16:50
2025-05-16 04:52:09 | 20.36 | 53.91 | 45793.7029 | 2025-05-16 16:52
2025-05-16 04:54:08 | 20.35 | 54.01 | 45793.7043 | 2025-05-16 16:54
2025-05-16 04:56:08 | 20.39 | 54.05 | 45793.7056 | 2025-05-16 16:56
2025-05-16 04:58:08 | 20.39 | 54.03 | 45793.707 | 2025-05-16 16:58
2025-05-16 05:00:09 | 20.39 | 54 | 45793.7084 | 2025-05-16 17:00
2025-05-16 05:02:08 | 20.38 | 53.94 | 45793.7098 | 2025-05-16 17:02
2025-05-16 05:04:09 | 20.37 | 53.88 | 45793.7112 | 2025-05-16 17:04
2025-05-16 05:06:08 | 20.35 | 53.82 | 45793.7126 | 2025-05-16 17:06
2025-05-16 05:08:09 | 20.34 | 53.91 | 45793.714 | 2025-05-16 17:08
2025-05-16 05:10:08 | 20.35 | 54 | 45793.7154 | 2025-05-16 17:10
2025-05-16 05:12:09 | 20.36 | 54.01 | 45793.7168 | 2025-05-16 17:12
2025-05-16 05:14:09 | 20.38 | 53.97 | 45793.7182 | 2025-05-16 17:14
2025-05-16 05:16:09 | 20.38 | 53.9 | 45793.7195 | 2025-05-16 17:16
2025-05-16 05:18:08 | 20.38 | 53.86 | 45793.7209 | 2025-05-16 17:18
2025-05-16 05:20:09 | 20.36 | 53.81 | 45793.7223 | 2025-05-16 17:20
2025-05-16 05:22:09 | 20.34 | 53.74 | 45793.7237 | 2025-05-16 17:22
2025-05-16 05:24:09 | 20.34 | 53.8 | 45793.7251 | 2025-05-16 17:24
2025-05-16 05:26:08 | 20.35 | 53.86 | 45793.7265 | 2025-05-16 17:26
2025-05-16 05:28:09 | 20.37 | 53.87 | 45793.7279 | 2025-05-16 17:28
2025-05-16 05:30:08 | 20.37 | 53.85 | 45793.7293 | 2025-05-16 17:30
2025-05-16 05:32:08 | 20.38 | 53.79 | 45793.7306 | 2025-05-16 17:32
2025-05-16 05:34:08 | 20.36 | 53.73 | 45793.732 | 2025-05-16 17:34
2025-05-16 05:36:09 | 20.36 | 53.67 | 45793.7334 | 2025-05-16 17:36
2025-05-16 05:38:08 | 20.35 | 53.64 | 45793.7348 | 2025-05-16 17:38
2025-05-16 05:40:09 | 20.33 | 53.7 | 45793.7362 | 2025-05-16 17:40
2025-05-16 05:42:09 | 20.36 | 53.74 | 45793.7376 | 2025-05-16 17:42
2025-05-16 05:44:09 | 20.38 | 53.74 | 45793.739 | 2025-05-16 17:44
2025-05-16 05:46:08 | 20.38 | 53.74 | 45793.7404 | 2025-05-16 17:46
2025-05-16 05:48:09 | 20.38 | 53.69 | 45793.7418 | 2025-05-16 17:48
2025-05-16 05:50:09 | 20.38 | 53.61 | 45793.7432 | 2025-05-16 17:50
2025-05-16 05:52:09 | 20.35 | 53.6 | 45793.7445 | 2025-05-16 17:52
2025-05-16 05:54:08 | 20.34 | 53.6 | 45793.7459 | 2025-05-16 17:54
2025-05-16 05:56:09 | 20.35 | 53.65 | 45793.7473 | 2025-05-16 17:56
2025-05-16 05:58:08 | 20.34 | 53.71 | 45793.7487 | 2025-05-16 17:58

## Historical
| Mea Date | 1995-09-29 | S No. | S21069 |  | 2000-08-11 | S21384 |  |  | 2005-02-25 | S21586 |  |  | 2009-07-03 | S21740 |  |  | 2014-10-09 | S21952 |  |  | 2020-11-11 | S22086 |  |  | 2025-05-15 | S22176
|  | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR
| A1 | 994.03 | 0.054572 | 5.7182 | 0.024721 | 994.07 | 0.054568 | 5.6995 | 0.02464 | 993.9916 | 0.11571 | 5.6916 | 0.024611 | 994.0646 | 0.065681 | 5.6775 | 0.024545 | 994.1995 | 0.070355 | 5.6915 | 0.024639 | 994.017 | 0.063 | 5.686 | 0.034 | 993.8652 | 0.063258 | 5.7358 | 0.033902
| A2 | 1980.3588 | 0.10366 | 8.4219 | 0.036409 | 1990.1204 | 0.10418 | 8.4939 | 0.03672 | 1980.4119 | 5.8869 | 8.3817 | 0.03664 | 1980.5079 | 0.12156 | 8.3615 | 0.036147 | 1980.7775 | 0.12681 | 8.3825 | 0.03629 | 1980.52 | 0.1 | 8.374 | 0.05 | 1980.1783 | 0.098608 | 8.4486 | 0.04986
| A3 | 2968.9979 | 0.15407 | 10.7121 | 0.04631 | 2973.7887 | 0.15432 | 10.7324 | 0.046397 | 2968.8226 | 2.0202 | 10.661 | 0.046165 | 2969.2816 | 0.17973 | 10.6351 | 0.045976 | 2969.7059 | 0.18622 | 10.6619 | 0.04616 | 2969.42 | 0.14 | 10.653 | 0.063 | 2968.8291 | 0.13834 | 10.7328 | 0.063326
| A4 | 3948.4356 | 0.20423 | 12.7063 | 0.054931 | 3948.9357 | 0.20425 | 12.6663 | 0.054758 | 3948.7118 | 0.179 | 12.6462 | 0.054673 | 3948.8537 | 0.23785 | 12.6152 | 0.054536 | 3949.4225 | 0.24566 | 12.647 | 0.054756 | 3949.15 | 0.18 | 12.638 | 0.075 | 3948.3254 | 0.1791 | 12.7391 | 0.075136
| A5 | 4931.9739 | 0.25468 | 14.4818 | 0.062607 | 4932.724 | 0.25472 | 14.4375 | 0.062415 | 4932.4416 | 0.21648 | 14.4139 | 0.062315 | 4932.5965 | 0.29641 | 14.3784 | 0.062158 | 4933.3499 | 0.30557 | 14.415 | 0.062413 | 4933.13 | 0.22 | 14.408 | 0.085 | 4932.0819 | 0.22055 | 14.5255 | 0.085658
| A6 | 5914.202 | 0.30511 | 16.1744 | 0.069924 | 5915.1822 | 0.30515 | 16.1228 | 0.069701 | 5914.8613 | 0.25412 | 16.0991 | 0.069599 | 5915.0491 | 0.35499 | 16.0588 | 0.069423 | 5915.9671 | 0.36551 | 16.1005 | 0.069713 | 5915.88 | 0.26 | 16.09 | 0.1 | 5914.5845 | 0.26221 | 16.2179 | 0.095651
| A7 | 6890.9693 | 0.35527 | 17.8224 | 0.077048 | 6892.0394 | 0.35532 | 17.7658 | 0.076804 | 6891.6899 | 0.2924 | 17.7417 | 0.0767 | 6891.8207 | 0.41329 | 17.6964 | 0.076503 | 6892.8931 | 0.42519 | 17.7431 | 0.076827 | 6892.98 | 0.3 | 17.74 | 0.1 | 6891.3862 | 0.30386 | 17.8708 | 0.1054
| A8 | 7863.1258 | 0.40521 | 19.3421 | 0.083619 | 7864.406 | 0.40527 | 19.2813 | 0.083356 | 7864.0478 | 0.33083 | 19.2569 | 0.08325 | 7864.1515 | 0.47135 | 19.2063 | 0.08303 | 7865.3283 | 0.48461 | 19.2576 | 0.083388 | 7865.64 | 0.35 | 19.25 | 0.11 | 7863.7411 | 0.34542 | 19.4016 | 0.11442
| A9 | 8832.1617 | 0.455 | 20.8018 | 0.089929 | 8833.702 | 0.45507 | 20.7375 | 0.089651 | 8833.3752 | 0.36924 | 20.7133 | 0.089546 | 8833.2017 | 0.52924 | 20.6573 | 0.089303 | 8834.5327 | 0.54385 | 20.7131 | 0.089693 | 8835.46 | 0.39 | 20.71 | 0.12 | 8833.1542 | 0.38692 | 20.8733 | 0.12309
| A10 | 9802.568 | 0.50486 | 22.2832 | 0.096333 | 9804.3182 | 0.50495 | 22.2153 | 0.096039 | 9803.9928 | 0.40859 | 22.196 | 0.095955 | 9803.6322 | 0.58722 | 22.1296 | 0.095668 | 9805.0675 | 0.60321 | 22.19 | 0.09609 | 9806.59 | 0.43 | 22.22 | 0.13 | 9803.7532 | 0.42854 | 22.3632 | 0.13187
| B1 | 99.162 | 0.0057279 | 1.1706 | 0.0050607 | 99.158 | 0.0057257 | 1.1675 | 0.0050472 | 99.1611 | 0.0061381 | 1.1672 | 0.0050456 | 99.1564 | 0.006623 | 1.1635 | 0.0050299 | 99.1589 | 0.0076874 | 1.1661 | 0.0050478 | 99.1417 | 0.0072 | 1.166 | 0.007 | 99.1331 | 0.007219 | 1.172 | 0.0069133
| B2 | 199.028 | 0.010709 | 1.7457 | 0.0075468 | 199.028 | 0.010706 | 1.741 | 0.0075267 | 199.0264 | 0.0099933 | 1.7402 | 0.007523 | 199.029 | 0.012292 | 1.7351 | 0.0075008 | 199.0329 | 0.013529 | 1.7389 | 0.0075276 | 199.001 | 0.011 | 1.738 | 0.01 | 198.9824 | 0.011174 | 1.7485 | 0.010309
| B3 | 298.294 | 0.015774 | 2.2147 | 0.0095743 | 298.278 | 0.015772 | 2.2088 | 0.0095488 | 298.2935 | 0.01444 | 2.2078 | 0.0095441 | 298.2944 | 0.018139 | 2.2012 | 0.0095158 | 298.3009 | 0.019545 | 2.2061 | 0.0095498 | 298.249 | 0.015 | 2.205 | 0.013 | 298.2047 | 0.015342 | 2.2202 | 0.013089
| B4 | 396.5528 | 0.020872 | 2.8022 | 0.012114 | 396.5508 | 0.020871 | 2.7977 | 0.012095 | 396.5635 | 0.018253 | 2.7933 | 0.012075 | 396.5757 | 0.023993 | 2.7851 | 0.01204 | 396.5866 | 0.025714 | 2.7912 | 0.012083 | 396.524 | 0.02 | 2.79 | 0.016 | 396.4813 | 0.019829 | 2.8113 | 0.016574
| B5 | 494.4916 | 0.025901 | 3.1871 | 0.013778 | 494.4936 | 0.025899 | 3.1818 | 0.013755 | 494.5144 | 0.022373 | 3.1765 | 0.013732 | 494.5359 | 0.029828 | 3.1676 | 0.013693 | 494.5412 | 0.031708 | 3.1744 | 0.013742 | 494.473 | 0.364 | 3.173 | 0.019 | 494.4313 | 0.024043 | 3.199 | 0.018856
| B6 | 593.5215 | 0.030991 | 3.5497 | 0.015346 | 593.5255 | 0.030989 | 3.5435 | 0.015319 | 594.1446 | 0.12518 | 3.5335 | 0.015276 | 593.6203 | 0.03574 | 3.5304 | 0.015262 | 593.5961 | 0.037781 | 3.5375 | 0.015314 | 593.512 | 0.028 | 3.535 | 0.021 | 594.8515 | 0.028355 | 3.5531 | 0.020946
| B7 | 691.3553 | 0.036026 | 3.9065 | 0.016888 | 691.3543 | 0.036025 | 3.8996 | 0.016858 | 691.3874 | 0.030514 | 3.8931 | 0.01683 | 691.4345 | 0.041582 | 3.8823 | 0.016783 | 691.4347 | 0.043794 | 3.8901 | 0.01684 | 691.343 | 0.033 | 3.893 | 0.023 | 691.3009 | 0.032573 | 3.9213 | 0.023114
| B8 | 789.146 | 0.041052 | 4.2106 | 0.018203 | 789.142 | 0.041049 | 4.2027 | 0.018169 | 789.1783 | 0.034488 | 4.1957 | 0.018138 | 789.2346 | 0.047424 | 4.1843 | 0.018089 | 789.2294 | 0.04978 | 4.1925 | 0.01815 | 789.131 | 0.037 | 4.197 | 0.025 | 789.0778 | 0.036779 | 4.2255 | 0.024909
| B9 | 887.3628 | 0.046106 | 4.5293 | 0.019581 | 887.3698 | 0.046104 | 4.5205 | 0.019543 | 887.3962 | 0.038437 | 4.513 | 0.01951 | 887.4889 | 0.053297 | 4.5009 | 0.019457 | 887.4941 | 0.055816 | 4.5095 | 0.019522 | 887.363 | 0.041 | 4.515 | 0.027 | 887.3028 | 0.041047 | 4.5464 | 0.026799
| B10 | 984.7705 | 0.05112 | 4.8376 | 0.020913 | 984.7775 | 0.051118 | 4.8279 | 0.020872 | 984.818 | 0.042525 | 4.8219 | 0.020845 | 984.942 | 0.059123 | 4.8072 | 0.020781 | 984.9506 | 0.061804 | 4.8163 | 0.02085 | 984.794 | 0.045 | 4.823 | 0.029 | 984.7797 | 0.045246 | 4.8226 | 0.028508
Note
2025-05-28 |  | L: ##; uL: No report for S21069 found; R: ##; uR: All figures from report S2****
b10
|  | L | uL | R | uR
| 1995-09-29 | L: 984.7705; uL: 0.05112; R: 4.8376; uR: 0.020913
| 2000-08-11 | L: 984.7775; uL: 0.051118; R: 4.8279; uR: 0.020872
| 2005-02-25 | L: 984.818; uL: 0.042525; R: 4.8219; uR: 0.020845
| 2009-07-03 | L: 984.942; uL: 0.059123; R: 4.8072; uR: 0.020781
| 2014-10-09 | L: 984.9506; uL: 0.061804; R: 4.8163; uR: 0.02085
| 2020-11-11 | L: 984.794; uL: 0.045; R: 4.823; uR: 0.029

## Report
checked by: |  |  | humidity | 43.0735 | %RH
|  |  | temperature | 19.485 |  |  |  |  | 2009-07-03 | [6 empty] | 2014-10-09 | [13 empty] | 2025-05-16
[6 empty] | temperature | 19.485 |  | S21740 | 2009 | [12 empty] | 09-10/11/2020
| dial code | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor
|  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  | Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor |  |  | uncertainty | factor |  | uncertainty | factor
| A1 | 1 | 994.1995 | 0.070355 | 2.069 | 5.6915 | 0.024639 | 2.1612 | 1 | 994.0646 | 0.065681 | 2.0503 | 5.6775 | 0.024545 | 2.1614 | 1 | 994.0646 | 0.065681 | 2.0503 | 5.6775 | 0.024545 | 2.1614 | 1 | 994.017 | 0.063 | 2 | 5.686 | 0.034 | 2.1 | 1 | (mH) | (mH) |  | (W) | (W)
| A2 | 2 | 1980.7775 | 0.12681 | 2.0919 | 8.3825 | 0.03629 | 2.1612 | 2 | 1980.5079 | 0.12156 | 2.065 | 8.3615 | 0.036147 | 2.1615 | 2 | 1980.5079 | 0.12156 | 2.065 | 8.3615 | 0.036147 | 2.1615 | 2 | 1980.52 | 0.1 | 2 | 8.374 | 0.05 | 2.1 | 2 | 993.8652 | 0.063258 | 1.9889 | 5.7358 | 0.033902 | 2.0825
| A3 | 3 | 2969.7059 | 0.18622 | 2.105 | 10.6619 | 0.04616 | 2.1611 | 3 | 2969.2816 | 0.17973 | 2.0714 | 10.6351 | 0.045976 | 2.1615 | 3 | 2969.2816 | 0.17973 | 2.0714 | 10.6351 | 0.045976 | 2.1615 | 3 | 2969.42 | 0.14 | 2 | 10.653 | 0.063 | 2.1 | 3 | 1980.1783 | 0.098608 | 2.0002 | 8.4486 | 0.04986 | 2.0836
| A4 | 4 | 3949.4225 | 0.24566 | 2.1114 | 12.647 | 0.054756 | 2.1611 | 4 | 3948.8537 | 0.23785 | 2.0741 | 12.6152 | 0.054536 | 2.1615 | 4 | 3948.8537 | 0.23785 | 2.0741 | 12.6152 | 0.054536 | 2.1615 | 4 | 3949.15 | 0.18 | 2 | 12.638 | 0.075 | 2.1 | 4 | 2968.8291 | 0.13834 | 2.0101 | 10.7328 | 0.063326 | 2.0844
| A5 | 5 | 4933.3499 | 0.30557 | 2.115 | 14.415 | 0.062413 | 2.161 | 5 | 4932.5965 | 0.29641 | 2.0755 | 14.3784 | 0.062158 | 2.1615 | 5 | 4932.5965 | 0.29641 | 2.0755 | 14.3784 | 0.062158 | 2.1615 | 5 | 4933.13 | 0.22 | 2 | 14.408 | 0.085 | 2.1 | 5 | 3948.3254 | 0.1791 | 2.0156 | 12.7391 | 0.075136 | 2.0844
| A6 | 6 | 5915.9671 | 0.36551 | 2.1173 | 16.1005 | 0.069713 | 2.161 | 6 | 5915.0491 | 0.35499 | 2.0763 | 16.0588 | 0.069423 | 2.1615 | 6 | 5915.0491 | 0.35499 | 2.0763 | 16.0588 | 0.069423 | 2.1615 | 6 | 5915.88 | 0.26 | 2 | 16.09 | 0.1 | 2.1 | 6 | 4932.0819 | 0.22055 | 2.0189 | 14.5255 | 0.085658 | 2.0843
| A7 | 7 | 6892.8931 | 0.42519 | 2.1188 | 17.7431 | 0.076827 | 2.1609 | 7 | 6891.8207 | 0.41329 | 2.0767 | 17.6964 | 0.076503 | 2.1615 | 7 | 6891.8207 | 0.41329 | 2.0767 | 17.6964 | 0.076503 | 2.1615 | 7 | 6892.98 | 0.3 | 2 | 17.74 | 0.1 | 2.1 | 7 | 5914.5845 | 0.26221 | 2.0211 | 16.2179 | 0.095651 | 2.0847
| A8 | 8 | 7865.3283 | 0.48461 | 2.12 | 19.2576 | 0.083388 | 2.1609 | 8 | 7864.1515 | 0.47135 | 2.0771 | 19.2063 | 0.08303 | 2.1615 | 8 | 7864.1515 | 0.47135 | 2.0771 | 19.2063 | 0.08303 | 2.1615 | 8 | 7865.64 | 0.35 | 2 | 19.25 | 0.11 | 2.1 | 8 | 6891.3862 | 0.30386 | 2.0225 | 17.8708 | 0.1054 | 2.0848
| A9 | 9 | 8834.5327 | 0.54385 | 2.1208 | 20.7131 | 0.089693 | 2.1609 | 9 | 8833.2017 | 0.52924 | 2.0773 | 20.6573 | 0.089303 | 2.1615 | 9 | 8833.2017 | 0.52924 | 2.0773 | 20.6573 | 0.089303 | 2.1615 | 9 | 8835.46 | 0.39 | 2 | 20.71 | 0.12 | 2.1 | 9 | 7863.7411 | 0.34542 | 2.0236 | 19.4016 | 0.11442 | 2.0847
| A10 | 10 | 9805.0675 | 0.60321 | 2.1214 | 22.19 | 0.09609 | 2.1608 | 10 | 9803.6322 | 0.58722 | 2.0775 | 22.1296 | 0.095668 | 2.1615 | 10 | 9803.6322 | 0.58722 | 2.0775 | 22.1296 | 0.095668 | 2.1615 | 10 | 9806.59 | 0.43 | 2 | 22.22 | 0.13 | 2.1 | 10 | 8833.1542 | 0.38692 | 2.0243 | 20.8733 | 0.12309 | 2.0846
[31 empty] | 9803.7532 | 0.42854 | 2.0249 | 22.3632 | 0.13187 | 2.0846
|  | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | [7 empty] | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor | [9 empty] | uncertainty | factor |  | uncertainty | factor
|  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  | Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor |  | (mH) | (mH) |  | (W) | (W)
| B1 | 1 | 99.1589 | 0.0076874 | 2.042 | 1.1661 | 0.0050478 | 2.1613 | 1 | 99.1564 | 0.006623 | 2.046 | 1.1635 | 0.0050299 | 2.1615 | 1 | 99.1564 | 0.006623 | 2.046 | 1.1635 | 0.0050299 | 2.1615 | 1 | 99.1417 | 0.0072 | 2 | 1.166 | 0.007 | 2.1 | 1 | 99.1331 | 0.007219 | 1.9977 | 1.172 | 0.0069133 | 2.0845
| B2 | 2 | 199.0329 | 0.013529 | 2.063 | 1.7389 | 0.0075276 | 2.1613 | 2 | 199.029 | 0.012292 | 2.0617 | 1.7351 | 0.0075008 | 2.1615 | 2 | 199.029 | 0.012292 | 2.0617 | 1.7351 | 0.0075008 | 2.1615 | 2 | 199.001 | 0.011 | 2 | 1.738 | 0.01 | 2.1 | 2 | 198.9824 | 0.011174 | 2.0028 | 1.7485 | 0.010309 | 2.0847
| B3 | 3 | 298.3009 | 0.019545 | 2.078 | 2.2061 | 0.0095498 | 2.1613 | 3 | 298.2944 | 0.018139 | 2.0688 | 2.2012 | 0.0095158 | 2.1615 | 3 | 298.2944 | 0.018139 | 2.0688 | 2.2012 | 0.0095158 | 2.1615 | 3 | 298.249 | 0.015 | 2 | 2.205 | 0.013 | 2.1 | 3 | 298.2047 | 0.015342 | 2.0065 | 2.2202 | 0.013089 | 2.0848
| B4 | 4 | 396.5866 | 0.025714 | 2.0836 | 2.7912 | 0.012083 | 2.1613 | 4 | 396.5757 | 0.023993 | 2.0715 | 2.7851 | 0.01204 | 2.1615 | 4 | 396.5757 | 0.023993 | 2.0715 | 2.7851 | 0.01204 | 2.1615 | 4 | 396.524 | 0.02 | 2 | 2.79 | 0.016 | 2.1 | 4 | 396.4813 | 0.019829 | 2.0094 | 2.8113 | 0.016574 | 2.0849
| B5 | 5 | 494.5412 | 0.031708 | 2.09 | 3.1744 | 0.013742 | 2.1613 | 5 | 494.5359 | 0.029828 | 2.0732 | 3.1676 | 0.013693 | 2.1615 | 5 | 494.5359 | 0.029828 | 2.0732 | 3.1676 | 0.013693 | 2.1615 | 5 | 494.473 | 0.364 | 2.2 | 3.173 | 0.019 | 2.1 | 5 | 494.4313 | 0.024043 | 2.0109 | 3.199 | 0.018856 | 2.0847
| B6 | 6 | 593.5961 | 0.037781 | 2.0945 | 3.5375 | 0.015314 | 2.1613 | 6 | 593.6203 | 0.03574 | 2.0743 | 3.5304 | 0.015262 | 2.1615 | 6 | 593.6203 | 0.03574 | 2.0743 | 3.5304 | 0.015262 | 2.1615 | 6 | 593.512 | 0.028 | 2 | 3.535 | 0.021 | 2.1 | 6 | 594.8515 | 0.028355 | 2.0123 | 3.5531 | 0.020946 | 2.0849
| B7 | 7 | 691.4347 | 0.043794 | 2.0978 | 3.8901 | 0.01684 | 2.1613 | 7 | 691.4345 | 0.041582 | 2.075 | 3.8823 | 0.016783 | 2.1615 | 7 | 691.4345 | 0.041582 | 2.075 | 3.8823 | 0.016783 | 2.1615 | 7 | 691.343 | 0.033 | 2 | 3.893 | 0.023 | 2.1 | 7 | 691.3009 | 0.032573 | 2.0134 | 3.9213 | 0.023114 | 2.0848
| B8 | 8 | 789.2294 | 0.04978 | 2.1007 | 4.1925 | 0.01815 | 2.1613 | 8 | 789.2346 | 0.047424 | 2.0755 | 4.1843 | 0.018089 | 2.1615 | 8 | 789.2346 | 0.047424 | 2.0755 | 4.1843 | 0.018089 | 2.1615 | 8 | 789.131 | 0.037 | 2 | 4.197 | 0.025 | 2.1 | 8 | 789.0778 | 0.036779 | 2.0144 | 4.2255 | 0.024909 | 2.0849
| B9 | 9 | 887.4941 | 0.055816 | 2.1028 | 4.5095 | 0.019522 | 2.1613 | 9 | 887.4889 | 0.053297 | 2.0758 | 4.5009 | 0.019457 | 2.1615 | 9 | 887.4889 | 0.053297 | 2.0758 | 4.5009 | 0.019457 | 2.1615 | 9 | 887.363 | 0.041 | 2 | 4.515 | 0.027 | 2.1 | 9 | 887.3028 | 0.041047 | 2.0152 | 4.5464 | 0.026799 | 2.0848
| B10 | 10 | 984.9506 | 0.061804 | 2.1045 | 4.8163 | 0.02085 | 2.1612 | 10 | 984.942 | 0.059123 | 2.0761 | 4.8072 | 0.020781 | 2.1615 | 10 | 984.942 | 0.059123 | 2.0761 | 4.8072 | 0.020781 | 2.1615 | 10 | 984.794 | 0.045 | 2 | 4.823 | 0.029 | 2.1 | 10 | 984.7797 | 0.045246 | 2.0161 | 4.8226 | 0.028508 | 2.0855

## Results
Date | 2014-10-09
Client | Babcock (NZ) Limited
Job number | 96040221
Report Number | S21952
For | 1592Hz
Results Using Universal Bridge
Dial | Label | Range | a | b | Inductance |  | Resistance |  | Frequency (kHz) | Temperature | Corrected T
[11 empty] | -0.03 | corrrection for GMH51
A | A0 | 3Z | -1 | -10 | 0 | milliH | 0 | ohm | 1.5921 | 19.43
A | A1 | 3Z | 99399 | 569099 | 0.9942 | milliH | 5.6915 | ohm | 1.5922 | 19.44 | 19.41
A | A2 | 3Z | 198037 | 838177 | 1.9808 | milliH | 8.3825 | ohm | 1.5923 | 19.42 | 19.39
A | A3 | 3Z | 296910 | 1066102 | 2.9697 | milliH | 10.6619 | ohm | 1.5924 | 19.46 | 19.43
A | A4 | 3Z | 394862 | 1264597 | 3.9494 | milliH | 12.647 | ohm | 1.5925 | 19.51 | 19.48
A | A5 | 3Z | 493235 | 1441386 | 4.9333 | milliH | 14.415 | ohm | 1.5926 | 19.52 | 19.49
A | A6 | 3Z | 591477 | 1609922 | 5.916 | milliH | 16.1005 | ohm | 1.5927 | 19.49 | 19.46
A | A7 | 3Z | 689150 | 1774172 | 6.8929 | milliH | 17.7431 | ohm | 1.5928 | 19.46 | 19.43
A | A8 | 3Z | 786374 | 1925612 | 7.8653 | milliH | 19.2576 | ohm | 1.5929 | 19.42 | 19.39
A | A9 | 3Z | 883275 | 2071149 | 8.8345 | milliH | 20.7131 | ohm | 1.5929 | 19.41 | 19.38
A | A10 | 3Z | 980309 | 2218833 | 9.8051 | milliH | 22.19 | ohm | 1.5913 | 19.44 | 19.41
B | B1 | 2Z | 99137 | 1166002 | 99.1589 | microH | 1.1661 | ohm | 1.5917 | 19.67 | 19.64
B | B2 | 2Z | 198991 | 1738805 | 199.0329 | microH | 1.7389 | ohm | 1.5918 | 19.65 | 19.62
B | B3 | 2Z | 298239 | 2205906 | 298.3009 | microH | 2.2061 | ohm | 1.5918 | 19.6 | 19.57
B | B4 | 2Z | 396505 | 2790983 | 396.5866 | microH | 2.7912 | ohm | 1.5919 | 19.56 | 19.53
B | B5 | 2Z | 494440 | 3174124 | 494.5412 | microH | 3.1744 | ohm | 1.592 | 19.5 | 19.47
B | B6 | 2Z | 593475 | 3537225 | 593.5961 | microH | 3.5375 | ohm | 1.5916 | 19.66 | 19.63
B | B7 | 2Z | 691294 | 3889800 | 691.4347 | microH | 3.8901 | ohm | 1.5922 | 19.59 | 19.56
B | B8 | 2Z | 789069 | 4192210 | 789.2294 | microH | 4.1925 | ohm | 1.5923 | 19.58 | 19.55
B | B9 | 2Z | 887314 | 4509200 | 887.4941 | microH | 4.5095 | ohm | 1.5923 | 19.5 | 19.47
B | B10 | 2Z | 984751 | 4815910 | 984.9506 | microH | 4.8163 | ohm | 1.5924 | 19.42 | 19.39
B | B0 | 2Z | -2 | -15 | 0 | microH | 0 | ohm | 1.5917 | 19.69 | 19.66
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
HP4284 check | [5 empty] | zero corrected | L | R | UB/HP (scale factor)
700 | 695.657 | 3.8876 | 695.6523 | 3.8877 |  | 700 | 693.7174 | 3.869 | 0.99671 | 1.0054
| 695.657 | 3.8876 | 0.0044987 | 6.694e-05 |  | 600 | 595.6647 | 3.5172 | 0.99653 | 1.0058
| 695.656 | 3.8877 |  |  |  | 500 | 496.2928 | 3.162 | 0.99647 | 1.0039
| 695.648 | 3.8877
| 695.65 | 3.8877 |  |  |  |  | average of 500 and 700 |  | 0.99659 | 1.0047
| 695.652 | 3.8877 |  |  |  |  | % difference at 600 |  | -0.006318 | 0.10703
| 695.646 | 3.8878
Conclude 600 result is good to at least 0.01% for L
600 | 597.6 | 3.5359 | 597.5997 | 3.5359
| 597.596 | 3.5359 | 0.0029439 | 2.2509e-05 |  |  | Reported U | 0.0061784 | %
| 597.599 | 3.5359
| 597.599 | 3.5359 |  |  |  |  | Conclude that HP4284 is not inconsistent with UB
| 597.605 | 3.5359
| 597.599 | 3.5359
500 | 498.219 | 3.1809 | 498.2277 | 3.1806
| 498.227 | 3.1808 | 0.0043861 | 0.00018154
| 498.226 | 3.1806
| 498.231 | 3.1806
| 498.23 | 3.1805
| 498.229 | 3.1804
| 498.232 | 3.1805
short | 1.9538 | 0.01795 | 1.9349 | 0.01864
| 1.9528 | 0.018 | 0.015132 | 0.00054327
| 1.9476 | 0.01818
| 1.9384 | 0.01852
| 1.9213 | 0.01919
| 1.9178 | 0.01917
| 1.9204 | 0.01917
| 1.9271 | 0.01894

## AH2700A results
from 'FinalDecadeCapacitor.xlsx' for S20885
Start with internal cal.
AH2700 will be calibrated 'on the day' against known reference capacitors.
Use these results to have a correction factor adequate for this calibration.
[7 empty] | subtract zero |  |  |  | Value from capacitance scale, out.csv |  |  | measured | NMIA
| Reference | C | sdC | G | sdG |  | C | sdC | G | sdG | C | mean | stdev | tan delta | G | tan delta
|  | pF | pF | nS | nS |  | pF | pF | nS | nS | pF | -2.483 | 0.13805 |  |  |  | correction
| zero | -1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06 |  | 0 | 1.9233e-07 | 0 | 2.7294e-06 |  | ppm correction |  | mrad | nS | mrad | mrad
| AH111A | 10 | 1.8e-07 | -0.00012269 | 2.05e-06 |  | 10 | 2.256e-07 | -0.00014989 | 2.8156e-06 | 9.9999 | -2.4644 |  | -0.0014989 | 0.00018 | 0.0018 | 0.0032989
| AH111B | 10 | 2.65e-07 | -6.61e-05 | 2.22e-06 |  | 10 | 2.9786e-07 | -9.33e-05 | 2.9416e-06 | 10 | -2.4315 |  | -0.000933 | 0.00024 | 0.0024 | 0.003333
| AH111C | 99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06 |  | 99.9998 | 2.1693e-07 | 0.0016097 | 2.4943e-06 | 99.9996 | -2.3617 |  | 0.0016097 | 0.0019 | 0.0019 | 0.0002903
| AH111D | 99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06 |  | 99.9998 | 2.7412e-07 | 0.0020538 | 2.8595e-06 | 99.9996 | -2.3869 |  | 0.0020538 | 0.0023 | 0.0023 | 0.00024625
| GR1000A | 1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05 |  | 1000.0317 | 6.4014e-06 | 0.17933 | 5.4534e-05 | 1000.0289 | -2.7431 |  | 0.017932
| GR1000B | 1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05 |  | 1000.0164 | 5.9216e-06 | 0.18442 | 5.7932e-05 | 1000.0139 | -2.5101 |  | 0.018442
2020-10-12 | UB1000 | 1000.1471 | 5.89e-06 | 0.142 | 9.76e-05 |  | 1000.1471 | 5.8916e-06 | 0.14197 | 9.7619e-05 | 1000.1444 | -2.6266 |  | 0.014195
2020-10-16 | UB1000 | 1000.1471 | 1.4353e-05 | 0.10825 | 0.0001556 |  | 1000.1471 | 1.4353e-05 | 0.10822 | 0.00015561 | 1000.1444 | -2.6266 |  | 0.01082
2020-10-19 | UB1000 | 1000.1393 | 7.4833e-06 | 0.10425 | 4.8629e-05 |  | 1000.1393 | 7.4846e-06 | 0.10422 | 4.8667e-05 | 1000.1367 | -2.6266 |  | 0.010421
2020-10-20 | UB1000 | 1000.1439 | 7.4162e-06 | 0.13882 | 5.3703e-05 |  | 1000.1439 | 7.4174e-06 | 0.13879 | 5.3738e-05 | 1000.1413 | -2.6266 |  | 0.013877
2020-10-23 | UB1000 | 1000.1511 | 3e-06 | 0.10824 | 5.0326e-05 |  | 1000.1511 | 3.0031e-06 | 0.10821 | 5.0363e-05 | 1000.1485 | -2.6266 |  | 0.010819
2020-10-23 | UB1000 | 1000.1506 | 5.3852e-06 | 0.1095 | 4.8936e-05 |  | 1000.1506 | 5.3869e-06 | 0.10947 | 4.8974e-05 | 1000.148 | -2.6266 |  | 0.010945
stdev
ppm
Average | UB1000 | 1000.1465 | 0.0044079 | 0.11851 | 0.017085 |  | 1000.1465 | 0.0044079 | 0.11848 | 0.017085 | 1000.1439 | -2.6266 | 4.4072 | 0.011846
Readings from p.51 of UBIII
from cap scale
2020-11-11 |  | C | G |  | C | Correction
|  | pF | nS |  | pF
| UB1000 | 1000.1593 | 0.13362 |  | 1000.1567 | -0.0026205
| GR1000A | 1000.0317 | 0.17503 |  | 1000.0289 | -0.002776
| GR1000B | 1000.0164 | 0.17621 |  | 1000.0139 | -0.002465

## Room data logger
Primary Lab
2014-10-09 14:00 | 20.41 | 42.67
2014-10-09 14:05 | 20.41 | 42.38
2014-10-09 14:10 | 20.45 | 42.55
2014-10-09 14:15 | 20.51 | 42.72
2014-10-09 14:20 | 20.57 | 42.77
2014-10-09 14:25 | 20.52 | 42.87
2014-10-09 14:30 | 20.53 | 42.67
2014-10-09 14:35 | 20.46 | 42.42
2014-10-09 14:40 | 20.41 | 42.55
2014-10-09 14:45 | 20.41 | 42.28
2014-10-09 14:50 | 20.4 | 42.47
2014-10-09 14:55 | 20.44 | 42.3
2014-10-09 15:00 | 20.4 | 42.65
2014-10-09 15:05 | 20.41 | 42.89
2014-10-09 15:10 | 20.41 | 42.74
2014-10-09 15:15 | 20.42 | 42.77
2014-10-09 15:20 | 20.4 | 42.72
2014-10-09 15:25 | 20.33 | 42.79
2014-10-09 15:30 | 20.35 | 43.19
2014-10-09 15:35 | 20.41 | 43.04
2014-10-09 15:40 | 20.45 | 42.96
2014-10-09 15:45 | 20.41 | 43.19
2014-10-09 15:50 | 20.44 | 42.92
2014-10-09 15:55 | 20.4 | 42.89
2014-10-09 16:00 | 20.47 | 43.23
2014-10-09 16:05 | 20.45 | 42.84
2014-10-09 16:10 | 20.41 | 42.92
2014-10-09 16:15 | 20.39 | 42.79
2014-10-09 16:20 | 20.51 | 43.06
2014-10-09 16:25 | 20.51 | 43.23
2014-10-09 16:30 | 20.5 | 43.41
2014-10-09 16:35 | 20.5 | 43.26
2014-10-09 16:40 | 20.55 | 43.48
2014-10-09 16:45 | 20.51 | 43.43
2014-10-09 16:50 | 20.5 | 43.43
2014-10-09 16:55 | 20.56 | 43.65
2014-10-09 17:00 | 20.56 | 43.58
2014-10-09 17:05 | 20.55 | 43.68
2014-10-09 17:10 | 20.47 | 43.63
2014-10-09 17:15 | 20.49 | 43.75
2014-10-09 17:20 | 20.42 | 43.77
2014-10-09 17:25 | 20.49 | 43.63
2014-10-09 17:30 | 20.52 | 43.73
2014-10-09 17:35 | 20.53 | 43.48
2014-10-09 17:40 | 20.5 | 43.38
2014-10-09 17:45 | 20.46 | 43.5
2014-10-09 17:50 | 20.46 | 43.38
2014-10-09 17:55 | 20.45 | 43.53
2014-10-09 18:00 | 20.49 | 43.43
average | 20.4612 | 43.0735
sd | 0.057686 | 0.43176

## Universal Bridge
New UB calibrations need to be entered in Data block and accessed through the Date |  |  | This table is the usual form for calculating spreadsheet corrections | [6 empty] | Date | C1 | R4A | R4B | R4C | G1 | G2 | c | d
|  |  | Component | Value | Date |  |  |  | 1.6 kHz | 1985-02-14 | 32 | 30.1 | -12 | 53 | 13 | 8.7 | <10 | <10 | except for range 4, up to 44? |  |  |  | -91.32 | 0 | 0 | 0 | 0 | 0
|  |  |  | (ppm) | 2014-07-07 |  |  |  | DC for resistors | 1988-11-30 | 123.32 | 35 | -12 | #N/A | 9 | 4 | 0 | 0 |  |  |  |  | 0 | 4.9 | 0 | #N/A | -4 | -4.7 | 0
Date | Record | Comments | C_1 | 143.1 | UBIII |  |  |  | DC for resistors | 1991-03-12 | 133 | 37 | -14 | 61 | 10 | 6 | 0 | 0 |  |  |  |  | 9.68 | 6.9 | -2 | 8 | -3 | -2.7 | 0
1981-05-22 | PEL Meas V | Check Y&Z ranges 6 & 4 @ 400 Hz | R_4A | 54.1261 | UBIII |  |  |  | DC for resistors | 1995-08-22 | 127.406 | 40.34 | -17.21 | 52 | 11.21 | 6.21 | 0 | 0 |  |  |  |  | 4.086 | 10.24 | -5.21 | -1 | -1.79 | -2.49 | 0
1983-12-15 | Universal Bridge II | Full calibration of standards and of all dial ratios and performance checks. | R_4B | -9.8547 | UBIII |  |  |  | DC for resistors | 1998-11-03 | 135.27 | 43.3741 | -16.3523 | 67.0039 | 11.1844 | 6.6891 | 0 | 0 |  |  |  |  | 11.95 | 13.2741 | -4.3523 | 14.0039 | -1.8156 | -2.0109 | 0
1984-02-10 |  | completion of above | R_4C | 67.1933 | UBIII |  |  |  | DC for resistors | 2001-09-08 | 124.1565 | 46.006 | -17.3198 | 64.091 | 12.0228 | 7.4763 | 0 | 0 |  |  |  |  | 0.83653 | 15.906 | -5.3198 | 11.091 | -0.97725 | -1.2237 | 0
1988-11-30 | UB II | component values measured | G_1 | 12.9485 | UBIII |  |  |  | DC for resistors | 2004-01-22 | 131.7148 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 0 | 0 |  |  |  |  | 8.3948 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 0
1991-03-12 | UB II | component values measured | G_2 | 7.4961 | UBIII |  |  |  | DC for resistors | 2004-02-18 | 123.9 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 5.2 | 0 |  |  |  |  | 0.58 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 5.2
1991-05-16 | UB II | problem with '6' dial, repaired! | [6 empty] | DC for resistors | 2005-02-18 | 129.4493 | 46.54 | -16.96 | 61.5 | 13.84 | 9 | 5.7 | 0 |  |  |  |  | 6.1293 | 16.44 | -4.96 | 8.5 | 0.84 | 0.3 | 5.7
1991-06-18 | AI III | problem with 0.01 mF on range 5Y, lack of DC path for current pump, AVOID | Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | assume negligible c,d |  |  | DC for resistors | 2009-08-13 | 130.2987 | 50.8747 | -14.1043 | 64.5247 | 12.5073 | 7.4753 | 2 | 0 |  |  |  |  | 6.9787 | 20.7747 | -2.1043 | 11.5247 | -0.4927 | -1.2247 | 2
1995-08-22 | UB II | bridge calibration | where Z3=R+jwC | correction_c | 10.97 |  |  |  | DC for resistors | 2014-07-07 | 143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 19.78 | 24.0261 | 2.1453 | 14.1933 | -0.051497 | -1.2039 | 10.97
1998-11-03 | UB II | bridge calibration, using HP3458 as transfer instrument | Y4=(G3/G2)(aG1+ßjwC1)(1+c+jd) |  |  | (G3=1/R4*)
2001-09-08 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem fixed | where Y4=G+jwC
2004-01-22 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem REALLY fixed
2004-02-18 | UBII | remeasure capacitor, same dc resistance values as for 22/01/04 | R4A is used for ranges 5 to 1 inclusive
2005-02-18 | UBII | bridge calibration, following MSLT.E.023.004. Uncertainties have increased. | R4B is used on range 6 only
2009-08-13 | UBIII | bridge calibration, following MSLT.E.023.004/5. Uncertainties have increased slightly from 2005. | R4C is used on range 7 only | [16 empty] | AH2700A | ppm | mrad
2014-07-07 | UBIII | bridge calibration, following MSLT.E.023.006. Uncertainties unchanged from 2009. | [9 empty] | Zeros | Range | a | b |  |  |  |  | R_4C | 15.0002 | 0.096
[13 empty] | 2Z | -2 | -15
[13 empty] | 3Z | -1 | -10
|  |  |  | Range | a correction | b correction
[5 empty] | (ppm) | (ppm)
|  |  |  | 1Z | 200.7 | 70.5485
|  |  |  | 2Z | 200.7 | 70.5485
|  |  |  | 3Z | 200.7 | 70.5485
|  |  |  | 4Z | 200.7 | 70.5485
|  |  |  | 5Z | 200.7 | 70.5485
|  |  |  | 6Z | 136.7192 | 6.5677
|  |  |  | 7Z | 213.7672 | 83.6157
|  |  |  | 1Y | 70.5079 | -59.6436
|  |  |  | 2Y | 70.5079 | -59.6436
|  |  |  | 3Y | 70.5079 | -59.6436
|  |  |  | 4Y | 70.5079 | -59.6436
|  |  |  | 5Y | 70.5079 | -59.6436
|  |  |  | 6Y | 134.4887 | 4.3372
|  |  |  | 7Y | 57.4407 | -72.7108
[19 empty] | Symbol font versions |  | Plain version
[5 empty] | Multiplier (C&L) | Unit Name | Unit Multiplier | Multiplier (G&R) | Unit Name | Unit Multiplier | CMC n for L | CMC m for R | [6 empty] | Unit Name | Unit Name | Unit Name | Unit Name
|  |  |  | 1Z | 0.0001 | microH | 1e-06 | 0.0001 | milliohm | 0.001 | 0.2 | 0.002 | [6 empty] | mH | mW | microH | milliohm
|  |  |  | 2Z | 0.001 | microH | 1e-06 | 1e-06 | ohm | 1 | 2 | 0.02 | [6 empty] | mH | W | microH | ohm
|  |  |  | 3Z | 1e-05 | milliH | 0.001 | 1e-05 | ohm | 1 | 20 | 0.2 | [6 empty] | mH | W | milliH | ohm
|  |  |  | 4Z | 0.0001 | milliH | 0.001 | 0.0001 | ohm | 1 | 200 | 2 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 5Z | 0.001 | milliH | 0.001 | 0.001 | ohm | 1 | 2000 | 20 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 6Z | 1e-05 | H | 1 | 1e-05 | kilohm | 1000 | 20000 | 200 | [6 empty] | H | kW | H | kilohm
|  |  |  | 7Z | 0.0001 | H | 1 | 0.0001 | kilohm | 1000 | 200000 | 2000 | [6 empty] | H | kW | H | kilohm
|  |  |  | 1Y | 0.0001 | microF | 1e-06 | 0.0001 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 2Y | 1e-05 | microF | 1e-06 | 1e-05 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 3Y | 1e-06 | microF | 1e-06 | 1e-06 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 4Y | 0.0001 | nF | 1e-09 | 0.0001 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 5Y | 1e-05 | nF | 1e-09 | 1e-05 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 6Y | 0.001 | pF | 1e-12 | 0.001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
|  |  |  | 7Y | 0.0001 | pF | 1e-12 | 0.0001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
Note on M9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on N9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on O9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on P9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on Q9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on R9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on V18: k.jones: AH2700A check of 1 M at 1.5915 k. AI VII, p.42
Note on M41: k.jones: The n/f or m/f factor for CMC calculation.

## Uncertainty
UNCERTAINTY |  |  |  |  | Information for calculating expanded uncertainties |  |  |  | Parameters could be dependent on component |  |  |  |  | This is a data table with B9 as the column input cell. | [12 empty] | CMC
GEUM uncertainty calculation template |  |  |  |  | k-factor from degrees of freedom | [8 empty] | This must be re-calculated with new data sets. | [12 empty] | L uncert = Q[0.000014L , 0.001R, n/f ] |  | R uncert = Q[0.000014R,0.0000001L , m/f ]
The component label is the key variable for the table |  |  |  |  | a | 1.96 |  | Model for uncertainty calculation. | Label | R alpha | u_R alpha | L/C alpha | u_L/C alpha | Label | L/C uncertainty | k | R/G uncertainty | k | Range | L/C | mult | n | R/G | mult | m | f | L in microH, R in ohm | CMC/U | L in microH, R in ohm | CMC/U
[5 empty] | b | 2.372 |  | Base on formula for Universal bridge....ie assign an uncertainty to each term. |  | 0.4%/degree |  |  |  |  | 7.0355e-05 | 2.069 | 0.024639 | 2.1612 | [8 empty] | microH |  | ohm
The 95% level of confidence expanded uncertainties are calculated using the Welch-Satterthwaite |  |  |  |  | c | 1.2737 |  |  | A1 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A1 | 7.0355e-05 | 2.069 | 0.024639 | 2.1612 | 3Z | 0.9942 | 0.001 | 20 | 5.6915 | 1 | 0.2 | 1592.22 | 0.019594 | 0.2785 | 0.00017892 | 0.0072614
formula to combine standard uncertainties. The individual degrees of freedom are derived either |  |  |  |  | d | 7.0962 |  |  | A2 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A2 | 0.00012681 | 2.0919 | 0.03629 | 2.1612 | 3Z | 1.9808 | 0.001 | 20 | 8.3825 | 1 | 0.2 | 1592.31 | 0.031576 | 0.24899 | 0.00026227 | 0.0072269
from the number of measurements used to derive the uncertainty or an estimate | [8 empty] | A3 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A3 | 0.00018622 | 2.105 | 0.04616 | 2.1611 | 3Z | 2.9697 | 0.001 | 20 | 10.6619 | 1 | 0.2 | 1592.41 | 0.044721 | 0.24016 | 0.00035531 | 0.0076974
of the percentage uncertainty in the uncertainty. | [8 empty] | A4 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A4 | 0.00024566 | 2.1114 | 0.054756 | 2.1611 | 3Z | 3.9494 | 0.001 | 20 | 12.647 | 1 | 0.2 | 1592.51 | 0.058094 | 0.23648 | 0.00045067 | 0.0082305
Component label | A1 | [7 empty] | A5 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A5 | 0.00030557 | 2.115 | 0.062413 | 2.161 | 3Z | 4.9333 | 0.001 | 20 | 14.415 | 1 | 0.2 | 1592.6 | 0.071664 | 0.23453 | 0.00054761 | 0.008774
Universal bridge reactance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | A6 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A6 | 0.00036551 | 2.1173 | 0.069713 | 2.161 | 3Z | 5.916 | 0.001 | 20 | 16.1005 | 1 | 0.2 | 1592.68 | 0.085303 | 0.23338 | 0.00064542 | 0.0092582
| Standard | for | 99399 | L/C dial setting |  |  |  |  | A7 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A7 | 0.00042519 | 2.1188 | 0.076827 | 2.1609 | 3Z | 6.8929 | 0.001 | 20 | 17.7431 | 1 | 0.2 | 1592.75 | 0.098918 | 0.23265 | 0.00074336 | 0.0096758
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | A8 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A8 | 0.00048461 | 2.12 | 0.083388 | 2.1609 | 3Z | 7.8653 | 0.001 | 20 | 19.2576 | 1 | 0.2 | 1592.85 | 0.11249 | 0.23212 | 0.00084088 | 0.010084
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment |  | A9 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A9 | 0.00054385 | 2.1208 | 0.089693 | 2.1609 | 3Z | 8.8345 | 0.001 | 20 | 20.7131 | 1 | 0.2 | 1592.92 | 0.12603 | 0.23174 | 0.00093827 | 0.010461
ac/dc error of standards | 0.2485 | 0.061751 | 30 | 5.5556 | 0.00068637 | 5ppm considered two sigma limit, essentially relying on dc values of resistors |  |  | A10 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A10 | 0.00060321 | 2.1214 | 0.09609 | 2.1608 | 3Z | 9.8051 | 0.001 | 20 | 22.19 | 1 | 0.2 | 1591.29 | 0.13962 | 0.23146 | 0.0010362 | 0.010784
dial linearity | 1 | 1 | 30 | 5.5556 | 0.18 | 1 ppm of full scale |  |  | B1 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B1 | 0.0076874 | 2.042 | 0.0050478 | 2.1613 | 2Z | 99.1589 | 1e-06 | 2 | 1.1661 | 1 | 0.02 | 1591.73 | 0.0022058 | 0.28694 | 2.2863e-05 | 0.0045293
amplifier/ratio | 1.6898 | 2.8554 | 30 | 5.5556 | 1.4676 | based on scatter of results in 1984, increased in 2004 and 2005 | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. |  | B2 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B2 | 0.013529 | 2.063 | 0.0075276 | 2.1613 | 2Z | 199.0329 | 1e-06 | 2 | 1.7389 | 1 | 0.02 | 1591.79 | 0.0035167 | 0.25994 | 3.3863e-05 | 0.0044985
temperature coefficient of resistors | 0.1988 | 0.039521 | 50 | 2 | 0.00078094 | depends on matching of Vishay components |  |  | B3 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B3 | 0.019545 | 2.078 | 0.0095498 | 2.1613 | 2Z | 298.3009 | 1e-06 | 2 | 2.2061 | 1 | 0.02 | 1591.84 | 0.0048873 | 0.25006 | 4.4739e-05 | 0.0046848
temperature coefficient of capacitor | 0.099399 | 0.0098802 | 30 | 5.5556 | 1.7571e-05 | 2 ppm per degree for capacitor |  |  | B4 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B4 | 0.025714 | 2.0836 | 0.012083 | 2.1613 | 2Z | 396.5866 | 1e-06 | 2 | 2.7912 | 1 | 0.02 | 1591.91 | 0.0063401 | 0.24656 | 5.7076e-05 | 0.0047237
component stability | 0.49699 | 0.247 | 50 | 2 | 0.030505 | some doubt on stability of capacitor in particular, increased in 2004 | Increase from 2 to 5 ppm to account for short-term instability of 1000 pF capacitor |  | B5 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B5 | 0.031708 | 2.09 | 0.013742 | 2.1613 | 2Z | 494.5412 | 1e-06 | 2 | 3.1744 | 1 | 0.02 | 1591.99 | 0.0077195 | 0.24345 | 6.7665e-05 | 0.0049241
L/C error due to phase error in R/G | 0.91056 | 0.82912 | - | 13 | 0.05288 | 16 ppm of resistance dial setting, see comment for resistance, reduced to 5 ppm after 2004 cal | Link to conductance setting, previously just 2 counts |  | B6 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B6 | 0.037781 | 2.0945 | 0.015314 | 2.1613 | 2Z | 593.5961 | 1e-06 | 2 | 3.5375 | 1 | 0.02 | 1591.64 | 0.0091189 | 0.24136 | 7.8321e-05 | 0.0051145
resolution | 1.2561 | 1.5778 | 50 | 2 | 1.2447 | this might be range and frequency dependent | Previous was 2, use latest value from MSLT.E.023.005, 200/f |  | B7 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B7 | 0.043794 | 2.0978 | 0.01684 | 2.1613 | 2Z | 691.4347 | 1e-06 | 2 | 3.8901 | 1 | 0.02 | 1592.21 | 0.010508 | 0.23994 | 8.8908e-05 | 0.0052795
| 0 | 0 | 1 | 5000 | 0 |  |  |  | B8 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B8 | 0.04978 | 2.1007 | 0.01815 | 2.1613 | 2Z | 789.2294 | 1e-06 | 2 | 4.1925 | 1 | 0.02 | 1592.28 | 0.011884 | 0.23874 | 9.9155e-05 | 0.0054632
| Sum of variance | 6.6204 |  | Sum of u^4/v | 2.9772 |  |  |  | B9 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B9 | 0.055816 | 2.1028 | 0.019522 | 2.1613 | 2Z | 887.4941 | 1e-06 | 2 | 4.5095 | 1 | 0.02 | 1592.34 | 0.013278 | 0.23788 | 0.00010964 | 0.005616
| u combined (counts) | 2.573 |  | effective v | 14.7221 |  |  |  | B10 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B10 | 0.061804 | 2.1045 | 0.02085 | 2.1612 | 2Z | 984.9506 | 1e-06 | 2 | 4.8163 | 1 | 0.02 | 1592.38 | 0.01466 | 0.2372 | 0.00012002 | 0.0057565
|  |  |  | k | 2.1292 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 5.4785 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ For all frequencies. Additional noise typical at 50 Hz needs to be added | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Phase angle assumed small, ie no dial interaction. | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3/ Consistent with 1984 calibration, the IEEE paper and recent complement checks | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
4/ Updated following 2004/5 UB calibration, see AUDIREGC | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Universal bridge resistance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | for | 569099 | conductance dial setting |  | all references to the conductance dial are now to cell D34 (previously separate lookups back to results page) | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
G/R error due to phase error in L/C | 15.9038 | 252.9321 | - | 13 | 4921.1277 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | Previous was 16, use latest value of 5 from MSLT.E.023.005 | Now return to 16 to stay consistent with CMCS | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
dial linearity | 10 | 100 | 30 | 5.5556 | 1800 |  | Previous value was 2, now decreased to 1 ppm following MSLT.E023.005 | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ac/dc error of standards | 14.2275 | 202.421 | 30 | 5.5556 | 7375.3703 | 5ppm considered two sigma limit, essentially relying on dc values of resistors | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
resolution | 12.5611 | 157.7807 | 30 | 5.5556 | 4481.0541 | this might be dependent on frequency and range, included below, increased from 0 to 1 in 2006, now 2000/f |  | Increase to 5 for 2011 measurements. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
amplifier/ratio | 9.6747 | 93.5995 | 30 | 5.5556 | 1576.9557 | based on scatter of results in 1984, increased following 2004 and 2005 calibration | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient of resistors | 11.382 | 129.5495 | 50 | 2 | 8391.5324 | depends on matching of Vishay components | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
component stability | 11.382 | 129.5495 | 30 | 5.5556 | 3020.9517 | stability of resistors only | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 50 | 2 | 0 |  | Previous value was 1, now increased to 2 ppm following MSLT.E023.005 | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 1 | 5000 | 0 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Sum of variance | 1065.8323 |  | Sum of u^4/v | 31566.9919 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u combined (counts) | 32.6471 |  | effective v | 35.9869 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  |  | k | 2.027 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 66.1772 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ This scatter is consistent with Thompson's phase standards. | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Updated following 200/54 UB calibration see AUDIREGC | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Resistance uncertainty for Sullivan inductor |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | [17 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ubridge uncertainty | 32.6471 | 1065.8323 | - | 35.9869 | 31566.9919 | ubridge count uncertainty | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
short term stability | 0 | 0 | 30 | 5.5556 | 0 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient | 1139.6199 | 1298733.4239 | 20 | 12.5 | 134936680512.975 | assume u_t = 0.5 degree
auxiliary balance senstivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget
| Sum of variance | 1299799.2562 |  | Sum of u^4/v | 134936712079.9669
| u combined (counts) | 1140.0874 |  | effective v | 12.5205
|  |  |  | k | 2.1612
Assumptions |  |  |  | expanded u (counts) | 2463.9447
1/ |  |  |  |  | 0.024639 | ohm
Inductance uncertainty for Sullivan inductor |  |  |  |  | k=a+b/v+c/v^2+d/v^3
Standard
Source of uncertainty | Uncertainties |  |  | Degrees of
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v
ubridge uncertainty | 2.573 | 6.6204 | - | 14.7221 | 2.9772
temperature coefficient | 0.99523 | 0.99049 | 30 | 5.5556 | 0.17659 | assume 0.5 deg at one sigma and look up the temp coefficient
short term stability | 1.988 | 3.9521 | 30 | 5.5556 | 2.8114 | 20 ppm of dial setting
auxiliary balance sensitivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget
| Sum of variance | 11.563 |  | Sum of u^4/v | 5.9651
| u combined (count) | 3.4004 |  | effective v | 22.414
|  |  |  | k | 2.069
Assumptions |  |  |  | expanded u (count) | 7.0355
1/ |  |  |  |  | 7.0355e-05 | milliH
Note on B40: k.jones: Modified as for L in keeping with CMCs
Note on B73: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.

## Previous History
[6 empty] | checked by: | Tom been here
| 2005-02-15 | S21586 |  |  | 2005-02-15 |  |  |  |  | 2000-08-11 | S21384 | [5 empty] | 2005-02-25
| 9171/1962 | [8 empty] | 9171/1962 | [6 empty] | dial code | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
Label | Range 2Z | [7 empty] | Label | Range 2Z | [9 empty] | uncertainty | factor |  | uncertainty | factor
B10 | 10 | 984646 | 4821626 | [5 empty] | B10 | 10 | 984617 | 4827668 | [6 empty] | (mH) | (mH)
B9 | 9 | 887241 | 4512760 | [5 empty] | B9 | 9 | 887225 | 4520316 |  |  |  |  | A1 | 1 | 993.9916 | 0.11571 | 2.1008 | 5.6916 | 0.024611 | 2.1611
B8 | 8 | 789040 | 4195500 | [5 empty] | B8 | 8 | 789013 | 4202536 |  |  |  |  | A2 | 2 | 1980.4119 | 5.8869 | 2.1615 | 8.3817 | 0.03664 | 2.1473
B7 | 7 | 691266 | 3892830 | [5 empty] | B7 | 7 | 691241 | 3899426 |  |  |  |  | A3 | 3 | 2968.8226 | 2.0202 | 2.1599 | 10.661 | 0.046165 | 2.1591
B6 | 6 | 593428 | 3543346 |  | 594040 | 3533330 |  |  | B6 | 6 | 593428 | 3543346 |  |  |  |  | A4 | 4 | 3948.7118 | 0.179 | 2.1114 | 12.6462 | 0.054673 | 2.1614
B5 | 5 | 494427 | 3176330 | [5 empty] | B5 | 5 | 494412 | 3181644 |  |  |  |  | A5 | 5 | 4932.4416 | 0.21648 | 2.1362 | 14.4139 | 0.062315 | 2.1614
B4 | 4 | 396493 | 2793140 | [5 empty] | B4 | 4 | 396485 | 2797540 |  |  |  |  | A6 | 6 | 5914.8613 | 0.25412 | 2.1586 | 16.0991 | 0.069599 | 2.1614
B3 | 3 | 298240 | 2207660 | [5 empty] | B3 | 3 | 298228 | 2208660 |  |  |  |  | A7 | 7 | 6891.6899 | 0.2924 | 2.1754 | 17.7417 | 0.0767 | 2.1614
B2 | 2 | 198990 | 1740140 | [5 empty] | B2 | 2 | 198994 | 1740961 |  |  |  |  | A8 | 8 | 7864.0478 | 0.33083 | 2.1889 | 19.2569 | 0.08325 | 2.1614
B1 | 1 | 99142 | 1167100 | [5 empty] | B1 | 1 | 99140 | 1167446 |  |  |  |  | A9 | 9 | 8833.3752 | 0.36924 | 2.2006 | 20.7133 | 0.089546 | 2.1614
| 0 | [8 empty] | 0 | [6 empty] | A10 | 10 | 9803.9928 | 0.40859 | 2.2066 | 22.196 | 0.095955 | 2.1615
| 9173/1962 | [8 empty] | 9173/1962
| Range 3Z | [8 empty] | Range 3Z
A10 | 10 | 980229 | 2219460 | [5 empty] | A10 | 10 | 980273 | 2221419
A9 | 9 | 883184 | 2071200 | [5 empty] | A9 | 9 | 883227 | 2073654 | [5 empty] | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
A8 | 8 | 786268 | 1925567 | [5 empty] | A8 | 8 | 786313 | 1928038 | [7 empty] | uncertainty | factor |  | uncertainty | factor
A7 | 7 | 689049 | 1774060 | [5 empty] | A7 | 7 | 689092 | 1776493 | [6 empty] | (mH) | (mH) |  | (W) | (W)
A6 | 6 | 591383 | 1609810 | [5 empty] | A6 | 6 | 591422 | 1612197 |  |  |  |  | B1 | 1 | 99.1611 | 0.0061381 | 2.0413 | 1.1672 | 0.0050456 | 2.1615
A5 | 5 | 493158 | 1441300 | [5 empty] | A5 | 5 | 493192 | 1443680 |  |  |  |  | B2 | 2 | 199.0264 | 0.0099933 | 2.07 | 1.7402 | 0.007523 | 2.1615
A4 | 4 | 394802 | 1264540 | [5 empty] | A4 | 4 | 394829 | 1266570 |  |  |  |  | B3 | 3 | 298.2935 | 0.01444 | 2.0766 | 2.2078 | 0.0095441 | 2.1615
A3 | 3 | 296830 | 1066030 | [5 empty] | A3 | 3 | 297330 | 1073180 |  |  |  |  | B4 | 4 | 396.5635 | 0.018253 | 2.1065 | 2.7933 | 0.012075 | 2.1615
A2 | 2 | 198006 | 838110 | [5 empty] | A2 | 2 | 198979 | 849340 |  |  |  |  | B5 | 5 | 494.5144 | 0.022373 | 2.1154 | 3.1765 | 0.013732 | 2.1615
A1 | 1 | 99381 | 569116 | [5 empty] | A1 | 1 | 99390 | 569920 |  |  |  |  | B6 | 6 | 594.1446 | 0.12518 | 2.1432 | 3.5335 | 0.015276 | 2.1615
| 0 | -1 | -10 | [6 empty] | 0 | [6 empty] | B7 | 7 | 691.3874 | 0.030514 | 2.1339 | 3.8931 | 0.01683 | 2.1615
[17 empty] | B8 | 8 | 789.1783 | 0.034488 | 2.143 | 4.1957 | 0.018138 | 2.1615
2000-08-11 | S21384 | [15 empty] | B9 | 9 | 887.3962 | 0.038437 | 2.1527 | 4.513 | 0.01951 | 2.1615
9171/1962 | [16 empty] | B10 | 10 | 984.818 | 0.042525 | 2.1556 | 4.8219 | 0.020845 | 2.1615
Range 2Z
10 | 984617 | 4827668 |  |  | 984617 | 4827668
9 | 887225 | 4520316 |  |  | 887225 | 4520316
8 | 789013 | 4202536 |  |  | 789013 | 4202536
7 | 691241 | 3899426 |  |  | 691241 | 3899426
6 | 593428 | 3543346 |  |  | 593428 | 3543346
5 | 494412 | 3181644 |  |  | 494412 | 3181644
4 | 396485 | 2797540 |  |  | 396485 | 2797540
3 | 298228 | 2208660 |  |  | 298228 | 2208660
2 | 198994 | 1740961 |  |  | 198994 | 1740961
1 | 99140 | 1167446 |  |  | 99140 | 1167446
0
9173/1962
Range 3Z
10 | 980273 | 2221419 |  |  | 980273 | 2221419
9 | 883227 | 2073654 |  |  | 883227 | 2073654
8 | 786313 | 1928038 |  |  | 786313 | 1928038
7 | 689092 | 1776493 |  |  | 689092 | 1776493
6 | 591422 | 1612197 |  |  | 591422 | 1612197
5 | 493192 | 1443680 |  |  | 493192 | 1443680
4 | 394829 | 1266570 |  |  | 394829 | 1266570
3 | 297330 | 1073180 |  |  | 297330 | 1073180
2 | 198979 | 849340 |  |  | 198979 | 849340
1 | 99390 | 569920 |  |  | 99390 | 569920
0
1995-09-29 | [7 empty] | 1995-09-29 | [7 empty] | 2005-02-23
9171/1962 | [7 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage |  | 9171/1962 |  |  |  | L % | R %
Range 2Z | [8 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor |  | Range 2Z |  |  |  |  | 3 Month Drift
10 | 984610 | 4837350 | [5 empty] | 1 | 994.03 | 0.054572 | 2.1524 | 5.7182 | 0.024721 | 2.1615 |  | 10 | 984646 | 4821626 |  | 0 | 0
9 | 887218 | 4529080 | [5 empty] | 2 | 1980.3588 | 0.10366 | 2.2108 | 8.4219 | 0.036409 | 2.1615 |  | 9 | 887241 | 4512760 |  | 0 | 0
8 | 789017 | 4210450 | [5 empty] | 3 | 2968.9979 | 0.15407 | 2.2298 | 10.7121 | 0.04631 | 2.1615 |  | 8 | 789040 | 4195500 |  | 0 | 0
7 | 691242 | 3906360 | [5 empty] | 4 | 3948.4356 | 0.20423 | 2.2381 | 12.7063 | 0.054931 | 2.1615 |  | 7 | 691266 | 3892830 |  | 0 | 0
6 | 593424 | 3549500 | [5 empty] | 5 | 4931.9739 | 0.25468 | 2.2428 | 14.4818 | 0.062607 | 2.1615 |  | 6 | 594040 | 3533330 |  | 1.1602 | -3.18
5 | 494410 | 3186950 | [5 empty] | 6 | 5914.202 | 0.30511 | 2.2456 | 16.1744 | 0.069924 | 2.1615 |  | 5 | 494427 | 3176330 |  | 0 | 0
4 | 396487 | 2802060 | [5 empty] | 7 | 6890.9693 | 0.35527 | 2.2475 | 17.8224 | 0.077048 | 2.1615 |  | 4 | 396493 | 2793140 |  | 0 | 0
3 | 298244 | 2214550 | [5 empty] | 8 | 7863.1258 | 0.40521 | 2.2488 | 19.3421 | 0.083619 | 2.1615 |  | 3 | 298240 | 2207660 |  | 0 | 0
2 | 198994 | 1745600 | [5 empty] | 9 | 8832.1617 | 0.455 | 2.2499 | 20.8018 | 0.089929 | 2.1615 |  | 2 | 198990 | 1740140 |  | 0 | 0
1 | 99144 | 1170560 | [5 empty] | 10 | 9802.568 | 0.50486 | 2.2506 | 22.2832 | 0.096333 | 2.1615 |  | 1 | 99142 | 1167100 |  | 0 | 0
0 | [15 empty] | 0
9173/1962 | [16 empty] | 9173/1962
Range 3Z | [16 empty] | Range 3Z
10 | 980098 | 2228210 | [13 empty] | 10 | 980229 | 2219460 |  | 0 | 0
9 | 883073 | 2080080 | [13 empty] | 9 | 883184 | 2071200 |  | 0 | 0
8 | 786185 | 1934120 | [5 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage |  | 8 | 786268 | 1925567 |  | 0 | 0
7 | 688985 | 1782150 | [6 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor |  | 7 | 689049 | 1774060 |  | 0 | 0
6 | 591324 | 1617360 | [5 empty] | 1 | 99.162 | 0.0057279 | 2.0996 | 1.1706 | 0.0050607 | 2.1615 |  | 6 | 591383 | 1609810 |  | 0 | 0
5 | 493117 | 1448110 | [5 empty] | 2 | 199.028 | 0.010709 | 2.1618 | 1.7457 | 0.0075468 | 2.1615 |  | 5 | 493158 | 1441300 |  | 0 | 0
4 | 394779 | 1270570 | [5 empty] | 3 | 298.294 | 0.015774 | 2.1888 | 2.2147 | 0.0095743 | 2.1615 |  | 4 | 394802 | 1264540 |  | 0 | 0
3 | 296852 | 1071150 | [5 empty] | 4 | 396.5528 | 0.020872 | 2.1975 | 2.8022 | 0.012114 | 2.1615 |  | 3 | 296830 | 1066030 |  | 0 | 0
2 | 198003 | 842150 | [5 empty] | 5 | 494.4916 | 0.025901 | 2.2074 | 3.1871 | 0.013778 | 2.1615 |  | 2 | 198006 | 838110 |  | 0 | 0
1 | 99384 | 571790 | [5 empty] | 6 | 593.5215 | 0.030991 | 2.2143 | 3.5497 | 0.015346 | 2.1615 |  | 1 | 99381 | 569116 |  | 0 | 0
0 | [7 empty] | 7 | 691.3553 | 0.036026 | 2.219 | 3.9065 | 0.016888 | 2.1615 |  | 0
[8 empty] | 8 | 789.146 | 0.041052 | 2.2231 | 4.2106 | 0.018203 | 2.1615
1991-05-30 | [7 empty] | 9 | 887.3628 | 0.046106 | 2.2261 | 4.5293 | 0.019581 | 2.1615 |  |  | 2000-08-11
9171/1962 | [7 empty] | 10 | 984.7705 | 0.05112 | 2.2284 | 4.8376 | 0.020913 | 2.1615
Range 2Z | L | R | [14 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage
10 | 1004838 | 4788256 | [15 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor
9 | 904811 | 4485090 | [14 empty] | 1 | 994.07 | 0.054568 | 2.1525 | 5.6995 | 0.02464 | 2.1615
8 | 803682 | 4172580 | [14 empty] | 2 | 1990.1204 | 0.10418 | 2.2107 | 8.4939 | 0.03672 | 2.1615
7 | 703119 | 3873825 | [14 empty] | 3 | 2973.7887 | 0.15432 | 2.2298 | 10.7324 | 0.046397 | 2.1615
6 | 602695 | 3522505 | [14 empty] | 4 | 3948.9357 | 0.20425 | 2.2382 | 12.6663 | 0.054758 | 2.1615
5 | 501222 | 3164915 | [14 empty] | 5 | 4932.724 | 0.25472 | 2.2429 | 14.4375 | 0.062415 | 2.1615
4 | 401408 | 2783495 | [14 empty] | 6 | 5915.1822 | 0.30515 | 2.2457 | 16.1228 | 0.069701 | 2.1615
3 | 301525 | 220245 | [14 empty] | 7 | 6892.0394 | 0.35532 | 2.2475 | 17.7658 | 0.076804 | 2.1615
2 | 200942 | 1734921 | [14 empty] | 8 | 7864.406 | 0.40527 | 2.2489 | 19.2813 | 0.083356 | 2.1615
1 | 99878 | 1163301 | [14 empty] | 9 | 8833.702 | 0.45507 | 2.2499 | 20.7375 | 0.089651 | 2.1615
0 | [16 empty] | 10 | 9804.3182 | 0.50495 | 2.2507 | 22.2153 | 0.096039 | 2.1615
9173/1962
Range 3Z | L | R
10 | 1000042 | 2185630 | [14 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage
9 | 900119 | 2043361 | [15 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor
8 | 800504 | 1903129 | [14 empty] | 1 | 99.158 | 0.0057257 | 2.0999 | 1.1675 | 0.0050472 | 2.1615
7 | 700510 | 1756857 | [14 empty] | 2 | 199.028 | 0.010706 | 2.1621 | 1.741 | 0.0075267 | 2.1615
6 | 600413 | 1596577 | [14 empty] | 3 | 298.278 | 0.015772 | 2.189 | 2.2088 | 0.0095488 | 2.1615
5 | 500100 | 1431956 | [14 empty] | 4 | 396.5508 | 0.020871 | 2.1977 | 2.7977 | 0.012095 | 2.1615
4 | 399841 | 1258189 | [14 empty] | 5 | 494.4936 | 0.025899 | 2.2076 | 3.1818 | 0.013755 | 2.1615
3 | 300168 | 1062034 | [14 empty] | 6 | 593.5255 | 0.030989 | 2.2145 | 3.5435 | 0.015319 | 2.1615
2 | 199991 | 836012 | [14 empty] | 7 | 691.3543 | 0.036025 | 2.2191 | 3.8996 | 0.016858 | 2.1615
1 | 100176 | 568494 | [14 empty] | 8 | 789.142 | 0.041049 | 2.2233 | 4.2027 | 0.018169 | 2.1615
0 | [16 empty] | 9 | 887.3698 | 0.046104 | 2.2262 | 4.5205 | 0.019543 | 2.1615
[17 empty] | 10 | 984.7775 | 0.051118 | 2.2286 | 4.8279 | 0.020872 | 2.1615
1984-02-16
9171/1962 |  |  | R % | L %
Range 2Z | L | R | 3 Month Drift
10 | 1005990 | 4806400 | 0.012772 | 0.0038745
9 | 905840 | 4507200 | 0.016598 | 0.0038435
8 | 804560 | 4198900 | 0.021209 | 0.0036923
7 | 704300 | 3900100 | 0.022794 | 0.0056735
6 | 604250 | 3550300 | 0.026489 | 0.0087071
5 | 503200 | 3193100 | 0.029865 | 0.0133
4 | 401510 | 2810300 | 0.032272 | 0.00085954
3 | 301630 | 2224200 | 3.0484 | 0.0011778
2 | 201130 | 1756160 | 0.04092 | 0.0031626
1 | 100040 | 1180150 | 0.048306 | 0.005479
0 | 358 | 5860 | 3.3835 | 3.3835
9173/1962
Range 3Z | L | R
10 | 1000616 | 2191510 | 0.0090781 | 0.0019409
9 | 900587 | 2049680 | 0.010431 | 0.0017583
8 | 800915 | 1910750 | 0.013495 | 0.0017363
7 | 701003 | 1766150 | 0.017803 | 0.0023795
6 | 601187 | 1605900 | 0.019643 | 0.004356
5 | 500064 | 1441700 | 0.022868 | -0.00024358
4 | 399832 | 1267980 | 0.026126 | -7.616e-05
3 | 300175 | 1071200 | 0.028951 | 7.8901e-05
2 | 199984 | 844290 | 0.033174 | -0.00011843
1 | 100176 | 574990 | 0.038225 | 0
0 | 30 | 602 | 3.3835 | 3.3835
[6 empty] | 1984-02-16 | 1005990 | 905840 | 804560 | 704300 | 604250 | 503200 | 401510 | 100040 | 358 | 1000616 | 900587 | 800915 | 701003 | 601187 | 500064 | 399832 | 300175 | 199984 | 100176 | 30
[6 empty] | 1991-05-30 | 1004838 | 904811 | 803682 | 703119 | 602695 | 501222 | 401408 | 99878 |  | 1000042 | 900119 | 800504 | 700510 | 600413 | 500100 | 399841 | 300168 | 199991 | 100176
[6 empty] | 1995-09-29 | 984610 | 887218 | 789017 | 691242 | 593424 | 494410 | 396487 | 99144 |  | 980098 | 883073 | 786185 | 688985 | 591324 | 493117 | 394779 | 296852 | 198003 | 99384
[6 empty] | 2000-08-11 | 984617 | 887225 | 789013 | 691241 | 593428 | 494412 | 396485 | 99140 |  | 980273 | 883227 | 786313 | 689092 | 591422 | 493192 | 394829 | 297330 | 198979 | 99390
[6 empty] | 2005-02-15 | 984646 | 887241 | 789040 | 691266 | 594040 | 494427 | 396493 | 99142 |  | 980229 | 883184 | 786268 | 689049 | 591383 | 493158 | 394802 | 296830 | 198006 | 99381 | -1
[6 empty] | 2009-07-03 | 984766 | 887330 | 789093 | 691310 | 593513 | 494446 | 396503 | 99136 |  | 980190 | 883164 | 786276 | 689060 | 591400 | 493172 | 394815 | 296875 | 198015 | 99388
[6 empty] | 1995-09-29 | 0.97987 | 0.98056 | 0.98175 | 0.98311 | 0.98462 | 0.98641 | 0.98774 | 0.99265 |  | 0.98006 | 0.98106 | 0.98211 | 0.98355 | 0.98486 | 0.98604 | 0.98734 | 0.98895 | 0.99006 | 0.99209
[6 empty] | 2000-08-11 | 0.97988 | 0.98056 | 0.98175 | 0.98311 | 0.98462 | 0.98641 | 0.98774 | 0.99261 |  | 0.98023 | 0.98123 | 0.98227 | 0.9837 | 0.98503 | 0.98619 | 0.98747 | 0.99055 | 0.99494 | 0.99215
|  | 984766 | 980190 |  |  | 2005-02-15 | 0.97991 | 0.98058 | 0.98178 | 0.98314 | 0.98564 | 0.98644 | 0.98776 | 0.99263 |  | 0.98019 | 0.98119 | 0.98222 | 0.98364 | 0.98496 | 0.98612 | 0.9874 | 0.98888 | 0.99007 | 0.99206
|  | 887330 | 883164 |  |  | 2009-07-03 | 0.98002 | 0.98068 | 0.98185 | 0.9832 | 0.98477 | 0.98648 | 0.98778 | 0.99257 |  | 0.98015 | 0.98116 | 0.98223 | 0.98365 | 0.98499 | 0.98615 | 0.98743 | 0.98903 | 0.99012 | 0.99213
|  | 789093 | 786276 |  |  | 1995-09-29 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | 691310 | 689060 |  |  | 2000-08-11 | 1 | 1 | 0.99999 | 1 | 1 | 1 | 0.99999 | 0.99996 |  | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0001 | 1.0016 | 1.0049 | 1.0001
|  | 593513 | 591400 |  |  | 2005-02-15 | 1 | 1 | 1 | 1 | 1.001 | 1 | 1 | 0.99998 |  | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 0.99993 | 1 | 0.99997
|  | 494446 | 493172 |  |  | 2009-07-03 | 1.0002 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1 | 0.99992 |  | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1
|  | 396503 | 394815
|  | 298239 | 296875
|  | 198991 | 198015 |  |  |  | 34971
|  | 99136 | 99388
500