## Commentary
This (S21950.xlsx) starts as an xlsx saving of S21814_TJS checked_final.xls, the previous report.
Commentary page is added.
Note that the 'OPEN' has been measured on range 7Y, rather than 6Y, with consequent changes in UB calculation
Replace 'Universal Bridge' with worksheet from S21952v1 and carefully redefine named blocks.
Correct some arial fonts for a couple of greek symbols in the report
Note that the uncertainty construction in S21814 is very close to the most recent approach
Small tidy up with UB back to 16 ppm dial interaction and temperature coefficient formula of 16074A fixed.
Realign report columns for tidier text layout
Improve 'Previous reports' page so that entering cell address in A50 will plot history for one item
Measured approx temperature coefficient of 100 mH and have updated the uncertainty budget as a result. Measured -300 ppm/degree rather than 10!
Added CMCFUDGE block so that the 53 Hz resistance uncertainties can be increased to match the CMC calculation.
In the end CMCFUDGE was not used as the R resolution uncertainty term was increased in line with the CMC formula.
Change resolution degrees of freedom to 30% accuracy. Otherwise this dominates at the zeros and gives k > 4

## Instructions
This template is created starting with 'Muirhead Decade L template II' and adding the reporting structure of 'HP16074A' from the E023 directory.
Note that an overestimation by a factor of 10 was found in the UB resistance uncertainty of the Muirhead decade L. Not significant as dominated by copper temperature.
Template to be completed
There are 4 Instructions |  | Test Data
1 | Universal bridge dial readings are entered on Results page | a | b
|  | 1 | 70
|  | 101503 | 193646
|  | 199897 | 278425
|  | 292665 | 342702
|  | 393478 | 403100
|  | 490494 | 456415
|  | 590914 | 507265
|  | 689345 | 554089
|  | 788688 | 599514
|  | 889705 | 644643
|  | 1010757 | 694910
|  | 0 | -11
|  | 103882 | 167100
|  | 207655 | 242290
|  | 304252 | 299420
|  | 398861 | 347800
|  | 494623 | 393060
|  | 592256 | 436120
|  | 689692 | 476900
|  | 782948 | 514970
|  | 878078 | 552860
|  | 984766 | 593284
2 | Previous values are entered on Results page | 2006-06-15 | Previous Calibration, S21637
| These are used in the uncertainty calculation | Inductance | Resistance
| The date must be entered to allow drift calculation | (mH) | (W)
|  | 10.1477 | 19.3261
|  | 19.985 | 27.7966
|  | 29.2602 | 34.218
|  | 39.3402 | 40.2543
|  | 49.0405 | 45.5806
|  | 59.082 | 50.6659
|  | 68.9246 | 55.3442
|  | 78.8583 | 59.8824
|  | 88.9616 | 64.3877
|  | 101.0768 | 69.482
|  | 103.9039 | 167.1075
|  | 207.6929 | 242.3458
|  | 304.3075 | 299.5391
|  | 398.9299 | 347.9799
|  | 494.7064 | 393.3124
|  | 592.3583 | 436.3849
|  | 689.8051 | 477.3772
|  | 783.0962 | 515.7194
|  | 878.2876 | 553.7116
|  | 985.1001 | 594.3189
3 | Check that Ubridge calibration is up to date on Universal Bridge page
This page will need to be unlocked to load new calibration data into the Data block.
4 | The Previous reports page is unprotected as it is not used in the final report.
It is strongly recommended that the history be reviewed before issuing the report
|  |  |  | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
[5 empty] | Change | Uncertainty | Factor | Change | Uncertainty | Factor
| Pasting the test data above into this sheet will give the following report tables |  |  |  | (mH) | (mH) |  | (W) | (W)
|  |  |  | 10 | 101.0934 | 0.0070 | 2.166 | 69.49 | 0.30 | 2.1577
|  |  |  | 9 | 88.9860 | 0.0065 | 2.1311 | 64.46 | 0.28 | 2.1574
|  |  |  | 8 | 78.8826 | 0.0059 | 2.1235 | 59.95 | 0.26 | 2.1574
|  |  |  | 7 | 68.9465 | 0.0052 | 2.1211 | 55.41 | 0.24 | 2.1574
|  |  |  | 6 | 59.1017 | 0.0045 | 2.118 | 50.72 | 0.22 | 2.1575
|  |  |  | 5 | 49.0579 | 0.0038 | 2.1142 | 45.64 | 0.20 | 2.1574
|  |  |  | 4 | 39.3546 | 0.0031 | 2.1116 | 40.31 | 0.17 | 2.1574
|  |  |  | 3 | 29.2715 | 0.0023 | 2.1068 | 34.27 | 0.15 | 2.1573
|  |  |  | 2 | 19.9931 | 0.0016 | 2.0989 | 27.84 | 0.12 | 2.1571
|  |  |  | 1 | 10.1520 | 0.0009 | 2.0722 | 19.36 | 0.08 | 2.1566
|  |  |  | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
[5 empty] | Change | Uncertainty | Factor | Change | Uncertainty | Factor
[5 empty] | (mH) | (mH) |  | (W) | (W)
|  |  |  | 10 | 984.939 | 0.068 | 2.1662 | 593.3 | 2.6 | 2.1567
|  |  |  | 9 | 878.232 | 0.060 | 2.1943 | 552.9 | 2.4 | 2.1569
|  |  |  | 8 | 783.086 | 0.053 | 2.1995 | 515.0 | 2.2 | 2.157
|  |  |  | 7 | 689.813 | 0.047 | 2.1988 | 476.9 | 2.1 | 2.1574
|  |  |  | 6 | 592.360 | 0.040 | 2.1978 | 436.2 | 1.9 | 2.1576
|  |  |  | 5 | 494.710 | 0.034 | 2.1957 | 393.1 | 1.7 | 2.1576
|  |  |  | 4 | 398.931 | 0.027 | 2.1924 | 347.8 | 1.5 | 2.1577
|  |  |  | 3 | 304.305 | 0.021 | 2.1855 | 299.4 | 1.3 | 2.1577
|  |  |  | 2 | 207.691 | 0.014 | 2.169 | 242.3 | 1.1 | 2.1576
|  |  |  | 1 | 103.900 | 0.008 | 2.1117 | 167.1 | 0.7 | 2.1571

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | rl |  | 4: Client; 10: Enter, say, 5 ppm as 5e-6
Date from |  |  | 4: Job number; 10: 1 sigma uncertainties from specifications
to |  |  | 4: S Number
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
53_0 ohm | 53 | 1Z | 90 | 38290 | 53 | 19.71 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
400_0 ohm | 400 | 1Z | 70 | 38261 | 400 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1000_0 ohm | 1000 | 1Z | 65 | 38470 | 1000 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_Short | 53 | 1Z | 40 | 10 | 53 | 19.7 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
400_Short | 400 | 1Z | -10 | 11 | 400 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1000_Short | 1000 | 1Z | -15 | -50 | 1000 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 53 | 1Z | 80 | -20 | 53 | 19.67 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 1Z | -11 | 1 | 400 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 1Z | -13 | 0 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_0.1 ohm | 53 | 1Z | 180 | 1050700 | 53 | 19.72 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10
400_0.1 ohm | 400 | 1Z | 62 | 1051191 | 400 | 19.74 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10
1000_0.1 ohm | 1000 | 1Z | 67 | 1051120 | 1000 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10
53_1 ohm | 53 | 1Z | 100 | 9879850 | 53 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
400_1 ohm | 400 | 1Z | 43 | 9879941 | 400 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
1000_1 ohm | 1000 | 1Z | 43 | 9879760 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
coax zero | 53 | 2Z | 100 | 10 | 53 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 2Z | 1 | -4 | 400 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 2Z | -2 | -8 | 1000 | 19.78 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_10 ohm | 53 | 2Z | 120 | 9913100 | 53 | 19.78 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
400_10 ohm | 400 | 2Z | 5 | 9913067 | 400 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
1000_10 ohm | 1000 | 2Z | 5 | 9912930 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10
coax zero | 53 | 3Z | 30 | -13 | 53 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 3Z | -1 | -9 | 400 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 3Z | -1 | -7 | 1000 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_100 ohm | 53 | 3Z | 50 | 10002950 | 53 | 19.67 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
400_100 ohm | 400 | 3Z | 5 | 10002950 | 400 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
1000_100 ohm | 1000 | 3Z | 5 | 10002954 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
coax zero | 53 | 4Z | 35 | -10 | 53 | 19.63 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 4Z | 0 | -11 | 400 | 19.64 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 4Z | -1 | -11 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_1000 ohm | 53 | 4Z | 10 | 10001710 | 53 | 19.63 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
400_1000 ohm | 400 | 4Z | 5 | 10001713 | 400 | 19.64 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
1000_1000 ohm | 1000 | 4Z | 5 | 10001670 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
coax zero | 53 | 5Z | 10 | -28 | 53 | 19.49 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 5Z | 0 | -26 | 400 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 5Z | 0 | -22 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_10 k ohm | 53 | 5Z | -67 | 10001342 | 53 | 19.5 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
400_10 k ohm | 400 | 5Z | -63 | 10001340 | 400 | 19.51 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
1000_10 k ohm | 1000 | 5Z | -61 | 10001312 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
coax zero | 53 | 6Z | 10 | -7 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 6Z | 0 | -10 | 400 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 6Z | -1 | -10 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_100 k ohm | 53 | 6Z | -800 | 10002700 | 53 | 19.54 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
400_100 k ohm | 400 | 6Z | -777 | 10002664 | 400 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
1000_100 k ohm | 1000 | 6Z | -769 | 10002551 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10
coax zero | 1600 | 4Z | -1 | -8 | 1600 | 19.56 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1600_100 micro | 1600 | 4Z | 1001 | 6324510 | 1600 | 19.56 | 0.5 | 10 | 0 | 0.0001575 | 10 | 0 | 5e-06 | 10
coax zero | 1600 | 5Z | -1 | -20 | 1600 | 19.56 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1600_100 milli | 1600 | 5Z | 99862 | 2000313 | 1600 | 19.56 | 0.5 | 10 | 0 | 0.0001575 | 10 | 0 | 5e-06 | 10
coax zero | 53 | 7Y | 30 | 10 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 400 | 7Y | 6 | 15 | 400 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
coax zero | 1000 | 7Y | 6 | 16 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
53_Open | 53 | 7Y | 30 | 10 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
400_Open | 400 | 7Y | 7 | 15 | 400 | 19.54 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1000_Open | 1000 | 7Y | 6 | 16 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
53_0 ohm | 53 | 1Z | 90 | 38290 | 53 | 19.71 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 1.0002e-09 | 3.9543e-09 | 2.0845 | 0.0038313 | 4.7927e-05 | 2.0256 | 3.7736e-09 | 3.7736e-05 | 1.0479 | 1.2701 | 53_0 ohm53 | 1000 | 3.8313 | 0.047927 | 2.0256 | 1000000 | 0.0010002 | 0.0039543 | 2.0845 | 53
400_0 ohm | 400 | 1Z | 70 | 38261 | 400 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 8.1017e-09 | 9.669e-10 | 2.064 | 0.0038263 | 3.0724e-05 | 2.0622 | 5.0001e-10 | 5.0003e-06 | 1.9337 | 6.1444 | 400_0 ohm400 | 1000 | 3.8263 | 0.030724 | 2.0622 | 1000000 | 0.0081017 | 0.0009669 | 2.064 | 400
1000_0 ohm | 1000 | 1Z | 65 | 38470 | 1000 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 7.8017e-09 | 5.5305e-10 | 2.0343 | 0.0038473 | 2.9909e-05 | 2.0767 | 2.0004e-10 | 2.0007e-06 | 2.7647 | 14.9492 | 1000_0 ohm1000 | 1000 | 3.8473 | 0.029909 | 2.0767 | 1000000 | 0.0078017 | 0.00055305 | 2.0343 | 1000
53_Short | 53 | 1Z | 40 | 10 | 53 | 19.7 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | -4.0009e-09 | 3.9543e-09 | 2.0845 | 3.0002e-06 | 4.7927e-05 | 2.0256 | 3.7736e-09 | 3.7736e-05 | 1.0479 | 1.2701 | 53_Short53 | 1000 | 0.0030002 | 0.047927 | 2.0256 | 1000000 | -0.0040009 | 0.0039543 | 2.0845 | 53
400_Short | 400 | 1Z | -10 | 11 | 400 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 1.0002e-10 | 9.6683e-10 | 2.064 | 1.0001e-06 | 3.0723e-05 | 2.0622 | 5e-10 | 5e-06 | 1.9337 | 6.1447 | 400_Short400 | 1000 | 0.0010001 | 0.030723 | 2.0622 | 1000000 | 0.00010002 | 0.00096683 | 2.064 | 400
1000_Short | 1000 | 1Z | -15 | -50 | 1000 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | -2.0004e-10 | 5.5292e-10 | 2.0344 | -5.0004e-06 | 2.9909e-05 | 2.0768 | 2e-10 | 2e-06 | 2.7646 | 14.9545 | 1000_Short1000 | 1000 | -0.0050004 | 0.029909 | 2.0768 | 1000000 | -0.00020004 | 0.00055292 | 2.0344 | 1000
coax zero | 53 | 1Z | 80 | -20 | 53 | 19.67 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7736e-09 | 3.7736e-05 | 0 | 0 | coax zero53 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 53
coax zero | 400 | 1Z | -11 | 1 | 400 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-10 | 5e-06 | 0 | 0 | coax zero400 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 400
coax zero | 1000 | 1Z | -13 | 0 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-10 | 2e-06 | 0 | 0 | coax zero1000 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1000
53_0.1 ohm | 53 | 1Z | 180 | 1050700 | 53 | 19.72 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10 | 1.0002e-08 | 3.9673e-09 | 2.0822 | 0.10508 | 4.8599e-05 | 2.0215 | 3.775e-09 | 3.7765e-05 | 1.0509 | 1.2869 | 53_0.1 ohm53 | 1 | 0.10508 | 4.8599e-05 | 2.0215 | 1000000 | 0.010002 | 0.0039673 | 2.0822 | 53
400_0.1 ohm | 400 | 1Z | 62 | 1051191 | 400 | 19.74 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10 | 7.3016e-09 | 1.0234e-09 | 2.0426 | 0.10513 | 3.1706e-05 | 2.0474 | 5.1093e-10 | 5.2121e-06 | 2.003 | 6.0832 | 400_0.1 ohm400 | 1 | 0.10513 | 3.1706e-05 | 2.0474 | 1000000 | 0.0073016 | 0.0010234 | 2.0426 | 400
1000_0.1 ohm | 1000 | 1Z | 67 | 1051120 | 1000 | 19.75 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-05 | 10 | 8.0017e-09 | 6.5539e-10 | 2.0181 | 0.10512 | 3.0891e-05 | 2.0588 | 2.2594e-10 | 2.4831e-06 | 2.9007 | 12.4403 | 1000_0.1 ohm1000 | 1 | 0.10512 | 3.0891e-05 | 2.0588 | 1000000 | 0.0080017 | 0.00065539 | 2.0181 | 1000
53_1 ohm | 53 | 1Z | 100 | 9879850 | 53 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 2.0004e-09 | 5.172e-09 | 2.046 | 0.98806 | 7.0921e-05 | 1.9935 | 3.9008e-09 | 4.0191e-05 | 1.3259 | 1.7646 | 53_1 ohm53 | 1 | 0.98806 | 7.0921e-05 | 1.9935 | 1000000 | 0.0020004 | 0.005172 | 2.046 | 53
400_1 ohm | 400 | 1Z | 43 | 9879941 | 400 | 19.73 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 5.4012e-09 | 3.7952e-09 | 2.1871 | 0.98807 | 6.1065e-05 | 2.0046 | 1.1074e-09 | 1.4709e-05 | 3.4272 | 4.1516 | 400_1 ohm400 | 1 | 0.98807 | 6.1065e-05 | 2.0046 | 1000000 | 0.0054012 | 0.0037952 | 2.1871 | 400
1000_1 ohm | 1000 | 1Z | 43 | 9879760 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 5.6012e-09 | 3.746e-09 | 2.2129 | 0.98805 | 6.0152e-05 | 2.0052 | 1.0081e-09 | 1.3976e-05 | 3.7159 | 4.3038 | 1000_1 ohm1000 | 1 | 0.98805 | 6.0152e-05 | 2.0052 | 1000000 | 0.0056012 | 0.003746 | 2.2129 | 1000
coax zero | 53 | 2Z | 100 | 10 | 53 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7736e-08 | 0.00037736 | 0 | 0 | coax zero53 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 53
coax zero | 400 | 2Z | 1 | -4 | 400 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-09 | 5e-05 | 0 | 0 | coax zero400 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 400
coax zero | 1000 | 2Z | -2 | -8 | 1000 | 19.78 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-09 | 2e-05 | 0 | 0 | coax zero1000 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1000
53_10 ohm | 53 | 2Z | 120 | 9913100 | 53 | 19.78 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 2.0004e-08 | 5.18e-08 | 2.0461 | 9.9138 | 0.00070423 | 1.9928 | 3.9016e-08 | 0.00040207 | 1.3276 | 1.7515 | 53_10 ohm53 | 1 | 9.9138 | 0.00070423 | 1.9928 | 1000000 | 0.020004 | 0.0518 | 2.0461 | 53
400_10 ohm | 400 | 2Z | 5 | 9913067 | 400 | 19.77 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 4.0009e-09 | 3.8074e-08 | 2.1873 | 9.9138 | 0.00060606 | 2.0036 | 1.1103e-08 | 0.00014752 | 3.4291 | 4.1082 | 400_10 ohm400 | 1 | 9.9138 | 0.00060606 | 2.0036 | 1000000 | 0.0040009 | 0.038074 | 2.1873 | 400
1000_10 ohm | 1000 | 2Z | 5 | 9912930 | 1000 | 19.76 | 0.5 | 10 | 0 | 0 | 10 | 0 | 3e-05 | 10 | 7.0015e-09 | 3.7584e-08 | 2.213 | 9.9137 | 0.0006031 | 2.0053 | 1.0113e-08 | 0.00014022 | 3.7163 | 4.3009 | 1000_10 ohm1000 | 1 | 9.9137 | 0.0006031 | 2.0053 | 1000000 | 0.0070015 | 0.037584 | 2.213 | 1000
coax zero | 53 | 3Z | 30 | -13 | 53 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7736e-07 | 0.0037736 | 0 | 0 | coax zero53 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 53
coax zero | 400 | 3Z | -1 | -9 | 400 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-08 | 0.0005 | 0 | 0 | coax zero400 | 1 | 0 | 0 | 0 | 1000001 | 0 | 0 | 0 | 400
coax zero | 1000 | 3Z | -1 | -7 | 1000 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-08 | 0.0002 | 0 | 0 | coax zero1000 | 1 | 0 | 0 | 0 | 1000002 | 0 | 0 | 0 | 1000
53_100 ohm | 53 | 3Z | 50 | 10002950 | 53 | 19.67 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | 2.0004e-07 | 5.2016e-07 | 2.0465 | 100.0368 | 0.0056598 | 2.0071 | 3.9039e-07 | 0.0040251 | 1.3324 | 1.4061 | 53_100 ohm53 | 1 | 100.0368 | 0.0056598 | 2.0071 | 1000003 | 0.20004 | 0.52016 | 2.0465 | 53
400_100 ohm | 400 | 3Z | 5 | 10002950 | 400 | 19.66 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | 6.0013e-08 | 3.8405e-07 | 2.188 | 100.0368 | 0.00434 | 2.029 | 1.1184e-07 | 0.0014871 | 3.4341 | 2.9184 | 400_100 ohm400 | 1 | 100.0368 | 0.00434 | 2.029 | 1000004 | 0.060013 | 0.38405 | 2.188 | 400
1000_100 ohm | 1000 | 3Z | 5 | 10002954 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | 6.0013e-08 | 3.7921e-07 | 2.2133 | 100.0368 | 0.0042807 | 2.0335 | 1.0202e-07 | 0.0014147 | 3.7171 | 3.0258 | 1000_100 ohm1000 | 1 | 100.0368 | 0.0042807 | 2.0335 | 1000005 | 0.060013 | 0.37921 | 2.2133 | 1000
coax zero | 53 | 4Z | 35 | -10 | 53 | 19.63 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7736e-06 | 0.037736 | 0 | 0 | coax zero53 | 0.001 | 0 | 0 | 0 | 1000006 | 0 | 0 | 0 | 53
coax zero | 400 | 4Z | 0 | -11 | 400 | 19.64 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-07 | 0.005 | 0 | 0 | coax zero400 | 0.001 | 0 | 0 | 0 | 1000007 | 0 | 0 | 0 | 400
coax zero | 1000 | 4Z | -1 | -11 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-07 | 0.002 | 0 | 0 | coax zero1000 | 0.001 | 0 | 0 | 0 | 1000008 | 0 | 0 | 0 | 1000
53_1000 ohm | 53 | 4Z | 10 | 10001710 | 53 | 19.63 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -2.5005e-06 | 5.2013e-06 | 2.0465 | 1000.2439 | 0.049575 | 2.031 | 3.9039e-06 | 0.04025 | 1.3323 | 1.2317 | 53_1000 ohm53 | 0.001 | 1.0002 | 4.9575e-05 | 2.031 | 1000000000000 | 2.4993 | 5.1987 | 2.0465 | 53 | Give result as parallel R and parallel C
400_1000 ohm | 400 | 4Z | 5 | 10001713 | 400 | 19.64 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | 5.0011e-07 | 3.8401e-06 | 2.188 | 1000.2443 | 0.033861 | 2.1089 | 1.1183e-06 | 0.014869 | 3.434 | 2.2772 | 400_1000 ohm400 | 0.001 | 1.0002 | 3.3861e-05 | 2.1089 | 1000000000000 | -0.49986 | 3.8382 | 2.188 | 400 | C=-L/R^2
1000_1000 ohm | 1000 | 4Z | 5 | 10001670 | 1000 | 19.65 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | 6.0013e-07 | 3.7916e-06 | 2.2133 | 1000.24 | 0.033179 | 2.1282 | 1.02e-06 | 0.014145 | 3.7171 | 2.3455 | 1000_1000 ohm1000 | 0.001 | 1.0002 | 3.3179e-05 | 2.1282 | 1000000000000 | -0.59984 | 3.7898 | 2.2133 | 1000
coax zero | 53 | 5Z | 10 | -28 | 53 | 19.49 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7736e-05 | 0.37736 | 0 | 0 | coax zero53 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 53
coax zero | 400 | 5Z | 0 | -26 | 400 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-06 | 0.05 | 0 | 0 | coax zero400 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 400
coax zero | 1000 | 5Z | 0 | -22 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-06 | 0.02 | 0 | 0 | coax zero1000 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 1000
53_10 k ohm | 53 | 5Z | -67 | 10001342 | 53 | 19.5 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -7.7016e-05 | 5.2012e-05 | 2.0465 | 10002.0891 | 0.49676 | 2.0304 | 3.9039e-05 | 0.4025 | 1.3323 | 1.2342 | 53_10 k ohm53 | 0.001 | 10.0021 | 0.00049676 | 2.0304 | 1000000000000 | 0.76984 | 0.5199 | 2.0465 | 53
400_10 k ohm | 400 | 5Z | -63 | 10001340 | 400 | 19.51 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -6.3013e-05 | 3.8399e-05 | 2.1879 | 10002.0851 | 0.3399 | 2.1058 | 1.1182e-05 | 0.14869 | 3.434 | 2.286 | 400_10 k ohm400 | 0.001 | 10.0021 | 0.0003399 | 2.1058 | 1000000000000 | 0.62987 | 0.38383 | 2.1879 | 400
1000_10 k ohm | 1000 | 5Z | -61 | 10001312 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -6.1013e-05 | 3.7915e-05 | 2.2133 | 10002.0531 | 0.33303 | 2.1245 | 1.02e-05 | 0.14145 | 3.7171 | 2.3544 | 1000_10 k ohm1000 | 0.001 | 10.0021 | 0.00033303 | 2.1245 | 1000000000000 | 0.60988 | 0.37899 | 2.2133 | 1000
coax zero | 53 | 6Z | 10 | -7 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00037736 | 3.7736 | 0 | 0 | coax zero53 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 53
coax zero | 400 | 6Z | 0 | -10 | 400 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5e-05 | 0.5 | 0 | 0 | coax zero400 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 400
coax zero | 1000 | 6Z | -1 | -10 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2e-05 | 0.2 | 0 | 0 | coax zero1000 | 0.001 | 0 | 0 | 0 | 1000000000000 | #DIV/0! | #DIV/0! | 0 | 1000
53_100 k ohm | 53 | 6Z | -800 | 10002700 | 53 | 19.54 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -0.0081012 | 0.00052012 | 2.0465 | 100027.9602 | 4.9641 | 2.0306 | 0.00039039 | 4.0251 | 1.3323 | 1.2333 | 53_100 k ohm53 | 0.001 | 100.028 | 0.0049641 | 2.0306 | 1000000000000 | 0.80967 | 0.051983 | 2.0465 | 53
400_100 k ohm | 400 | 6Z | -777 | 10002664 | 400 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -0.0077712 | 0.00038402 | 2.188 | 100027.6302 | 3.3969 | 2.1064 | 0.00011183 | 1.487 | 3.434 | 2.2845 | 400_100 k ohm400 | 0.001 | 100.0276 | 0.0033969 | 2.1064 | 1000000000000 | 0.77669 | 0.038381 | 2.188 | 400
1000_100 k ohm | 1000 | 6Z | -769 | 10002551 | 1000 | 19.52 | 0.5 | 10 | 0 | 0 | 10 | 0 | 5e-06 | 10 | -0.0076812 | 0.00037917 | 2.2133 | 100026.5002 | 3.3305 | 2.1245 | 0.00010201 | 1.4146 | 3.7171 | 2.3544 | 1000_100 k ohm1000 | 0.001 | 100.0265 | 0.0033305 | 2.1245 | 1000000000000 | 0.76771 | 0.037897 | 2.2133 | 1000
coax zero | 1600 | 4Z | -1 | -8 | 1600 | 19.56 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.25e-07 | 0.00125 | 0 | 0 | coax zero1600 | 1 | 0 | 0 | 0 | 1000024 | 0 | 0 | 0 | 1600
1600_100 micro | 1600 | 4Z | 1001 | 6324510 | 1600 | 19.56 | 0.5 | 10 | 0 | 0.0001575 | 10 | 0 | 5e-06 | 10 | 0.00010022 | 2.4068e-06 | 2.2048 | 632.4973 | 0.021162 | 2.1194 | 6.4473e-07 | 0.0089427 | 3.7331 | 2.3664 | 1600_100 micro1600 | 1 | 632.4973 | 0.021162 | 2.1194 | 1000025 | 100.2239 | 2.4069 | 2.2048 | 1600
coax zero | 1600 | 5Z | -1 | -20 | 1600 | 19.56 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.25e-06 | 0.0125 | 0 | 0 | coax zero1600 | 1 | 0 | 0 | 0 | 1000026 | 0 | 0 | 0 | 1600
1600_100 milli | 1600 | 5Z | 99862 | 2000313 | 1600 | 19.56 | 0.5 | 10 | 0 | 0.0001575 | 10 | 0 | 5e-06 | 10 | 0.099884 | 2.833e-05 | 2.0572 | 2000.4768 | 0.083822 | 2.0233 | 2.7422e-06 | 0.032255 | 10.3311 | 2.5987 | 1600_100 milli1600 | 1 | 2000.4768 | 0.083822 | 2.0233 | 1000027 | 99887.0185 | 28.3309 | 2.0572 | 1600
coax zero | 53 | 7Y | 30 | 10 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 3.7737e-15 | 8.9939e-14 | 0 | 0 | coax zero53 | 1000000000 | 0 | 0 | 0 | 1000000000000 | 0 | 0 | 0 | 53 | Y reported as nS and pF
coax zero | 400 | 7Y | 6 | 15 | 400 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 5.0001e-16 | 6.7862e-13 | 0 | 0 | coax zero400 | 1000000000 | 0 | 0 | 0 | 1000000000000 | 0 | 0 | 0 | 400
coax zero | 1000 | 7Y | 6 | 16 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 2.0001e-16 | 1.6966e-12 | 0 | 0 | coax zero1000 | 1000000000 | 0 | 0 | 0 | 1000000000000 | 0 | 0 | 0 | 1000
53_Open | 53 | 7Y | 30 | 10 | 53 | 19.55 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 3.9538e-15 | 2.0845 | 0 | 3.9533e-11 | 2.0845 | 3.7737e-15 | 8.9939e-14 | 1.0477 | 439.5477 | 53_Open53 | 1000000000 | 0 | 0.039533 | 2.0845 | 1000000000000 | 0 | 0.0039538 | 2.0845 | 53
400_Open | 400 | 7Y | 7 | 15 | 400 | 19.54 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 1.0001e-16 | 9.6671e-16 | 2.064 | 0 | 9.6657e-12 | 2.064 | 5.0002e-16 | 6.7863e-13 | 1.9334 | 14.243 | 400_Open400 | 1000000000 | 0 | 0.0096657 | 2.064 | 1000000000000 | 0.00010001 | 0.00096671 | 2.064 | 400
1000_Open | 1000 | 7Y | 6 | 16 | 1000 | 19.53 | 0.5 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 5.5285e-16 | 2.0344 | 0 | 5.5277e-12 | 2.0344 | 2.0001e-16 | 1.6966e-12 | 2.7641 | 3.2582 | 1000_Open1000 | 1000000000 | 0 | 0.0055277 | 2.0344 | 1000000000000 | 0 | 0.00055285 | 2.0344 | 1000

## pyReport
41948
[9 empty] | Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Inductance |  | Expanded uncertainty |  | Coverage factor
| Module | Frequency (Hz) | Resistance (mΩ) | Expanded Uncertainty (mΩ) | Coverage Factor | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor |  | (Hz) | (Ω) |  | (Ω) |  |  | (μH) |  | (μH)
53_Short | SHORT | 53 | 0.0030002 | 0.047927 | 2.0256 | -0.0040009 | 0.0039543 | 2.0845 | SHORT | 53 | 0.0030002 | m | 0.09319 | m | 2.4549 | -0.0040008 |  | 0.0093197 |  | 2.4475
400_Short |  | 400 | 0.0010001 | 0.030723 | 2.0622 | 0.00010002 | 0.00096683 | 2.064 |  | 400 | 0.0010001 | m | 0.014189 | m | 2.1895 | 0.00010002 |  | 0.0015565 |  | 2.1795
1000_Short |  | 1000 | -0.0050004 | 0.029909 | 2.0768 | -0.00020004 | 0.00055292 | 2.0344 |  | 1 k | -0.0050004 | m | 0.01032 | m | 2.2521 | -0.00020004 |  | 0.0012601 |  | 2.3006
53_0 ohm | 0 Ω | 53 | 3.8313 | 0.047927 | 2.0256 | 0.0010002 | 0.0039543 | 2.0845 | 0 Ω | 53 | 3.8313 | m | 0.09319 | m | 2.4549 | 0.0010002 |  | 0.0093197 |  | 2.4475
400_0 ohm |  | 400 | 3.8263 | 0.030724 | 2.0622 | 0.0081017 | 0.0009669 | 2.064 |  | 400 | 3.8263 | m | 0.014192 | m | 2.1892 | 0.0081016 |  | 0.0015565 |  | 2.1795
1000_0 ohm |  | 1000 | 3.8473 | 0.029909 | 2.0767 | 0.0078017 | 0.00055305 | 2.0343 |  | 1 k | 3.8473 | m | 0.010323 | m | 2.2513 | 0.0078016 |  | 0.0012601 |  | 2.3005
| Module | Frequency (Hz) | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor
53_0.1 ohm | 0.1 Ω | 53 | 0.10508 | 4.8599e-05 | 2.0215 | 0.010002 | 0.0039673 | 2.0822 | 0.1 Ω | 53 | 0.10508 |  | 9.3213e-05 |  | 2.4338 | 0.010002 |  | 0.0093188 |  | 2.4608
400_0.1 ohm |  | 400 | 0.10513 | 3.1706e-05 | 2.0474 | 0.0073016 | 0.0010234 | 2.0426 |  | 400 | 0.10513 |  | 1.6957e-05 |  | 2.0587 | 0.0073015 |  | 0.0013094 |  | 2.2108
1000_0.1 ohm |  | 1000 | 0.10512 | 3.0891e-05 | 2.0588 | 0.0080017 | 0.00065539 | 2.0181 |  | 1 k | 0.10512 |  | 1.3997e-05 |  | 2.0451 | 0.0080016 |  | 0.00078531 |  | 2.093
53_1 ohm | 1 Ω | 53 | 0.98806 | 7.0921e-05 | 1.9935 | 0.0020004 | 0.005172 | 2.046 | 1 Ω | 53 | 0.98806 |  | 0.00013295 |  | 2.0485 | 0.0020004 |  | 0.009459 |  | 2.3106
400_1 ohm |  | 400 | 0.98807 | 6.1065e-05 | 2.0046 | 0.0054012 | 0.0037952 | 2.1871 |  | 400 | 0.98806 |  | 0.00010854 |  | 2.0463 | 0.0054011 |  | 0.0035264 |  | 2.1192
1000_1 ohm |  | 1000 | 0.98805 | 6.0152e-05 | 2.0052 | 0.0056012 | 0.003746 | 2.2129 |  | 1 k | 0.98805 |  | 0.0001082 |  | 2.0476 | 0.0056011 |  | 0.0034293 |  | 2.1438
53_10 ohm | 10 Ω | 53 | 9.9138 | 0.00070423 | 1.9928 | 0.020004 | 0.0518 | 2.0461 | 10 Ω | 53 | 9.9138 |  | 0.0013302 |  | 2.0489 | 0.020004 |  | 0.094599 |  | 2.3103
400_10 ohm |  | 400 | 9.9138 | 0.00060606 | 2.0036 | 0.0040009 | 0.038074 | 2.1873 |  | 400 | 9.9138 |  | 0.0010864 |  | 2.0474 | 0.0040008 |  | 0.03533 |  | 2.1206
1000_10 ohm |  | 1000 | 9.9137 | 0.0006031 | 2.0053 | 0.0070015 | 0.037584 | 2.213 |  | 1 k | 9.9136 |  | 0.001083 |  | 2.0487 | 0.0070014 |  | 0.034364 |  | 2.1454
53_100 ohm | 100 Ω | 53 | 100.0368 | 0.0056598 | 2.0071 | 0.20004 | 0.52016 | 2.0465 | 100 Ω | 53 | 100.0367 |  | 0.011776 |  | 2.0994 | 0.20004 |  | 0.94646 |  | 2.3083
400_100 ohm |  | 400 | 100.0368 | 0.00434 | 2.029 | 0.060013 | 0.38405 | 2.188 |  | 400 | 100.0366 |  | 0.0088538 |  | 2.1181 | 0.060012 |  | 0.35628 |  | 2.121
1000_100 ohm |  | 1000 | 100.0368 | 0.0042807 | 2.0335 | 0.060013 | 0.37921 | 2.2133 |  | 1 k | 100.0367 |  | 0.008817 |  | 2.1221 | 0.060012 |  | 0.34672 |  | 2.1455
[9 empty] | Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Capacitance |  | Expanded uncertainty |  | Coverage factor
| Module | Frequency (Hz) | Resistance (kΩ) | Expanded Uncertainty (kΩ) | Coverage Factor | Capacitance (pF) | Expanded Uncertainty (pF) | Coverage Factor |  | (Hz) | (kΩ) |  | (kΩ) |  |  | (pF) |  | (pF)
53_1000 ohm | 1 kΩ | 53 | 1.0002 | 4.9575e-05 | 2.031 | 2.4993 | 5.1987 | 2.0465 | 1 kΩ | 53 | 1.0002 |  | 0.00011775 |  | 2.0994 | 2.4993 |  | 9.4599 |  | 2.3084
400_1000 ohm |  | 400 | 1.0002 | 3.3861e-05 | 2.1089 | -0.49986 | 3.8382 | 2.188 |  | 400 | 1.0002 |  | 8.8527e-05 |  | 2.1181 | -0.49986 |  | 3.5606 |  | 2.121
1000_1000 ohm |  | 1000 | 1.0002 | 3.3179e-05 | 2.1282 | -0.59984 | 3.7898 | 2.2133 |  | 1 k | 1.0002 |  | 8.8159e-05 |  | 2.1221 | -0.59983 |  | 3.4651 |  | 2.1455
53_10 k ohm | 10 kΩ | 53 | 10.0021 | 0.00049676 | 2.0304 | 0.76984 | 0.5199 | 2.0465 | 10 kΩ | 53 | 10.0021 |  | 0.0011775 |  | 2.0994 | 0.76983 |  | 0.94606 |  | 2.3084
400_10 k ohm |  | 400 | 10.0021 | 0.0003399 | 2.1058 | 0.62987 | 0.38383 | 2.1879 |  | 400 | 10.0021 |  | 0.00088524 |  | 2.1181 | 0.62987 |  | 0.35607 |  | 2.121
1000_10 k ohm |  | 1000 | 10.0021 | 0.00033303 | 2.1245 | 0.60988 | 0.37899 | 2.2133 |  | 1 k | 10.002 |  | 0.00088156 |  | 2.1221 | 0.60987 |  | 0.34652 |  | 2.1455
53_100 k ohm | 100 kΩ | 53 | 100.028 | 0.0049641 | 2.0306 | 0.80967 | 0.051983 | 2.0465 | 100 kΩ | 53 | 100.0277 |  | 0.011776 |  | 2.0994 | 0.80966 |  | 0.094593 |  | 2.3083
400_100 k ohm |  | 400 | 100.0276 | 0.0033969 | 2.1064 | 0.77669 | 0.038381 | 2.188 |  | 400 | 100.0274 |  | 0.0088535 |  | 2.1181 | 0.77668 |  | 0.035607 |  | 2.121
1000_100 k ohm |  | 1000 | 100.0265 | 0.0033305 | 2.1245 | 0.76771 | 0.037897 | 2.2133 |  | 1 k | 100.0263 |  | 0.0088167 |  | 2.1221 | 0.7677 |  | 0.034652 |  | 2.1455
[9 empty] | Module | Frequency | Conductance |  | Expanded uncertainty |  | Coverage factor | Capacitance |  | Expanded uncertainty |  | Coverage factor
| Module | Frequency (Hz) | Conductance (nS) | Expanded Uncertainty (nS) | Coverage Factor | Capacitance (pF) | Expanded Uncertainty (pF) | Coverage Factor |  | (Hz) | (nS) |  | (nS) |  |  | (pF) |  | (pF)
53_Open | OPEN | 53 | 0 | 0.039533 | 2.0845 | 0 | 0.0039538 | 2.0845 | OPEN | 53 | 0 |  | 0.093192 |  | 2.4687 | 0 |  | 0.0093192 |  | 2.4687
400_Open |  | 400 | 0 | 0.0096657 | 2.064 | 0.00010001 | 0.00096671 | 2.064 |  | 400 | 0 |  | 0.012366 |  | 2.4252 | 0.00010001 |  | 0.0012358 |  | 2.4236
1000_Open |  | 1000 | 0 | 0.0055277 | 2.0344 | 0 | 0.00055285 | 2.0344 |  | 1 k | 0 |  | 0.0051911 |  | 2.3215 | 0 |  | 0.00051037 |  | 2.2824
[9 empty] | Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Inductance |  | Expanded uncertainty |  | Coverage factor
| Module | Frequency (Hz) | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor |  | (kHz) | (Ω) |  | (Ω) |  |  | (H) |  | (H)
1600_100 micro | 100 μH | 1600 | 632.4973 | 0.021162 | 2.1194 | 100.2239 | 2.4069 | 2.2048 | 100 μH | 1.6 | 632.4964 |  | 0.055762 |  | 2.1218 | 100.2201 | μ | 2.1961 | μ | 2.1433
1600_100 milli | 100 mH | 1600 | 2000.4768 | 0.083822 | 2.0233 | 99887.0185 | 28.3309 | 2.0572 | 100 mH | 1.6 | 2.0005 | k | 0.00018066 | k | 2.1009 | 99.883 | m | 0.039144 | m | 2.3982

## Results
Results Using Universal Bridge
HP16074A Calibration R-L Standard
Babcock Fitzroy
2014-11-05 | Job No | 96040220 | Snumber | 21950 | [5 empty] | see p.74 of Kibble and Rayner
Item | Label | Range | a | b | Inductance |  | Resistance |  | Nominal Frequency (Hz) | Convert | want wRC<<1
[10 empty] | L to C |  | GMH51 | Temperature | -0.03
=-L/R^2, pF
0 ohm | 53_0 ohm | 1Z | 90 | 38290 | 0.0010002 | microH | 3.8313 | milliohm | 53 |  |  | 19.74 | 19.71
0 ohm | 400_0 ohm | 1Z | 70 | 38261 | 0.0081016 | microH | 3.8263 | milliohm | 400 |  |  | 19.78 | 19.75
0 ohm | 1000_0 ohm | 1Z | 65 | 38470 | 0.0078016 | microH | 3.8473 | milliohm | 1000 |  |  | 19.78 | 19.75
Short | 53_Short | 1Z | 40 | 10 | -0.0040008 | microH | 0.0030002 | milliohm | 53 |  |  | 19.73 | 19.7
Short | 400_Short | 1Z | -10 | 11 | 0.00010002 | microH | 0.0010001 | milliohm | 400 |  |  | 19.78 | 19.75
Short | 1000_Short | 1Z | -15 | -50 | -0.00020004 | microH | -0.0050004 | milliohm | 1000 |  |  | 19.76 | 19.73
coax 0 | 53_coax 0 | 1Z | 80 | -20 | 0 | microH | 0 | milliohm | 53 |  |  | 19.7 | 19.67
coax 0 | 400_coax 0 | 1Z | -11 | 1 | 0 | microH | 0 | milliohm | 400 |  |  | 19.79 | 19.76
coax 0 | 1000_coax 0 | 1Z | -13 | 0 | 0 | microH | 0 | milliohm | 1000 |  |  | 19.79 | 19.76
0.1 ohm | 53_0.1 ohm | 1Z | 180 | 1050700 | 0.010002 | microH | 105.0794 | milliohm | 53 |  |  | 19.75 | 19.72
0.1 ohm | 400_0.1 ohm | 1Z | 62 | 1051191 | 0.0073015 | microH | 105.1264 | milliohm | 400 |  |  | 19.77 | 19.74
0.1 ohm | 1000_0.1 ohm | 1Z | 67 | 1051120 | 0.0080016 | microH | 105.1194 | milliohm | 1000 |  |  | 19.78 | 19.75
1 ohm | 53_1 ohm | 1Z | 100 | 9879850 | 0.0020004 | microH | 988.0567 | milliohm | 53 |  |  | 19.76 | 19.73
1 ohm | 400_1 ohm | 1Z | 43 | 9879941 | 0.0054011 | microH | 988.0637 | milliohm | 400 |  |  | 19.76 | 19.73
1 ohm | 1000_1 ohm | 1Z | 43 | 9879760 | 0.0056011 | microH | 988.0457 | milliohm | 1000 |  |  | 19.79 | 19.76
coax 0 | 53_coax 0 | 2Z | 100 | 10 | 0 | microH | 0 | ohm | 53 |  |  | 19.8 | 19.77
coax 0 | 400_coax 0 | 2Z | 1 | -4 | 0 | microH | 0 | ohm | 400 |  |  | 19.8 | 19.77
coax 0 | 1000_coax 0 | 2Z | -2 | -8 | 0 | microH | 0 | ohm | 1000 |  |  | 19.81 | 19.78
10 ohm | 53_10 ohm | 2Z | 120 | 9913100 | 0.020004 | microH | 9.9138 | ohm | 53 |  |  | 19.81 | 19.78
10 ohm | 400_10 ohm | 2Z | 5 | 9913067 | 0.0040008 | microH | 9.9138 | ohm | 400 |  |  | 19.8 | 19.77
10 ohm | 1000_10 ohm | 2Z | 5 | 9912930 | 0.0070014 | microH | 9.9136 | ohm | 1000 |  |  | 19.79 | 19.76
coax 0 | 53_coax 0 | 3Z | 30 | -13 | 0 | milliH | 0 | ohm | 53 |  |  | 19.69 | 19.66
coax 0 | 400_coax 0 | 3Z | -1 | -9 | 0 | milliH | 0 | ohm | 400 |  |  | 19.69 | 19.66
coax 0 | 1000_coax 0 | 3Z | -1 | -7 | 0 | milliH | 0 | ohm | 1000 |  |  | 19.69 | 19.66
100 ohm | 53_100 ohm | 3Z | 50 | 10002950 | 0.00020004 | milliH | 100.0367 | ohm | 53 |  |  | 19.7 | 19.67
100 ohm | 400_100 ohm | 3Z | 5 | 10002950 | 6.0012e-05 | milliH | 100.0366 | ohm | 400 |  |  | 19.69 | 19.66
100 ohm | 1000_100 ohm | 3Z | 5 | 10002954 | 6.0012e-05 | milliH | 100.0367 | ohm | 1000 |  |  | 19.68 | 19.65
coax 0 | 53_coax 0 | 4Z | 35 | -10 | 0 | milliH | 0 | ohm | 53 |  |  | 19.66 | 19.63
coax 0 | 400_coax 0 | 4Z | 0 | -11 | 0 | milliH | 0 | ohm | 400 |  |  | 19.67 | 19.64
coax 0 | 1000_coax 0 | 4Z | -1 | -11 | 0 | milliH | 0 | ohm | 1000 |  |  | 19.68 | 19.65
1000 ohm | 53_1000 ohm | 4Z | 10 | 10001710 | -0.0025005 | milliH | 1000.2426 | ohm | 53 | 2.4993 | 8.3249e-07 | 19.66 | 19.63
1000 ohm | 400_1000 ohm | 4Z | 5 | 10001713 | 0.0005001 | milliH | 1000.243 | ohm | 400 | -0.49986 | -1.2566e-06 | 19.67 | 19.64
1000 ohm | 1000_1000 ohm | 4Z | 5 | 10001670 | 0.00060012 | milliH | 1000.2387 | ohm | 1000 | -0.59983 | -3.7698e-06 | 19.68 | 19.65
coax 0 | 53_coax 0 | 5Z | 10 | -28 | 0 | milliH | 0 | ohm | 53 |  |  | 19.52 | 19.49
coax 0 | 400_coax 0 | 5Z | 0 | -26 | 0 | milliH | 0 | ohm | 400 |  |  | 19.55 | 19.52
coax 0 | 1000_coax 0 | 5Z | 0 | -22 | 0 | milliH | 0 | ohm | 1000 |  |  | 19.55 | 19.52
10 k ohm | 53_10 k ohm | 5Z | -67 | 10001342 | -0.077015 | milliH | 10002.0756 | ohm | 53 | 0.76983 | 2.5642e-06 | 19.53 | 19.5
10 k ohm | 400_10 k ohm | 5Z | -63 | 10001340 | -0.063013 | milliH | 10002.0716 | ohm | 400 | 0.62987 | 1.5834e-05 | 19.54 | 19.51
10 k ohm | 1000_10 k ohm | 5Z | -61 | 10001312 | -0.061012 | milliH | 10002.0396 | ohm | 1000 | 0.60987 | 3.8327e-05 | 19.55 | 19.52
coax 0 | 53_coax 0 | 6Z | 10 | -7 | 0 | H | 0 | kilohm | 53 |  |  | 19.58 | 19.55
coax 0 | 400_coax 0 | 6Z | 0 | -10 | 0 | H | 0 | kilohm | 400 |  |  | 19.56 | 19.53
coax 0 | 1000_coax 0 | 6Z | -1 | -10 | 0 | H | 0 | kilohm | 1000 |  |  | 19.56 | 19.53
100 k ohm | 53_100 k ohm | 6Z | -800 | 10002700 | -0.0081011 | H | 100.0277 | kilohm | 53 | 0.80966 | 2.697e-08 | 19.57 | 19.54
100 k ohm | 400_100 k ohm | 6Z | -777 | 10002664 | -0.0077711 | H | 100.0274 | kilohm | 400 | 0.77668 | 1.9525e-07 | 19.56 | 19.53
100 k ohm | 1000_100 k ohm | 6Z | -769 | 10002551 | -0.0076811 | H | 100.0263 | kilohm | 1000 | 0.7677 | 4.8249e-07 | 19.55 | 19.52
coax 0 | 1600_coax 0 | 4Z | -1 | -8 | 0 | milliH | 0 | ohm | 1600 |  |  | 19.59 | 19.56
100 micro | 1600_100 micro | 4Z | 1001 | 6324510 | 0.10022 | milliH | 632.4964 | ohm | 1600 |  |  | 19.59 | 19.56
coax 0 | 1600_coax 0 | 5Z | -1 | -20 | 0 | milliH | 0 | ohm | 1600 |  |  | 19.59 | 19.56
100 milli | 1600_100 milli | 5Z | 99862 | 2000313 | 99.883 | milliH | 2000.4741 | ohm | 1600 |  |  | 19.59 | 19.56
coax open | 53_coax open | 7Y | 30 | 10 | 0 | pF | 0 | nS | 53 |  |  | 19.58 | 19.55
coax open | 400_coax open | 7Y | 6 | 15 | 0 | pF | 0 | nS | 400 |  |  | 19.58 | 19.55
coax open | 1000_coax open | 7Y | 6 | 16 | 0 | pF | 0 | nS | 1000 |  |  | 19.56 | 19.53
Open | 53_Open | 7Y | 30 | 10 | 0 | pF | 0 | nS | 53 |  |  | 19.58 | 19.55
Open | 400_Open | 7Y | 7 | 15 | 0.00010001 | pF | 0 | nS | 400 |  |  | 19.57 | 19.54
Open | 1000_Open | 7Y | 6 | 16 | 0 | pF | 0 | nS | 1000 |  |  | 19.56 | 19.53
19.6398
0.096619
Additional | 2014-11-07 | [5 empty] | ppm/degC
| 1600_100_milli | 5Z | 99862 | 2000313 | 19.59 |  | 315.3151
|  |  | 99847 | 2000360 | 20.11
|  |  | 99858 | 2000358 | 19.79
|  |  | 99871 | 2000366 | 19.37

## Universal Bridge
New UB calibrations need to be entered in Data block and accessed through the Date |  |  | This table is the usual form for calculating spreadsheet corrections | [6 empty] | Date | C1 | R4A | R4B | R4C | G1 | G2 | c | d
|  |  | Component | Value | Date |  |  |  | 1.6 kHz | 1985-02-14 | 32 | 30.1 | -12 | 53 | 13 | 8.7 | <10 | <10 | except for range 4, up to 44? |  |  |  | -91.32 | 0 | 0 | 0 | 0 | 0
|  |  |  | (ppm) | 2014-07-07 |  |  |  | DC for resistors | 1988-11-30 | 123.32 | 35 | -12 | #N/A | 9 | 4 | 0 | 0 |  |  |  |  | 0 | 4.9 | 0 | #N/A | -4 | -4.7 | 0
Date | Record | Comments | C_1 | 143.1 | UBIII |  |  |  | DC for resistors | 1991-03-12 | 133 | 37 | -14 | 61 | 10 | 6 | 0 | 0 |  |  |  |  | 9.68 | 6.9 | -2 | 8 | -3 | -2.7 | 0
1981-05-22 | PEL Meas V | Check Y&Z ranges 6 & 4 @ 400 Hz | R_4A | 54.1261 | UBIII |  |  |  | DC for resistors | 1995-08-22 | 127.406 | 40.34 | -17.21 | 52 | 11.21 | 6.21 | 0 | 0 |  |  |  |  | 4.086 | 10.24 | -5.21 | -1 | -1.79 | -2.49 | 0
1983-12-15 | Universal Bridge II | Full calibration of standards and of all dial ratios and performance checks. | R_4B | -9.8547 | UBIII |  |  |  | DC for resistors | 1998-11-03 | 135.27 | 43.3741 | -16.3523 | 67.0039 | 11.1844 | 6.6891 | 0 | 0 |  |  |  |  | 11.95 | 13.2741 | -4.3523 | 14.0039 | -1.8156 | -2.0109 | 0
1984-02-10 |  | completion of above | R_4C | 67.1933 | UBIII |  |  |  | DC for resistors | 2001-09-08 | 124.1565 | 46.006 | -17.3198 | 64.091 | 12.0228 | 7.4763 | 0 | 0 |  |  |  |  | 0.83653 | 15.906 | -5.3198 | 11.091 | -0.97725 | -1.2237 | 0
1988-11-30 | UB II | component values measured | G_1 | 12.9485 | UBIII |  |  |  | DC for resistors | 2004-01-22 | 131.7148 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 0 | 0 |  |  |  |  | 8.3948 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 0
1991-03-12 | UB II | component values measured | G_2 | 7.4961 | UBIII |  |  |  | DC for resistors | 2004-02-18 | 123.9 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 5.2 | 0 |  |  |  |  | 0.58 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 5.2
1991-05-16 | UB II | problem with '6' dial, repaired! | [6 empty] | DC for resistors | 2005-02-18 | 129.4493 | 46.54 | -16.96 | 61.5 | 13.84 | 9 | 5.7 | 0 |  |  |  |  | 6.1293 | 16.44 | -4.96 | 8.5 | 0.84 | 0.3 | 5.7
1991-06-18 | AI III | problem with 0.01 mF on range 5Y, lack of DC path for current pump, AVOID | Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | assume negligible c,d |  |  | DC for resistors | 2009-08-13 | 130.2987 | 50.8747 | -14.1043 | 64.5247 | 12.5073 | 7.4753 | 2 | 0 |  |  |  |  | 6.9787 | 20.7747 | -2.1043 | 11.5247 | -0.4927 | -1.2247 | 2
1995-08-22 | UB II | bridge calibration | where Z3=R+jwC | correction_c | 10.97 |  |  |  | DC for resistors | 2014-07-07 | 143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 19.78 | 24.0261 | 2.1453 | 14.1933 | -0.051497 | -1.2039 | 10.97
1998-11-03 | UB II | bridge calibration, using HP3458 as transfer instrument | Y4=(G3/G2)(aG1+ßjwC1)(1+c+jd) |  |  | (G3=1/R4*)
2001-09-08 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem fixed | where Y4=G+jwC
2004-01-22 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem REALLY fixed
2004-02-18 | UBII | remeasure capacitor, same dc resistance values as for 22/01/04 | R4A is used for ranges 5 to 1 inclusive
2005-02-18 | UBII | bridge calibration, following MSLT.E.023.004. Uncertainties have increased. | R4B is used on range 6 only
2009-08-13 | UBIII | bridge calibration, following MSLT.E.023.004/5. Uncertainties have increased slightly from 2005. | R4C is used on range 7 only | [16 empty] | AH2700A | ppm | mrad
2014-07-07 | UBIII | bridge calibration, following MSLT.E.023.006. Uncertainties unchanged from 2009. | [9 empty] | Zeros | Range | Frequency | R&F | a | b |  |  | R_4C | 15.0002 | 0.096
[13 empty] | 1Z | 53 | 1Z53 | 80 | -20
[13 empty] | 1Z | 400 | 1Z400 | -11 | 1
[13 empty] | 1Z | 1000 | 1Z1000 | -13 | 0
[13 empty] | 2Z | 53 | 2Z53 | 100 | 10
|  |  |  | Range | a correction | b correction | [6 empty] | 2Z | 400 | 2Z400 | 1 | -4
[5 empty] | (ppm) | (ppm) | [6 empty] | 2Z | 1000 | 2Z1000 | -2 | -8
|  |  |  | 1Z | 200.7 | 70.5485 | [6 empty] | 3Z | 53 | 3Z53 | 30 | -13
|  |  |  | 2Z | 200.7 | 70.5485 | [6 empty] | 3Z | 400 | 3Z400 | -1 | -9
|  |  |  | 3Z | 200.7 | 70.5485 | [6 empty] | 3Z | 1000 | 3Z1000 | -1 | -7
|  |  |  | 4Z | 200.7 | 70.5485 | [6 empty] | 4Z | 53 | 4Z53 | 35 | -10
|  |  |  | 5Z | 200.7 | 70.5485 | [6 empty] | 4Z | 400 | 4Z400 | 0 | -11
|  |  |  | 6Z | 136.7192 | 6.5677 | [6 empty] | 4Z | 1000 | 4Z1000 | -1 | -11
|  |  |  | 7Z | 213.7672 | 83.6157 | [6 empty] | 4Z | 1600 | 4Z1600 | -1 | -8
|  |  |  | 1Y | 70.5079 | -59.6436 | [6 empty] | 5Z | 53 | 5Z53 | 10 | -28
|  |  |  | 2Y | 70.5079 | -59.6436 | [6 empty] | 5Z | 400 | 5Z400 | 0 | -26
|  |  |  | 3Y | 70.5079 | -59.6436 | [6 empty] | 5Z | 1000 | 5Z1000 | 0 | -22
|  |  |  | 4Y | 70.5079 | -59.6436 | [6 empty] | 5Z | 1600 | 5Z1600 | -1 | -20
|  |  |  | 5Y | 70.5079 | -59.6436 | [6 empty] | 6Z | 53 | 6Z53 | 10 | -7
|  |  |  | 6Y | 134.4887 | 4.3372 | [6 empty] | 6Z | 400 | 6Z400 | 0 | -10
|  |  |  | 7Y | 57.4407 | -72.7108 | [6 empty] | 6Z | 1000 | 6Z1000 | -1 | -10
[13 empty] | 7Y | 53 | 7Y53 | 30 | 10 |  | Symbol font versions |  | Plain version
[5 empty] | Multiplier (C&L) | Unit Name | Unit Multiplier | Multiplier (G&R) | Unit Name | Unit Multiplier | CMC n for L | CMC m for R | 7Y | 400 | 7Y400 | 6 | 15 |  | Unit Name | Unit Name | Unit Name | Unit Name
|  |  |  | 1Z | 0.0001 | microH | 1e-06 | 0.0001 | milliohm | 0.001 | 0.2 | 0.002 | 7Y | 1000 | 7Y1000 | 6 | 16 |  | mH | mW | microH | milliohm
|  |  |  | 2Z | 0.001 | microH | 1e-06 | 1e-06 | ohm | 1 | 2 | 0.02 | [6 empty] | mH | W | microH | ohm
|  |  |  | 3Z | 1e-05 | milliH | 0.001 | 1e-05 | ohm | 1 | 20 | 0.2 | [6 empty] | mH | W | milliH | ohm
|  |  |  | 4Z | 0.0001 | milliH | 0.001 | 0.0001 | ohm | 1 | 200 | 2 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 5Z | 0.001 | milliH | 0.001 | 0.001 | ohm | 1 | 2000 | 20 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 6Z | 1e-05 | H | 1 | 1e-05 | kilohm | 1000 | 20000 | 200 | [6 empty] | H | kW | H | kilohm
|  |  |  | 7Z | 0.0001 | H | 1 | 0.0001 | kilohm | 1000 | 200000 | 2000 | [6 empty] | H | kW | H | kilohm
|  |  |  | 1Y | 0.0001 | microF | 1e-06 | 0.0001 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 2Y | 1e-05 | microF | 1e-06 | 1e-05 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 3Y | 1e-06 | microF | 1e-06 | 1e-06 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 4Y | 0.0001 | nF | 1e-09 | 0.0001 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 5Y | 1e-05 | nF | 1e-09 | 1e-05 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 6Y | 0.001 | pF | 1e-12 | 0.001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
|  |  |  | 7Y | 0.0001 | pF | 1e-12 | 0.0001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
Note on M9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on N9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on O9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on P9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on Q9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on R9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on V18: k.jones: AH2700A check of 1 M at 1.5915 k. AI VII, p.42
Note on M41: k.jones: The n/f or m/f factor for CMC calculation.

## Uncertainty
UNCERTAINTY |  |  |  |  | Information for calculating expanded uncertainties |  |  |  | Parameters dependent on component | [6 empty] | This is a data table with B9 as the column input cell. | [14 empty] | CMC |  |  |  | IANZ
GEUM uncertainty calculation template |  |  |  |  | k-factor from degrees of freedom | [10 empty] | This must be re-calculated with new data sets. | [14 empty] | L uncert = Q[0.000014L,0.001R,n/f] |  | R uncert = Q[0.000014R, 0.0000001L, m/f]
The component label is the key variable for the table |  |  |  |  | a | 1.96 |  | Model for uncertainty calculation. | Label | R alpha | u_R alpha | L/C alpha | u_L/C alpha | Aux L | Aux R | Label | L/C uncertainty | k | R uncertainty | k | L to C | Range | L/C | mult | n | R/G | mult | m | f | CMC |  | CMC/U |  | CMC/U | L |  | R
[5 empty] | b | 2.372 |  | Base on formula for Universal bridge....ie assign an uncertainty to each term. | [8 empty] | 0.0093197 | 2.4475 | 0.09319 | 2.4549 | [9 empty] | Fudge
The 95% level of confidence expanded uncertainties are calculated using the Welch-Satterthwaite |  |  |  |  | c | 1.2737 |  |  | 53_0 ohm | 0 | 0 | 0 | 0 | 5 | 40 | 53_0 ohm | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0.0010002 | 1e-06 | 0.2 | 3.8313 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 7.6625e-08 | 0.00082225
formula to combine standard uncertainties. The individual degrees of freedom are derived either |  |  |  |  | d | 7.0962 |  |  | 400_0 ohm | 0 | 0 | 0 | 0 | 5 | 40 | 400_0 ohm | 0.0015565 | 2.1795 | 0.014192 | 2.1892 |  | 1Z | 0.0081016 | 1e-06 | 0.2 | 3.8263 | 0.001 | 0.002 | 400 | 1 | 5.0001e-10 | 0.32124 | 5.0003e-06 | 0.35234 | 5.0011e-10 | 0.3213 | 7.6525e-08 | 0.0053923
from the number of measurements used to derive the uncertainty or an estimate | [8 empty] | 1000_0 ohm | 0 | 0 | 0 | 0 | 5 | 40 | 1000_0 ohm | 0.0012601 | 2.3005 | 0.010323 | 2.2513 |  | 1Z | 0.0078016 | 1e-06 | 0.2 | 3.8473 | 0.001 | 0.002 | 1000 | 1 | 2.0004e-10 | 0.15875 | 2.0007e-06 | 0.19382 | 2.0011e-10 | 0.1588 | 7.6945e-08 | 0.0074539
of the percentage uncertainty in the uncertainty. | [8 empty] | 53_Short | 0 | 0 | 0 | 0 | 5 | 40 | 53_Short | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | -0.0040008 | 1e-06 | 0.2 | 0.0030002 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7735e-09 | 0.4049 | 6.0004e-11 | 6.4389e-07
Component label | 53_Short | [7 empty] | 400_Short | 0 | 0 | 0 | 0 | 5 | 40 | 400_Short | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0.00010002 | 1e-06 | 0.2 | 0.0010001 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 2.0001e-11 | 1.4096e-06
Universal bridge reactance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | 1000_Short | 0 | 0 | 0 | 0 | 5 | 40 | 1000_Short | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | -0.00020004 | 1e-06 | 0.2 | -0.0050004 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | -1.0001e-10 | -9.6904e-06
| Standard | for | 40 | inductance dial setting |  |  |  |  | 53_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | 400_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment |  | 1000_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
ac/dc error of standards | 0.0001 | 1e-08 | 30 | 5.5556 | 1.8e-17 | 5ppm considered two sigma limit, essentially relying on dc values of resistors |  |  | 53_0.1 ohm | 0 | 50 | 0 | 0 | 2.5 | 40 | 53_0.1 ohm | 0.0093188 | 2.4608 | 0.093213 | 2.4338 |  | 1Z | 0.010002 | 1e-06 | 0.2 | 105.0794 | 0.001 | 0.002 | 53 | 1 | 3.775e-09 | 0.4051 | 3.7765e-05 | 0.40514 | 3.7737e-09 | 0.40496 | 2.1016e-06 | 0.022546
dial linearity | 1 | 1 | 30 | 5.5556 | 0.18 | 1 ppm of full scale |  |  | 400_0.1 ohm | 0 | 50 | 0 | 0 | 2.5 | 40 | 400_0.1 ohm | 0.0013094 | 2.2108 | 0.016957 | 2.0587 |  | 1Z | 0.0073015 | 1e-06 | 0.2 | 105.1264 | 0.001 | 0.002 | 400 | 1 | 5.1093e-10 | 0.39021 | 5.2121e-06 | 0.30736 | 5.001e-10 | 0.38194 | 2.1025e-06 | 0.12399
amplifier/ratio | 0.00068 | 4.624e-07 | 30 | 5.5556 | 3.8486e-14 | based on scatter of results in 1984, increased in 2004 and 2005 | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. |  | 1000_0.1 ohm | 0 | 50 | 0 | 0 | 2.5 | 40 | 1000_0.1 ohm | 0.00078531 | 2.093 | 0.013997 | 2.0451 |  | 1Z | 0.0080016 | 1e-06 | 0.2 | 105.1194 | 0.001 | 0.002 | 1000 | 1 | 2.2594e-10 | 0.28771 | 2.4831e-06 | 0.1774 | 2.0011e-10 | 0.25482 | 2.1024e-06 | 0.1502
temperature coefficient of resistors | 8e-05 | 6.4e-09 | 50 | 2 | 2.048e-17 | depends on matching of Vishay components |  |  | 53_1 ohm | -60 | 30 | 0 | 0 | 1 | 40 | 53_1 ohm | 0.009459 | 2.3106 | 0.13295 | 2.0485 |  | 1Z | 0.0020004 | 1e-06 | 0.2 | 988.0567 | 0.001 | 0.002 | 53 | 1 | 3.9008e-09 | 0.41239 | 4.0191e-05 | 0.30231 | 3.7736e-09 | 0.39894 | 1.9761e-05 | 0.14864
temperature coefficient of capacitor | 4e-05 | 1.6e-09 | 30 | 5.5556 | 4.608e-19 | 2 ppm per degree for capacitor |  |  | 400_1 ohm | -60 | 30 | 0 | 0 | 1 | 40 | 400_1 ohm | 0.0035264 | 2.1192 | 0.10854 | 2.0463 |  | 1Z | 0.0054011 | 1e-06 | 0.2 | 988.0637 | 0.001 | 0.002 | 400 | 1 | 1.1074e-09 | 0.31403 | 1.4709e-05 | 0.13552 | 5.0008e-10 | 0.14181 | 1.9761e-05 | 0.18207
component stability | 0.0002 | 4e-08 | 50 | 2 | 8e-16 | some doubt on stability of capacitor in particular, increased in 2004 | Increase from 2 to 5 ppm to account for short-term instability of 1000 pF capacitor |  | 1000_1 ohm | -60 | 30 | 0 | 0 | 1 | 40 | 1000_1 ohm | 0.0034293 | 2.1438 | 0.1082 | 2.0476 |  | 1Z | 0.0056011 | 1e-06 | 0.2 | 988.0457 | 0.001 | 0.002 | 1000 | 1 | 1.0081e-09 | 0.29396 | 1.3976e-05 | 0.12917 | 2.0008e-10 | 0.058344 | 1.9761e-05 | 0.18264
inductance error due to phase error in resistance | 1.6e-05 | 2.56e-10 | - | 13 | 5.0412e-21 | 16 ppm of resistance dial setting, see comment for resistance, reduced to 5 ppm after 2004 cal | Link to conductance setting, previously just 2 counts |  | 53_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
resolution | 37.7358 | 1423.9943 | 30 | 5.5556 | 364996.76 | this might be range and frequency dependent | Previous was 2, use latest value from MSLT.E.023.005, 200/f |  | 400_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
| 0 | 0 | 1 | 5000 | 0 |  |  |  | 1000_coax 0 | 0 | 0 | 0 | 0 | 5 | 40 | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
| Sum of variance | 1424.9943 |  | Sum of u^4/v | 364996.94 |  |  |  | 53_10 ohm | -60 | 30 | 0 | 0 |  |  | 53_10 ohm | 0.094599 | 2.3103 | 0.0013302 | 2.0489 |  | 2Z | 0.020004 | 1e-06 | 2 | 9.9138 | 1 | 0.02 | 53 | 1 | 3.9016e-08 | 0.41244 | 0.00040207 | 0.30227 | 3.7739e-09 | 0.039893 | 0.00019828 | 0.14906
| u combined (counts) | 37.7491 |  | effective v | 5.5634 |  |  |  | 400_10 ohm | -60 | 30 | 0 | 0 |  |  | 400_10 ohm | 0.03533 | 2.1206 | 0.0010864 | 2.0474 |  | 2Z | 0.0040008 | 1e-06 | 2 | 9.9138 | 1 | 0.02 | 400 | 1 | 1.1103e-08 | 0.31427 | 0.00014752 | 0.13579 | 5.0006e-10 | 0.014154 | 0.00019828 | 0.1825
|  |  |  | k | 2.4687 |  |  |  | 1000_10 ohm | -60 | 30 | 0 | 0 |  |  | 1000_10 ohm | 0.034364 | 2.1454 | 0.001083 | 2.0487 |  | 2Z | 0.0070014 | 1e-06 | 2 | 9.9136 | 1 | 0.02 | 1000 | 1 | 1.0113e-08 | 0.2943 | 0.00014022 | 0.12947 | 2.001e-10 | 0.0058229 | 0.00019827 | 0.18307
Assumptions |  |  |  | expanded u (counts) | 93.1921 |  |  |  | 53_coax 0 | 0 | 0 | 0 | 0 |  |  | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
1/ For all frequencies. Additional noise typical at 50 Hz needs to be added | [8 empty] | 400_coax 0 | 0 | 0 | 0 | 0 |  |  | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
2/ Phase angle assumed small, ie no dial interaction. | [8 empty] | 1000_coax 0 | 0 | 0 | 0 | 0 |  |  | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
3/ Consistent with 1984 calibration, the IEEE paper and recent complement checks | [8 empty] | 53_100 ohm | 0 | 5 | 0 | 0 |  |  | 53_100 ohm | 0.00094646 | 2.3083 | 0.011776 | 2.0994 |  | 3Z | 0.00020004 | 0.001 | 20 | 100.0367 | 1 | 0.2 | 53 | 1 | 3.9039e-07 | 0.41248 | 0.0040251 | 0.34181 | 3.7764e-09 | 0.00399 | 0.0020007 | 0.1699
4/ Updated following 2004/5 UB calibration, see AUDIREGC | [8 empty] | 400_100 ohm | 0 | 5 | 0 | 0 |  |  | 400_100 ohm | 0.00035628 | 2.121 | 0.0088538 | 2.1181 |  | 3Z | 6.0012e-05 | 0.001 | 20 | 100.0366 | 1 | 0.2 | 400 | 1 | 1.1184e-07 | 0.3139 | 0.0014871 | 0.16796 | 5.0084e-10 | 0.0014058 | 0.0020007 | 0.22597
[9 empty] | 1000_100 ohm | 0 | 5 | 0 | 0 |  |  | 1000_100 ohm | 0.00034672 | 2.1455 | 0.008817 | 2.1221 |  | 3Z | 6.0012e-05 | 0.001 | 20 | 100.0367 | 1 | 0.2 | 1000 | 1 | 1.0202e-07 | 0.29423 | 0.0014147 | 0.16045 | 2.0084e-10 | 0.00057926 | 0.0020007 | 0.22692
[9 empty] | 53_coax 0 | 0 | 0 | 0 | 0 |  |  | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
Universal bridge resistance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | 400_coax 0 | 0 | 0 | 0 | 0 |  |  | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
| Standard | for | 10 | conductance dial setting |  | all references to the conductance dial are now to cell D34 (previously separate lookups back to results page) |  |  | 1000_coax 0 | 0 | 0 | 0 | 0 |  |  | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | 53_1000 ohm | 0 | 5 | 0 | 0 |  |  | 53_1000 ohm | 0.0094645 | 2.3084 | 0.11775 | 2.0994 | 9.4599 | 4Z | -0.0025005 | 0.001 | 200 | 1000.2426 | 1 | 2 | 53 | 1 | 3.9039e-06 | 0.41248 | 0.04025 | 0.34183 | 3.7386e-09 | 0.00039501 | 0.020005 | 0.16989
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment |  | 400_1000 ohm | 0 | 5 | 0 | 0 |  |  | 400_1000 ohm | 0.0035623 | 2.121 | 0.088527 | 2.1181 | 3.5606 | 4Z | 0.0005001 | 0.001 | 200 | 1000.243 | 1 | 2 | 400 | 1 | 1.1183e-06 | 0.31391 | 0.014869 | 0.16796 | 5.07e-10 | 0.00014232 | 0.020005 | 0.22597
resistance error due to phase error in inductance | 0.0064 | 4.096e-05 | - | 13 | 1.2906e-10 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | Previous was 16, use latest value of 5 from MSLT.E.023.005 | Now return to 16 to stay consistent with CMCs | 1000_1000 ohm | 0 | 5 | 0 | 0 |  |  | 1000_1000 ohm | 0.0034667 | 2.1455 | 0.088159 | 2.1221 | 3.4651 | 4Z | 0.00060012 | 0.001 | 200 | 1000.2387 | 1 | 2 | 1000 | 1 | 1.02e-06 | 0.29423 | 0.014145 | 0.16045 | 2.084e-10 | 6.0114e-05 | 0.020005 | 0.22692
dial linearity | 10 | 100 | 50 | 2 | 5000 |  | Previous value was 2, now decreased to 1 ppm following MSLT.E023.005 |  | 53_coax 0 | 0 | 0 | 0 | 0 |  |  | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
ac/dc error of standards | 0.00025 | 6.25e-08 | 30 | 5.5556 | 7.0312e-16 | 5ppm considered two sigma limit, essentially relying on dc values of resistors |  |  | 400_coax 0 | 0 | 0 | 0 | 0 |  |  | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
resolution | 377.3585 | 142399.4304 | 30 | 5.5556 | 3649967600.2008 | this might be dependent on frequency and range, included below, increased from 0 to 1 in 2006, now 2000/f |  |  | 1000_coax 0 | 0 | 0 | 0 | 0 |  |  | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
amplifier/ratio | 0.00017 | 2.89e-08 | 30 | 5.5556 | 1.5034e-16 | based on scatter of results in 1984, increased following 2004 and 2005 calibration | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. |  | 53_10 k ohm | 0 | 5 | 0 | 0 |  |  | 53_10 k ohm | 0.094645 | 2.3084 | 1.1775 | 2.0994 | 0.94606 | 5Z | -0.077015 | 0.001 | 2000 | 10002.0756 | 1 | 20 | 53 | 1 | 3.9039e-05 | 0.41248 | 0.4025 | 0.34183 | 2.6954e-09 | 2.8479e-05 | 0.20004 | 0.16989
temperature coefficient of resistors | 0.0002 | 4e-08 | 50 | 2 | 8e-16 | depends on matching of Vishay components |  |  | 400_10 k ohm | 0 | 5 | 0 | 0 |  |  | 400_10 k ohm | 0.035622 | 2.121 | 0.88524 | 2.1181 | 0.35607 | 5Z | -0.063013 | 0.001 | 2000 | 10002.0716 | 1 | 20 | 400 | 1 | 1.1182e-05 | 0.31391 | 0.14869 | 0.16796 | -3.8218e-10 | -1.0729e-05 | 0.20004 | 0.22597
component stability | 0.0002 | 4e-08 | 30 | 5.5556 | 2.88e-16 | stability of resistors only |  |  | 1000_10 k ohm | 0 | 5 | 0 | 0 |  |  | 1000_10 k ohm | 0.034666 | 2.1455 | 0.88156 | 2.1221 | 0.34652 | 5Z | -0.061012 | 0.001 | 2000 | 10002.0396 | 1 | 20 | 1000 | 1 | 1.02e-05 | 0.29424 | 0.14145 | 0.16045 | -6.5417e-10 | -1.8871e-05 | 0.20004 | 0.22692
| 0 | 0 | 50 | 2 | 0 |  | Previous value was 1, now increased to 2 ppm following MSLT.E023.005 |  | 53_coax 0 | 0 | 0 | 0 | 0 |  |  | 53_coax 0 | 0.0093197 | 2.4475 | 0.09319 | 2.4549 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 53 | 1 | 3.7736e-09 | 0.4049 | 3.7736e-05 | 0.40494 | 3.7736e-09 | 0.4049 | 0 | 0
| 0 | 0 | 1 | 5000 | 0 |  |  |  | 400_coax 0 | 0 | 0 | 0 | 0 |  |  | 400_coax 0 | 0.0015565 | 2.1795 | 0.014189 | 2.1895 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 400 | 1 | 5e-10 | 0.32124 | 5e-06 | 0.35238 | 5e-10 | 0.32124 | 0 | 0
| Sum of variance | 142499.4304 |  | Sum of u^4/v | 3649972600.2008 |  |  |  | 1000_coax 0 | 0 | 0 | 0 | 0 |  |  | 1000_coax 0 | 0.0012601 | 2.3006 | 0.01032 | 2.2521 |  | 1Z | 0 | 1e-06 | 0.2 | 0 | 0.001 | 0.002 | 1000 | 1 | 2e-10 | 0.15872 | 2e-06 | 0.19379 | 2e-10 | 0.15872 | 0 | 0
| u combined (counts) | 377.491 |  | effective v | 5.5634 |  |  |  | 53_100 k ohm | 0 | 5 | 0 | 0 |  |  | 53_100 k ohm | 0.00094646 | 2.3083 | 0.011776 | 2.0994 | 0.094593 | 6Z | -0.0081011 | 1 | 20000 | 100.0277 | 1000 | 200 | 53 | 1 | 0.00039039 | 0.41248 | 4.0251 | 0.34181 | -1.0964e-07 | -0.00011584 | 2.0006 | 0.16989
|  |  |  | k | 2.4687 |  |  |  | 400_100 k ohm | 0 | 5 | 0 | 0 |  |  | 400_100 k ohm | 0.00035627 | 2.121 | 0.0088535 | 2.1181 | 0.035607 | 6Z | -0.0077711 | 1 | 20000 | 100.0274 | 1000 | 200 | 400 | 1 | 0.00011183 | 0.31389 | 1.487 | 0.16795 | -1.0829e-07 | -0.00030397 | 2.0005 | 0.22596
Assumptions |  |  |  | expanded u (counts) | 931.9214 |  |  |  | 1000_100 k ohm | 0 | 5 | 0 | 0 |  |  | 1000_100 k ohm | 0.00034671 | 2.1455 | 0.0088167 | 2.1221 | 0.034652 | 6Z | -0.0076811 | 1 | 20000 | 100.0263 | 1000 | 200 | 1000 | 1 | 0.00010201 | 0.29422 | 1.4146 | 0.16044 | -1.0733e-07 | -0.00030959 | 2.0005 | 0.2269
1/ This scatter is consistent with Thompson's phase standards. | [8 empty] | 1600_coax 0 | 0 | 0 | 0 | 0 |  |  | 1600_coax 0 | 0.00035234 | 2.2011 | 0.0037558 | 2.3462 |  | 4Z | 0 | 0.001 | 200 | 0 | 1 | 2 | 1600 | 1 | 1.25e-07 | 0.35477 | 0.00125 | 0.33282 | 1.25e-10 | 0.00035477 | 0 | 0
2/ Updated following 200/54 UB calibration see AUDIREGC | [8 empty] | 1600_100 micro | 0 | 5 | 315.3151 | 31.5315 |  |  | 1600_100 micro | 0.0021961 | 2.1433 | 0.055762 | 2.1218 |  | 4Z | 0.10022 | 0.001 | 200 | 632.4964 | 1 | 2 | 1600 | 1 | 6.4473e-07 | 0.29358 | 0.0089427 | 0.16037 | 1.5281e-09 | 0.00069581 | 0.01265 | 0.22685
[9 empty] | 1600_coax 0 | 0 | 0 | 0 | 0 |  |  | 1600_coax 0 | 0.00035234 | 2.2011 | 0.0037558 | 2.3462 |  | 4Z | 0 | 0.001 | 200 | 0 | 1 | 2 | 1600 | 1 | 1.25e-07 | 0.35477 | 0.00125 | 0.33282 | 1.25e-10 | 0.00035477 | 0 | 0
[9 empty] | 1600_100 milli | 0 | 5 | 315.3151 | 31.5315 |  |  | 1600_100 milli | 0.039144 | 2.3982 | 0.18066 | 2.1009 |  | 5Z | 99.883 | 0.001 | 2000 | 2000.4741 | 1 | 20 | 1600 | 1 | 2.7422e-06 | 0.070056 | 0.032255 | 0.17854 | 1.3985e-06 | 0.035727 | 0.040009 | 0.22146
Resistance uncertainty for 16074A |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | 53_coax open | 0 | 0 | 0 | 0 |  |  | 53_coax open | 0.0093192 | 2.4687 | 0.093192 | 2.4687 |  | 7Y | 0 | 1e-12 | 0 | 0 | 1e-09 | 0 | 53 | 1 | 3.7736e-15 | 0.40493 |  | 0
| Standard | [7 empty] | 400_coax open | 0 | 0 | 0 | 0 |  |  | 400_coax open | 0.0012358 | 2.4236 | 0.012366 | 2.4252 |  | 7Y | 0 | 1e-12 | 0 | 0 | 1e-09 | 0 | 400 | 1 | 5e-16 | 0.40459 |  | 0
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | 1000_coax open | 0 | 0 | 0 | 0 |  |  | 1000_coax open | 0.00051037 | 2.2824 | 0.0051911 | 2.3215 |  | 7Y | 0 | 1e-12 | 0 | 0 | 1e-09 | 0 | 1000 | 1 | 2e-16 | 0.39188 |  | 0
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  |  |  | 53_Open | 0 | 0 | 0 | 0 |  |  | 53_Open | 0.0093192 | 2.4687 | 0.093192 | 2.4687 |  | 7Y | 0 | 1e-12 | 0 | 0 | 1e-09 | 0 | 53 | 1 | 3.7736e-15 | 0.40493 |  | 0
ubridge uncertainty | 377.491 | 142499.4304 | - | 5.5634 | 3649972600.2008 | ubridge count uncertainty |  |  | 400_Open | 0 | 0 | 0 | 0 |  |  | 400_Open | 0.0012358 | 2.4236 | 0.012366 | 2.4252 |  | 7Y | 0.00010001 | 1e-12 | 0 | 0 | 1e-09 | 0 | 400 | 1 | 5e-16 | 0.40459 |  | 0
3 month drift in resistance | 0 | 0 | 30 | 5.5556 | 0 |  |  |  | 1000_Open | 0 | 0 | 0 | 0 |  |  | 1000_Open | 0.00051037 | 2.2824 | 0.0051911 | 2.3215 |  | 7Y | 0 | 1e-12 | 0 | 0 | 1e-09 | 0 | 1000 | 1 | 2e-16 | 0.39188 |  | 0
temperature coefficient | 0 | 0 | 20 | 12.5 | 0 | assume u_t = 0.5 degree
auxiliary balance senstivity | 40 | 1600 | 30 | 5.5556 | 460800 | in Ubridge budget | [24 empty] | C uncert = (0.2/f +22C) pF with C in microF, f in Hz
| Sum of variance | 144099.4304 |  | Sum of u^4/v | 3650433400.2008
| u combined (counts) | 379.6043 |  | effective v | 5.6883
|  |  |  | k | 2.4549
Assumptions |  |  |  | expanded u (counts) | 931.8977
1/ |  |  |  |  | 0.09319 | milliohm
Reactance uncertainty for 16074A |  |  |  |  | k=a+b/v+c/v^2+d/v^3
Standard
Source of uncertainty | Uncertainties |  |  | Degrees of
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v
ubridge uncertainty | 37.7491 | 1424.9943 | - | 5.5634 | 364996.94
temperature coefficient | 0 | 0 | 30 | 5.5556 | 0 | assume 0.5 deg at one sigma and 50 ppm for the temp coefficient
3 month drift of standard | 0 | 0 | 30 | 5.5556 | 0
auxiliary balance sensitivity | 5 | 25 | 30 | 5.5556 | 112.5 | in Ubridge budget
| Sum of variance | 1449.9943 |  | Sum of u^4/v | 365109.44
| u combined (count) | 38.0788 |  | effective v | 5.7585
|  |  |  | k | 2.4475
Assumptions |  |  |  | expanded u (count) | 93.1973
1/ |  |  |  |  | 0.0093197 | microH
Note on B18: t.stewart: Assume 0.5 degree as have below? KJ: agreed and modified
Note on B40: k.jones: say 1/f noise of 1 ppm at 100 Hz TJS: or 1000 Hz? KJ:in counts so 10x100/f
Note on B41: k.jones: was with x10 but!!! the dial setting takes care of the fact that there are 7 dials!!!!
Note on B57: k.jones: Was done in % for the Muirhead inductors.
Note on B59: k.jones: set to zero until rest of calculation is correct. Leave at zero.
Note on B60: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B72: k.jones: Was % for Muirheads
Note on B73: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B74: k.jones: set to zero for now Leave at zero.

## Report
Previous report
| 2014-11-05 | [13 empty] | 2011-05-05 | S21814
| Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Inductance |  | Expanded uncertainty |  | Coverage factor | ABS((new-old)/U) |  | Module | Frequency | Resistance |  | Uncertainty |  | Coverage | Inductance |  | Uncertainty |  | Coverage
|  | (Hz) | (Ω) |  | (Ω) |  |  | (μH) |  | (μH) |  |  | R | L |  | (Hz) | (Ω) |  | (Ω) |  | Factor | (μH) |  | (μH) |  | Factor
53_Short | SHORT | 53 | 0.0030002 | m | 0.09319 | m | 2.4549 | -0.0040008 |  | 0.0093197 |  | 2.4475 | 0.096583 | 0.3005 | SHORT | 53 | -0.0060003 | m | 0.010186 | m | 2.2464 | -0.0012002 |  | 0.015979 |  | 4.1963
400_Short |  | 400 | 0.0010001 | m | 0.014189 | m | 2.1895 | 0.00010002 |  | 0.0015565 |  | 2.1795 | 0.21144 | 0.12852 |  | 400 | -0.0020001 | m | 0.0099185 | m | 2.4012 | 0.00030005 |  | 0.0017231 |  | 2.4128
1000_Short |  | 1 k | -0.0050004 | m | 0.01032 | m | 2.2521 | -0.00020004 |  | 0.0012601 |  | 2.3006 | 0.29071 | 0.15875 |  | 1 k | -0.0020001 | m | 0.0099168 | m | 2.4045 | 0 |  | 0.0012697 |  | 2.3181
53_0 ohm | 0 Ω | 53 | 3.8313 | m | 0.09319 | m | 2.4549 | 0.0010002 |  | 0.0093197 |  | 2.4475 | 0.010212 | 0.78342 | 0 Ω | 53 | 3.8322 | m | 0.010188 | m | 2.2456 | 0.0083015 |  | 0.015979 |  | 4.1963
400_0 ohm |  | 400 | 3.8263 | m | 0.014192 | m | 2.1892 | 0.0081016 |  | 0.0015565 |  | 2.1795 | 0.70127 | 0.064388 |  | 400 | 3.8362 | m | 0.0099194 | m | 2.3996 | 0.0080014 |  | 0.0017231 |  | 2.4128
1000_0 ohm |  | 1 k | 3.8473 | m | 0.010323 | m | 2.2513 | 0.0078016 |  | 0.0012601 |  | 2.3005 | 1.5548 | 0.079527 |  | 1 k | 3.8312 | m | 0.0099176 | m | 2.4029 | 0.0077014 |  | 0.0012697 |  | 2.3181
53_0.1 ohm | 0.1 Ω | 53 | 0.10508 |  | 9.3213e-05 |  | 2.4338 | 0.010002 |  | 0.0093188 |  | 2.4608 | 6.5159 | 0.18249 | 0.1 Ω | 53 | 0.10447 |  | 1.073e-05 |  | 2.0448 | 0.0083015 |  | 0.016287 |  | 4.3048
400_0.1 ohm |  | 400 | 0.10513 |  | 1.6957e-05 |  | 2.0587 | 0.0073015 |  | 0.0013094 |  | 2.2108 | 34.6375 | 1.2984 |  | 400 | 0.10454 |  | 1.0165e-05 |  | 2.0722 | 0.0090016 |  | 0.0017048 |  | 2.9893
1000_0.1 ohm |  | 1 k | 0.10512 |  | 1.3997e-05 |  | 2.0451 | 0.0080016 |  | 0.00078531 |  | 2.093 | 41.1779 | 0.89178 |  | 1 k | 0.10454 |  | 1.0157e-05 |  | 2.0727 | 0.0073013 |  | 0.00076692 |  | 2.2592
53_1 ohm | 1 Ω | 53 | 0.98806 |  | 0.00013295 |  | 2.0485 | 0.0020004 |  | 0.009459 |  | 2.3106 | 1.1094 | 0.52868 | 1 Ω | 53 | 0.98791 |  | 0.00010391 |  | 2.0515 | 0.0070012 |  | 0.015971 |  | 4.1937
400_1 ohm |  | 400 | 0.98806 |  | 0.00010854 |  | 2.0463 | 0.0054011 |  | 0.0035264 |  | 2.1192 | 1.4511 | 0.17014 |  | 400 | 0.98791 |  | 0.00010385 |  | 2.0518 | 0.0060011 |  | 0.0016656 |  | 2.3233
1000_1 ohm |  | 1 k | 0.98805 |  | 0.0001082 |  | 2.0476 | 0.0056011 |  | 0.0034293 |  | 2.1438 | 1.4094 | 0.1167 |  | 1 k | 0.98789 |  | 0.00010384 |  | 2.0518 | 0.0052009 |  | 0.0011604 |  | 2.1048
53_10 ohm | 10 Ω | 53 | 9.9138 |  | 0.0013302 |  | 2.0489 | 0.020004 |  | 0.094599 |  | 2.3103 | 0.61297 | 0.63437 | 10Ω | 53 | 9.913 |  | 0.0010424 |  | 2.0516 | -0.040007 |  | 0.15985 |  | 4.1985
400_10 ohm |  | 400 | 9.9138 |  | 0.0010864 |  | 2.0474 | 0.0040008 |  | 0.03533 |  | 2.1206 | 0.70356 | 0.028307 |  | 400 | 9.913 |  | 0.0010419 |  | 2.0519 | 0.0050009 |  | 0.016625 |  | 2.3379
1000_10 ohm |  | 1 k | 9.9136 |  | 0.001083 |  | 2.0487 | 0.0070014 |  | 0.034364 |  | 2.1454 | 0.55708 | 0.11643 |  | 1 k | 9.913 |  | 0.0010419 |  | 2.0519 | 0.0030005 |  | 0.011499 |  | 2.1149
53_100 ohm | 100 Ω | 53 | 100.0367 |  | 0.011776 |  | 2.0994 | 0.20004 |  | 0.94646 |  | 2.3083 | 0.075964 | 0.10567 | 100 Ω | 53 | 100.0358 |  | 0.0088059 |  | 2.1234 | 0.30005 |  | 1.5978 |  | 4.1959
400_100 ohm |  | 400 | 100.0366 |  | 0.0088538 |  | 2.1181 | 0.060012 |  | 0.35628 |  | 2.121 | 0.096516 | 0.05615 |  | 400 | 100.0358 |  | 0.0087999 |  | 2.1241 | 0.040007 |  | 0.16658 |  | 2.3322
1000_100 ohm |  | 1 k | 100.0367 |  | 0.008817 |  | 2.1221 | 0.060012 |  | 0.34672 |  | 2.1455 | 0.13095 | 0.11539 |  | 1 k | 100.0355 |  | 0.0087998 |  | 2.1241 | 0.020004 |  | 0.11587 |  | 2.115
| Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Capacitance |  | Expanded uncertainty |  | Coverage factor |  |  | Module | Frequency | Resistance |  | Uncertainty |  | Coverage | Capacitance |  | Uncertainty |  | Coverage
|  | (Hz) | (kΩ) |  | (kΩ) |  |  | (pF) |  | (pF) | [5 empty] | (Hz) | (kΩ) |  | (kΩ) |  | Factor | (pF) |  | (pF) |  | Factor
53_1000 ohm | 1 kΩ | 53 | 1.0002 |  | 0.00011775 |  | 2.0994 | 2.4993 |  | 9.4599 |  | 2.3084 | 0.11587 | 0.25363 | 1 kΩ | 53 | 1.0002 | #N/A | 8.8048e-05 |  | 2.1234 | 0.099972 |  | 15.9704 |  | 4.196
400_1000 ohm |  | 400 | 1.0002 |  | 8.8527e-05 |  | 2.1181 | -0.49986 |  | 3.5606 |  | 2.121 | 0.15864 | 0.056154 |  | 400 | 1.0002 | #N/A | 8.7987e-05 |  | 2.1241 | -0.29992 |  | 1.665 |  | 2.3323
1000_1000 ohm |  | 1 k | 1.0002 |  | 8.8159e-05 |  | 2.1221 | -0.59983 |  | 3.4651 |  | 2.1455 | 0.14909 | 0.1154 |  | 1 k | 1.0002 | #N/A | 8.7986e-05 |  | 2.1241 | -0.19994 |  | 1.1581 |  | 2.115
53_10 k ohm | 10 kΩ | 53 | 10.0021 |  | 0.0011775 |  | 2.0994 | 0.76983 |  | 0.94606 |  | 2.3084 | 0.3248 | 0.052884 | 10 kΩ | 53 | 10.0017 | #N/A | 0.00088042 |  | 2.1234 | 0.81987 |  | 1.5972 |  | 4.196
400_10 k ohm |  | 400 | 10.0021 |  | 0.00088524 |  | 2.1181 | 0.62987 |  | 0.35607 |  | 2.121 | 0.42977 | 0.02817 |  | 400 | 10.0017 | #N/A | 0.00087982 |  | 2.1241 | 0.6399 |  | 0.16652 |  | 2.3324
1000_10 k ohm |  | 1 k | 10.002 |  | 0.00088156 |  | 2.1221 | 0.60987 |  | 0.34652 |  | 2.1455 | 0.41228 | 0.1155 |  | 1 k | 10.0017 | #N/A | 0.00087981 |  | 2.1241 | 0.6499 |  | 0.11581 |  | 2.115
53_100 k ohm | 100 kΩ | 53 | 100.0277 |  | 0.011776 |  | 2.0994 | 0.80966 |  | 0.094593 |  | 2.3083 | 0.25768 | 0.10538 | 100 kΩ | 53 | 100.0247 | #N/A | 0.0088055 |  | 2.1234 | 0.79969 |  | 0.1597 |  | 4.1959
400_100 k ohm |  | 400 | 100.0274 |  | 0.0088535 |  | 2.1181 | 0.77668 |  | 0.035607 |  | 2.121 | 0.33708 | 0.1411 |  | 400 | 100.0244 | #N/A | 0.0087994 |  | 2.1241 | 0.7817 |  | 0.01665 |  | 2.3323
1000_100 k ohm |  | 1 k | 100.0263 |  | 0.0088167 |  | 2.1221 | 0.7677 |  | 0.034652 |  | 2.1455 | 0.32034 | 0.23146 |  | 1 k | 100.0234 | #N/A | 0.0087993 |  | 2.1241 | 0.77572 |  | 0.011581 |  | 2.115
| Module | Frequency | Conductance |  | Expanded uncertainty |  | Coverage factor | Capacitance |  | Expanded uncertainty |  | Coverage factor |  |  | Module | Frequency | Conductance |  | Uncertainty |  | Coverage | Capacitance |  | Uncertainty |  | Coverage
|  | (Hz) | (nS) |  | (nS) |  |  | (pF) |  | (pF) | [5 empty] | (Hz) | (nS) |  | (nS) |  | Factor | (pF) |  | (pF) |  | Factor
53_Open | OPEN | 53 | 0 |  | 0.093192 |  | 2.4687 | 0 |  | 0.0093192 |  | 2.4687 | 0 | 0 | OPEN | 53 | 0 | #N/A | 0.043096 |  | 2.0182 | 0 |  | 0.16403 |  | 4.3452
400_Open |  | 400 | 0 |  | 0.012366 |  | 2.4252 | 0.00010001 |  | 0.0012358 |  | 2.4236 | 0 | 0.080923 |  | 400 | 0 | #N/A | 0.039963 |  | 3.877 | 0 |  | 0.020559 |  | 4.0319
1000_Open |  | 1 k | 0 |  | 0.0051911 |  | 2.3215 | 0 |  | 0.00051037 |  | 2.2824 | 0 | 0 |  | 1 k | 0 | #N/A | 0.042855 |  | 4.2642 | 0 |  | 0.0069789 |  | 3.1211
| Module | Frequency | Resistance |  | Expanded uncertainty |  | Coverage factor | Inductance |  | Expanded uncertainty |  | Coverage factor |  |  | Module | Frequency | Resistance |  | Uncertainty |  | Coverage | Inductance |  | Uncertainty |  | Coverage
|  | (kHz) | (Ω) |  | (Ω) |  |  | (H) |  | (H) | [5 empty] | (kHz) | (Ω) |  | (Ω) |  | Factor | (H) |  | (H) |  | Factor
1600_100 micro | 100 μH | 1.6 | 632.4964 |  | 0.055762 |  | 2.1218 | 100.2201 | μ | 2.1961 | μ | 2.1433 | 0.089579 | 0.092227 | 100 μH | 1.6 | 632.4914 |  | 0.055654 |  | 2.1238 | 100.0176 | μ | 0.74512 | μ | 2.1022
1600_100 milli | 100 mH | 1.6 | 2.0005 | k | 0.00018066 | k | 2.1009 | 99.883 | m | 0.039144 | m | 2.3982 | 0.059305 | 0.49816 | 100 mH | 1.6 | 2.0005 | k | 0.00017705 | k | 2.1184 | 99.8635 | m | 0.0056049 | m | 2.1099

## Previous reports
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44
| 2014-11-05 | [8 empty] | 2011-05-05 | S21814 | [7 empty] | 2005-02-15 | [8 empty] | 2001-12-06 | 2001 Values | [6 empty] | 1998-12-14
| Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage
|  | (Hz) | (W) | (W) | Factor | (mH) | (mH) | Factor |  |  | (Hz) | (Ω) | (Ω) | Factor | (μH) | (μH) | Factor |  |  | (Hz) | (W) | (W) | Factor | (mH) | (mH) | Factor |  | (Hz) |  |  | Factor |  |  | Factor |  |  | (Hz) | (W) | (W) | Factor | (mH) | (mH) | Factor
53_Short | SHORT | 53 | 0.0030002 | 0.09319 | 2.4549 | -0.0040008 | 0.0093197 | 2.4475 |  | SHORT | 53 | -0.0060003 | 0.010186 | 2.2464 | -0.0012002 | 0.015979 | 4.1963 |  | SHORT | 53 | -0.0020001 | 0.014746 | 2.217 | 0.0030005 | 0.0094675 | 2.1788 |  | 53 | -0.0030002 | 0.0063362 | 2.3282 | -0.012002 | 0.010411 | 1.9633 |  | SHORT | 53 | 0.001 | 0.0061738 | 2.4604 | 0.0050009 | 0.012348 | 2.4561
400_Short |  | 400 | 0.0010001 | 0.014189 | 2.1895 | 0.00010002 | 0.0015565 | 2.1795 |  |  | 400 | -0.0020001 | 0.0099185 | 2.4012 | 0.00030005 | 0.0017231 | 2.4128 |  |  | 400 | -0.0030002 | 0.0049281 | 2.2112 | 0.00010002 | 0.001229 | 2.2166 |  | 400 | -0.00010001 | 0.0034474 | 2.4085 | -0.00010002 | 0.0024696 | 2.4452 |  |  | 400 | 0.001 | 0.0024738 | 2.4185 | 0 | 0.00030983 | 2.1874
1000_Short |  | 1 k | -0.0050004 | 0.01032 | 2.2521 | -0.00020004 | 0.0012601 | 2.3006 |  |  | 1 k | -0.0020001 | 0.0099168 | 2.4045 | 0 | 0.0012697 | 2.3181 |  |  | 1 k | -0.0010001 | 0.0049213 | 2.2142 | -0.00010002 | 0.00049719 | 2.1937 |  | 1 k | -0.0030002 | 0.0036365 | 2.2364 | -0.00070011 | 0.0024697 | 2.444 |  |  | 1 k | 0.0061003 | 0.0025591 | 2.268 | -0.00030005 | 0.00030983 | 2.1874
53_0 ohm | 0 W | 53 | 3.8313 | 0.09319 | 2.4549 | 0.0010002 | 0.0093197 | 2.4475 |  | 0 Ω | 53 | 3.8322 | 0.010188 | 2.2456 | 0.0083015 | 0.015979 | 4.1963 |  | 0 W | 53 | 3.8122 | 0.014746 | 2.2168 | 0 | 0.009274 | 2.217 |  | 53 | 3.8132 | 0.0063241 | 2.3411 | 0 | 0.010091 | 1.9612 |  | 0 W | 53 | 3.8142 | 0.0061747 | 2.4425 | 0.0080014 | 0.012348 | 2.4678
400_0 ohm |  | 400 | 3.8263 | 0.014192 | 2.1892 | 0.0081016 | 0.0015565 | 2.1795 |  |  | 400 | 3.8362 | 0.0099194 | 2.3996 | 0.0080014 | 0.0017231 | 2.4128 |  |  | 400 | 3.8252 | 0.0069063 | 2.1407 | 0.0082014 | 0.0012349 | 2.2058 |  | 400 | 3.8551 | 0.0073734 | 2.2834 | 0.0072012 | 0.0024698 | 2.4411 |  |  | 400 | 3.8212 | 0.0024918 | 2.3476 | 0.0079014 | 0.00031695 | 2.1563
1000_0 ohm |  | 1 k | 3.8473 | 0.010323 | 2.2513 | 0.0078016 | 0.0012601 | 2.3005 |  |  | 1 k | 3.8312 | 0.0099176 | 2.4029 | 0.0077014 | 0.0012697 | 2.3181 |  |  | 1 k | 3.8222 | 0.0080978 | 2.1809 | 0.0074013 | 0.00049541 | 2.1985 |  | 1 k | 3.8602 | 0.006417 | 2.2497 | 0.0079013 | 0.0024697 | 2.4439 |  |  | 1 k | 3.8313 | 0.0029501 | 2.1822 | 0.0075013 | 0.0003108 | 2.182
53_0.1 ohm | 0.1 W | 53 | 0.10508 | 9.3213e-05 | 2.4338 | 0.010002 | 0.0093188 | 2.4608 |  | 0.1 Ω | 53 | 0.10447 | 1.073e-05 | 2.0448 | 0.0083015 | 0.016287 | 4.3048 |  | 0.1 W | 53 | 0.10295 | 1.5295e-05 | 2.1344 | 0.0030005 | 0.0093572 | 2.1946 |  | 53 | 0.10204 | 0.00013546 | 2.4558 | -0.0070011 | 0.010096 | 1.9612 |  | 0.1 W | 53 | 0.1014 | 0.00021712 | 2.4644 | 0.0010002 | 0.012363 | 2.4163
400_0.1 ohm |  | 400 | 0.10513 | 1.6957e-05 | 2.0587 | 0.0073015 | 0.0013094 | 2.2108 |  |  | 400 | 0.10454 | 1.0165e-05 | 2.0722 | 0.0090016 | 0.0017048 | 2.9893 |  |  | 400 | 0.10295 | 8.4941e-06 | 2.1673 | 0.0075013 | 0.0013174 | 2.0969 |  | 400 | 0.10205 | 0.00013531 | 2.4579 | 0.0078013 | 0.0024718 | 2.4131 |  |  | 400 | 0.10141 | 0.00021961 | 2.4653 | 0.0084014 | 0.00044976 | 2.0552
1000_0.1 ohm |  | 1 k | 0.10512 | 1.3997e-05 | 2.0451 | 0.0080016 | 0.00078531 | 2.093 |  |  | 1 k | 0.10454 | 1.0157e-05 | 2.0727 | 0.0073013 | 0.00076692 | 2.2592 |  |  | 1 k | 0.10295 | 8.4941e-06 | 2.1673 | 0.0073013 | 0.00083263 | 2.078 |  | 1 k | 0.10206 | 0.00013605 | 2.4581 | 0.0075012 | 0.0024715 | 2.4158 |  |  | 1 k | 0.10141 | 0.00022105 | 2.4654 | 0.0075013 | 0.00044465 | 2.0603
53_1 ohm | 1 W | 53 | 0.98806 | 0.00013295 | 2.0485 | 0.0020004 | 0.009459 | 2.3106 |  | 1 Ω | 53 | 0.98791 | 0.00010391 | 2.0515 | 0.0070012 | 0.015971 | 4.1937 |  | 1 W | 53 | 0.98745 | 8.0382e-05 | 2.3335 | 0 | 0.010636 | 2.0752 |  | 53 | 0.98738 | 7.1302e-05 | 2.0948 | -0.0050008 | 0.010562 | 1.9623 |  | 1 W | 53 | 0.98719 | 7.1115e-05 | 2.0954 | 0.0030005 | 0.012425 | 2.3552
400_1 ohm |  | 400 | 0.98806 | 0.00010854 | 2.0463 | 0.0054011 | 0.0035264 | 2.1192 |  |  | 400 | 0.98791 | 0.00010385 | 2.0518 | 0.0060011 | 0.0016656 | 2.3233 |  |  | 400 | 0.98745 | 8.1195e-05 | 2.3673 | 0.0061011 | 0.0073492 | 2.1434 |  | 400 | 0.98738 | 7.0982e-05 | 2.0966 | 0.0056009 | 0.0039214 | 2.0911 |  |  | 400 | 0.98719 | 7.0833e-05 | 2.0972 | 0.006001 | 0.0034103 | 2.1496
1000_1 ohm |  | 1 k | 0.98805 | 0.0001082 | 2.0476 | 0.0056011 | 0.0034293 | 2.1438 |  |  | 1 k | 0.98789 | 0.00010384 | 2.0518 | 0.0052009 | 0.0011604 | 2.1048 |  |  | 1 k | 0.98745 | 8.1195e-05 | 2.3673 | 0.0054009 | 0.0073288 | 2.1515 |  | 1 k | 0.98736 | 6.9917e-05 | 2.0999 | 0.0050008 | 0.0039214 | 2.0911 |  |  | 1 k | 0.98719 | 7.0651e-05 | 2.0977 | 0.0054009 | 0.0034115 | 2.1491
53_10 ohm | 10 W | 53 | 9.9138 | 0.0013302 | 2.0489 | 0.020004 | 0.094599 | 2.3103 |  | 10Ω | 53 | 9.913 | 0.0010424 | 2.0516 | -0.040007 | 0.15985 | 4.1985 |  | 10 W | 53 | 9.9107 | 0.00096579 | 2.1062 | 0.13702 | 0.1162 | 2.0425 |  | 53 | 9.9123 | 0.00065255 | 2.1419 | -0.18003 | 0.10566 | 1.9623 |  | 10 W | 53 | 9.9117 | 0.00065528 | 2.1375 | -0.10002 | 0.12412 | 2.3646
400_10 ohm |  | 400 | 9.9138 | 0.0010864 | 2.0474 | 0.0040008 | 0.03533 | 2.1206 |  |  | 400 | 9.913 | 0.0010419 | 2.0519 | 0.0050009 | 0.016625 | 2.3379 |  |  | 400 | 9.9107 | 0.00096955 | 2.1124 | 0.0070012 | 0.073761 | 2.1435 |  | 400 | 9.9124 | 0.00065479 | 2.1385 | 0.0030005 | 0.039322 | 2.0912 |  |  | 400 | 9.9117 | 0.00065467 | 2.1387 | 0.006001 | 0.03423 | 2.15
1000_10 ohm |  | 1 k | 9.9136 | 0.001083 | 2.0487 | 0.0070014 | 0.034364 | 2.1454 |  |  | 1 k | 9.913 | 0.0010419 | 2.0519 | 0.0030005 | 0.011499 | 2.1149 |  |  | 1 k | 9.9107 | 0.00096995 | 2.112 | 0.0040007 | 0.073558 | 2.1516 |  | 1 k | 9.9124 | 0.00065529 | 2.1377 | 0.0010002 | 0.039326 | 2.0911 |  |  | 1 k | 9.9117 | 0.00065432 | 2.1392 | 0.006001 | 0.03423 | 2.15
53_100 ohm | 100 W | 53 | 100.0367 | 0.011776 | 2.0994 | 0.20004 | 0.94646 | 2.3083 |  | 100 Ω | 53 | 100.0358 | 0.0088059 | 2.1234 | 0.30005 | 1.5978 | 4.1959 |  | 100 W | 53 | 100.0336 | 0.0085339 | 2.2116 | 0.41007 | 1.0677 | 2.0748 |  | 53 | 100.0136 | 0.001387 | 2.0438 | 0 | 0.5318 | 1.9684 |  | 100 W | 53 | 100.0121 | 0.0013617 | 2.0506 | 0.20003 | 0.61737 | 2.4606
400_100 ohm |  | 400 | 100.0366 | 0.0088538 | 2.1181 | 0.060012 | 0.35628 | 2.121 |  |  | 400 | 100.0358 | 0.0087999 | 2.1241 | 0.040007 | 0.16658 | 2.3322 |  |  | 400 | 100.0336 | 0.0085955 | 2.2279 | 0.06001 | 0.74437 | 2.1437 |  | 400 | 100.0132 | 0.0013723 | 2.0478 | 0.040007 | 0.24701 | 2.4335 |  |  | 400 | 100.0125 | 0.0013669 | 2.049 | 0.06001 | 0.036037 | 2.0785
1000_100 ohm |  | 1k | 100.0367 | 0.008817 | 2.1221 | 0.060012 | 0.34672 | 2.1455 |  |  | 1 k | 100.0355 | 0.0087998 | 2.1241 | 0.020004 | 0.11587 | 2.115 |  |  | 1k | 100.0333 | 0.0085865 | 2.2293 | 0.040007 | 0.74237 | 2.1516 |  | 1k | 100.013 | 0.0013726 | 2.0477 | 0.040007 | 0.24701 | 2.4338 |  |  | 1k | 100.0123 | 0.0013667 | 2.0491 | 0.050009 | 0.036124 | 2.077
| Module | Frequency | Resistance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Frequency | Resistance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage
|  | (Hz) | (kW) | (kW) | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (kΩ) | (kΩ) | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (kW) | (kW) | Factor | (pF) | (pF) | Factor |  | (Hz) |  |  | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (kW) | (kW) | Factor | (pF) | (pF) | Factor
53_1000 ohm | 1 kW | 53 | 1.0002 | 0.00011775 | 2.0994 | 2.4993 | 9.4599 | 2.3084 |  | 1 kΩ | 53 | 1.0002 | 8.8048e-05 | 2.1234 | 0.099972 | 15.9704 | 4.196 |  | 1 kW | 53 | 1.0002 | 8.1437e-05 | 2.3344 | -3.9989 | 10.6755 | 2.0749 |  | 53 | 1.0002 | 1.3989e-05 | 2.0413 | 0 | 2.9377 | 2.0679 |  | 1 kW | 53 | 1.0002 | 1.3846e-05 | 2.0447 | 0 | 2.4697 | 2.4273
400_1000 ohm |  | 400 | 1.0002 | 8.8527e-05 | 2.1181 | -0.49986 | 3.5606 | 2.121 |  |  | 400 | 1.0002 | 8.7987e-05 | 2.1241 | -0.29992 | 1.665 | 2.3323 |  |  | 400 | 1.0002 | 8.2245e-05 | 2.3674 | -0.6998 | 7.445 | 2.1433 |  | 400 | 1.0002 | 1.3903e-05 | 2.0431 | 0.59984 | 2.4691 | 2.4335 |  |  | 400 | 1.0002 | 1.3855e-05 | 2.0445 | 0.79982 | 0.3611 | 2.077
1000_1000 ohm |  | 1 k | 1.0002 | 8.8159e-05 | 2.1221 | -0.59983 | 3.4651 | 2.1455 |  |  | 1 k | 1.0002 | 8.7986e-05 | 2.1241 | -0.19994 | 1.1581 | 2.115 |  |  | 1 k | 1.0002 | 8.2244e-05 | 2.3674 | -0.39989 | 7.4246 | 2.1513 |  | 1 k | 1.0002 | 1.3908e-05 | 2.043 | 0.79979 | 2.4691 | 2.4338 |  |  | 1 k | 1.0002 | 1.3855e-05 | 2.0445 | 0.69985 | 0.35995 | 2.079
53_10 k ohm | 10 kW | 53 | 10.0021 | 0.0011775 | 2.0994 | 0.76983 | 0.94606 | 2.3084 |  | 10 kΩ | 53 | 10.0017 | 0.00088042 | 2.1234 | 0.81987 | 1.5972 | 4.196 |  | 10 kW | 53 | 10.0013 | 0.00081429 | 2.3344 | 0.69993 | 1.0736 | 2.0716 |  | 53 | 10.0011 | 0.00013827 | 2.0449 | 1.5879 | 0.33499 | 2.0474 |  | 10 kW | 53 | 10.001 | 0.00013667 | 2.0491 | 0.63801 | 0.24694 | 2.4395
400_10 k ohm |  | 400 | 10.0021 | 0.00088524 | 2.1181 | 0.62987 | 0.35607 | 2.121 |  |  | 400 | 10.0017 | 0.00087982 | 2.1241 | 0.6399 | 0.16652 | 2.3324 |  |  | 400 | 10.0013 | 0.00082237 | 2.3674 | 0.62994 | 0.75627 | 2.1261 |  | 400 | 10.0011 | 0.00013839 | 2.0446 | 1.6179 | 0.28047 | 2.1946 |  |  | 400 | 10.001 | 0.00013606 | 2.0509 | 0.70801 | 0.03603 | 2.0785
1000_10 k ohm |  | 1 k | 10.002 | 0.00088156 | 2.1221 | 0.60987 | 0.34652 | 2.1455 |  |  | 1 k | 10.0017 | 0.00087981 | 2.1241 | 0.6499 | 0.11581 | 2.115 |  |  | 1 k | 10.0013 | 0.00082236 | 2.3674 | 0.63994 | 0.75387 | 2.1334 |  | 1 k | 10.0011 | 0.00013829 | 2.0449 | 1.6179 | 0.28147 | 2.1935 |  |  | 1 k | 10.001 | 0.00013585 | 2.0516 | 0.69801 | 0.036002 | 2.079
53_100 k ohm | 100 kW | 53 | 100.0277 | 0.011776 | 2.0994 | 0.80966 | 0.094593 | 2.3083 |  | 100 kΩ | 53 | 100.0247 | 0.0088055 | 2.1234 | 0.79969 | 0.1597 | 4.1959 |  | 100 kW | 53 | 100.0222 | 0.0081744 | 2.312 | 0.80578 | 0.10666 | 2.0755 |  | 53 | 100.0186 | 0.0013841 | 2.0446 | 0.78183 | 0.029439 | 2.0667 |  | 100 kW | 53 | 100.0173 | 0.0014169 | 2.0395 | 0.7938 | 0.024689 | 2.4414
400_100 k ohm |  | 400 | 100.0274 | 0.0088535 | 2.1181 | 0.77668 | 0.035607 | 2.121 |  |  | 400 | 100.0244 | 0.0087994 | 2.1241 | 0.7817 | 0.01665 | 2.3323 |  |  | 400 | 100.0219 | 0.0082487 | 2.3423 | 0.78579 | 0.074457 | 2.1433 |  | 400 | 100.0183 | 0.0013806 | 2.0455 | 0.77184 | 0.024693 | 2.4326 |  |  | 400 | 100.0171 | 0.001416 | 2.0396 | 0.77581 | 0.0036333 | 2.0735
1000_100 k ohm |  | 1 k | 100.0263 | 0.0088167 | 2.1221 | 0.7677 | 0.034652 | 2.1455 |  |  | 1 k | 100.0234 | 0.0087993 | 2.1241 | 0.77572 | 0.011581 | 2.115 |  |  | 1 k | 100.0211 | 0.0082494 | 2.3416 | 0.77981 | 0.074259 | 2.1512 |  | 1 k | 100.0172 | 0.0013741 | 2.0473 | 0.76486 | 0.024693 | 2.4326 |  |  | 1 k | 100.0164 | 0.0013977 | 2.0422 | 0.76882 | 0.0036336 | 2.0734
| Module | Frequency | Conductance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Conductance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Conductance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Frequency | Conductance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage |  | Module | Frequency | Conductance | Uncertainty | Coverage | Capacitance | Uncertainty | Coverage
|  | (Hz) | (nS) | (nS) | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (nS) | (nS) | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (nS) | (nS) | Factor | (pF) | (pF) | Factor |  | (Hz) | (nS) | (nS) | Factor | (pF) | (pF) | Factor |  |  | (Hz) | (nS) | (nS) | Factor | (pF) | (pF) | Factor
53_Open | OPEN | 53 | 0 | 0.093192 | 2.4687 | 0 | 0.0093192 | 2.4687 |  | OPEN | 53 | 0 | 0.043096 | 2.0182 | 0 | 0.16403 | 4.3452 |  | OPEN | 53 | -0.0050001 | 0.14746 | 2.217 | -0.0050001 | 0.09279 | 2.2158 |  | 53 | 0 | 0.034467 | 2.4132 | 0 | 0.02939 | 2.0679 |  | OPEN | 53 | 0 | 0.024733 | 2.4252 | 0.0010001 | 0.024726 | 2.413
400_Open |  | 400 | 0 | 0.012366 | 2.4252 | 0.00010001 | 0.0012358 | 2.4236 |  |  | 400 | 0 | 0.039963 | 3.877 | 0 | 0.020559 | 4.0319 |  |  | 400 | -0.002 | 0.049153 | 2.217 | -0.002 | 0.01229 | 2.2166 |  | 400 | 0 | 0.034475 | 2.4079 | 0 | 0.024696 | 2.4453 |  |  | 400 | 0.001 | 0.024742 | 2.4146 | 0 | 0.0030957 | 2.189
1000_Open |  | 1 k | 0 | 0.0051911 | 2.3215 | 0 | 0.00051037 | 2.2824 |  |  | 1 k | 0 | 0.042855 | 4.2642 | 0 | 0.0069789 | 3.1211 |  |  | 1 k | -0.002 | 0.049153 | 2.217 | -0.002 | 0.0049213 | 2.2142 |  | 1 k | 0 | 0.034467 | 2.4132 | 0 | 0.024696 | 2.4453 |  |  | 1 k | 0 | 0.024733 | 2.4252 | 0 | 0.005362 | 2.189
| Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage |  | Module | Frequency | Resistance | Uncertainty | Coverage | Inductance | Uncertainty | Coverage
|  | (kHz) | (W) | (W) | Factor | (H) | (H) | Factor |  |  | (kHz) | (Ω) | (Ω) | Factor | (H) | (H) | Factor |  |  | (kHz) | (W) | (W) | Factor | (H) | (H) | Factor |  | (kHz) | (W) | (W) | Factor | (H) | (H) | Factor |  |  | (kHz) | (W) | (W) | Factor | (H) | (H) | Factor
1600_100 micro | 100 mH | 1.6 | 632.4964 | 0.055762 | 2.1218 | 100.2201 | 2.1961 | 2.1433 |  | 100 μH | 1.6 | 632.4914 | 0.055654 | 2.1238 | 100.0176 | 0.74512 | 2.1022 |  | 100 mH | 1.6 | 632.4956 | 0.052018 | 2.3415 | 100.1173 | 0.11853 | 2.1609 |  | 1.6 | 632.4884 | 0.0089215 | 2.0413 | 100.8164 | 3.0237 | 2.1146 |  | 100 mH | 1.6 | 632.4855 | 0.0087257 | 2.0462 | 101.1174 | 2.2023 | 2.14
1600_100 milli | 100 mH | 1.6 | 2.0005 | 0.00018066 | 2.1009 | 99.883 | 0.039144 | 2.3982 |  | 100 mH | 1.6 | 2.0005 | 0.00017705 | 2.1184 | 99.8635 | 0.0056049 | 2.1099 |  | 100 mH | 1.6 | 2.0004 | 0.00016063 | 2.2639 | 99.8812 | 0.015719 | 2.0882 |  | 1.6 | 2.0004 | 5.1286e-05 | 2.0384 | 99.8822 | 0.024882 | 2.3447 |  | 100 mH | 1.6 | 2.0004 | 4.6822e-05 | 2.0278 | 99.8782 | 0.0073967 | 2.0931
1600_100 milli | Resistance |  | Reactance
2014-11-05 | 2.0005 | 0.00018066 | 99.883 | 0.039144
2011-05-05 | 2.0005 | 0.00017705 | 99.8635 | 0.0056049
2005-02-15 | 2.0004 | 0.00016063 | 99.8812 | 0.015719
2001-12-06 | 2.0004 | 5.1286e-05 | 99.8822 | 0.024882
1998-12-14 | 2.0004 | 4.6822e-05 | 99.8782 | 0.0073967

## Sheet1
variability in short measurements(counts) |  | L | R
| 1 kHz | -10 | -16
|  | -21 | 54
|  | -15 | -20
|  | -14 | -27
|  | -14 | -157
|  | 4.5461 | 37.7746
|  | 3.9623 | 76.5487

## Conditions
2014-11-03 13:00 | 20.6 | 43.28
2014-11-03 13:05 | 20.69 | 42.77
2014-11-03 13:10 | 20.69 | 43.16
2014-11-03 13:15 | 20.67 | 43.06
2014-11-03 13:20 | 20.69 | 43.48
2014-11-03 13:25 | 20.61 | 43.16
2014-11-03 13:30 | 20.61 | 43.65
2014-11-03 13:35 | 20.69 | 43.21
2014-11-03 13:40 | 20.67 | 43.06
2014-11-03 13:45 | 20.67 | 43.16
2014-11-03 13:50 | 20.67 | 43.28
2014-11-03 13:55 | 20.74 | 43.09
2014-11-03 14:00 | 20.74 | 43.28
2014-11-03 14:05 | 20.76 | 43.06
2014-11-03 14:10 | 20.74 | 43.31
2014-11-03 14:15 | 20.71 | 42.94
2014-11-03 14:20 | 20.72 | 43.5
2014-11-03 14:25 | 20.73 | 43.01
2014-11-03 14:30 | 20.67 | 43.5
2014-11-03 14:35 | 20.74 | 42.94
2014-11-03 14:40 | 20.68 | 43.14
2014-11-03 14:45 | 20.71 | 43.11
2014-11-03 14:50 | 20.68 | 43.26
2014-11-03 14:55 | 20.58 | 43.11
2014-11-03 15:00 | 20.58 | 43.46
2014-11-03 15:05 | 20.6 | 43.01
2014-11-03 15:10 | 20.61 | 43.11
2014-11-03 15:15 | 20.57 | 43.26
2014-11-03 15:20 | 20.63 | 42.92
2014-11-03 15:25 | 20.61 | 43.16
2014-11-03 15:30 | 20.68 | 42.5
2014-11-03 15:35 | 20.67 | 43.23
2014-11-03 15:40 | 20.69 | 42.94
2014-11-03 15:45 | 20.71 | 42.74
2014-11-03 15:50 | 20.69 | 43.16
2014-11-03 15:55 | 20.69 | 42.96
2014-11-03 16:00 | 20.55 | 43.16
2014-11-03 16:05 | 20.58 | 42.94
2014-11-03 16:10 | 20.63 | 42.69
2014-11-03 16:15 | 20.56 | 43.01
2014-11-03 16:20 | 20.51 | 43.11
2014-11-03 16:25 | 20.61 | 42.6
2014-11-03 16:30 | 20.69 | 42.52
2014-11-03 16:35 | 20.69 | 42.65
2014-11-03 16:40 | 20.67 | 42.55
2014-11-03 16:45 | 20.63 | 42.52
2014-11-03 16:50 | 20.58 | 42.38
2014-11-03 16:55 | 20.6 | 42.52
2014-11-03 17:00 | 20.55 | 42.57
2014-11-03 17:05 | 20.51 | 42.4
2014-11-03 17:10 | 20.55 | 42.28
2014-11-03 17:15 | 20.55 | 42.18
2014-11-03 17:20 | 20.46 | 42.5
2014-11-03 17:25 | 20.61 | 42.25
2014-11-03 17:30 | 20.58 | 42.3
2014-11-03 17:35 | 20.6 | 42.28
2014-11-03 17:40 | 20.56 | 42.2
2014-11-03 17:45 | 20.62 | 42.11
2014-11-03 17:50 | 20.53 | 41.79
2014-11-03 17:55 | 20.56 | 41.91
2014-11-03 18:00 | 20.58 | 42.3
2014-11-03 18:05 | 20.56 | 41.81
2014-11-03 18:10 | 20.55 | 41.64
2014-11-03 18:15 | 20.49 | 42.15
2014-11-03 18:20 | 20.55 | 41.84
2014-11-03 18:25 | 20.53 | 41.93
2014-11-03 18:30 | 20.52 | 41.88
2014-11-03 18:35 | 20.57 | 41.91
2014-11-03 18:40 | 20.62 | 42.03
2014-11-03 18:45 | 20.62 | 42.2
2014-11-03 18:50 | 20.63 | 42.13
2014-11-03 18:55 | 20.57 | 42.01
2014-11-03 19:00 | 20.52 | 42.03
2014-11-03 19:05 | 20.51 | 41.79
2014-11-03 19:10 | 20.47 | 41.93
2014-11-03 19:15 | 20.51 | 42.23
2014-11-03 19:20 | 20.53 | 42.18
2014-11-03 19:25 | 20.56 | 42.03
2014-11-03 19:30 | 20.53 | 42.11
2014-11-03 19:35 | 20.58 | 42.01
2014-11-03 19:40 | 20.57 | 42.01
2014-11-03 19:45 | 20.6 | 42.08
2014-11-03 19:50 | 20.67 | 41.84
2014-11-03 19:55 | 20.6 | 41.81
2014-11-03 20:00 | 20.58 | 41.96
2014-11-03 20:05 | 20.57 | 42.03
2014-11-03 20:10 | 20.55 | 41.93
2014-11-03 20:15 | 20.61 | 41.93
2014-11-03 20:20 | 20.61 | 41.74
2014-11-03 20:25 | 20.6 | 41.69
2014-11-03 20:30 | 20.57 | 41.76
2014-11-03 20:35 | 20.66 | 41.86
2014-11-03 20:40 | 20.63 | 41.74
2014-11-03 20:45 | 20.68 | 41.91
2014-11-03 20:50 | 20.65 | 41.91
2014-11-03 20:55 | 20.62 | 41.86
2014-11-03 21:00 | 20.6 | 41.93
2014-11-03 21:05 | 20.6 | 42.06
2014-11-03 21:10 | 20.69 | 41.86
2014-11-03 21:15 | 20.57 | 41.91
2014-11-03 21:20 | 20.57 | 41.59
2014-11-03 21:25 | 20.55 | 41.25
2014-11-03 21:30 | 20.51 | 41.2
2014-11-03 21:35 | 20.42 | 41.34
2014-11-03 21:40 | 20.38 | 41.52
2014-11-03 21:45 | 20.47 | 41.42
2014-11-03 21:50 | 20.4 | 41.47
2014-11-03 21:55 | 20.45 | 41.47
2014-11-03 22:00 | 20.46 | 41.59
2014-11-03 22:05 | 20.52 | 41.37
2014-11-03 22:10 | 20.56 | 41.42
2014-11-03 22:15 | 20.5 | 41.42
2014-11-03 22:20 | 20.5 | 41.2
2014-11-03 22:25 | 20.5 | 41.27
2014-11-03 22:30 | 20.5 | 41.22
2014-11-03 22:35 | 20.49 | 41.37
2014-11-03 22:40 | 20.47 | 41.42
2014-11-03 22:45 | 20.44 | 41.37
2014-11-03 22:50 | 20.4 | 41.61
2014-11-03 22:55 | 20.47 | 41.42
2014-11-03 23:00 | 20.44 | 41.34
2014-11-03 23:05 | 20.41 | 41.15
2014-11-03 23:10 | 20.4 | 41.32
2014-11-03 23:15 | 20.41 | 41.1
2014-11-03 23:20 | 20.47 | 41.42
2014-11-03 23:25 | 20.49 | 41.34
2014-11-03 23:30 | 20.42 | 41.44
2014-11-03 23:35 | 20.47 | 41.49
2014-11-03 23:40 | 20.45 | 41.49
2014-11-03 23:45 | 20.49 | 41.49
2014-11-03 23:50 | 20.52 | 41.34
2014-11-03 23:55 | 20.49 | 41.74
2014-11-04 | 20.55 | 41.37
2014-11-04 00:05 | 20.5 | 41.64
2014-11-04 00:10 | 20.52 | 41.71
2014-11-04 00:15 | 20.6 | 41.32
2014-11-04 00:20 | 20.56 | 41.88
2014-11-04 00:25 | 20.67 | 41.79
2014-11-04 00:30 | 20.65 | 41.59
2014-11-04 00:35 | 20.61 | 41.49
2014-11-04 00:40 | 20.53 | 41.34
2014-11-04 00:45 | 20.56 | 41.27
2014-11-04 00:50 | 20.57 | 41.3
2014-11-04 00:55 | 20.53 | 41.57
2014-11-04 01:00 | 20.52 | 41.57
2014-11-04 01:05 | 20.55 | 41.52
2014-11-04 01:10 | 20.51 | 41.47
2014-11-04 01:15 | 20.47 | 41.59
2014-11-04 01:20 | 20.53 | 41.49
2014-11-04 01:25 | 20.6 | 41.37
2014-11-04 01:30 | 20.58 | 41.34
2014-11-04 01:35 | 20.63 | 41.22
2014-11-04 01:40 | 20.65 | 41.07
2014-11-04 01:45 | 20.65 | 41.3
2014-11-04 01:50 | 20.66 | 41.2
2014-11-04 01:55 | 20.65 | 41.12
2014-11-04 02:00 | 20.56 | 41.1
2014-11-04 02:05 | 20.47 | 41.05
2014-11-04 02:10 | 20.49 | 40.98
2014-11-04 02:15 | 20.55 | 41.07
2014-11-04 02:20 | 20.53 | 40.88
2014-11-04 02:25 | 20.5 | 40.93
2014-11-04 02:30 | 20.53 | 41.05
2014-11-04 02:35 | 20.6 | 41.05
2014-11-04 02:40 | 20.56 | 41.15
2014-11-04 02:45 | 20.58 | 41.12
2014-11-04 02:50 | 20.56 | 41.07
2014-11-04 02:55 | 20.62 | 40.85
2014-11-04 03:00 | 20.61 | 40.85
2014-11-04 03:05 | 20.65 | 40.9
2014-11-04 03:10 | 20.65 | 40.98
2014-11-04 03:15 | 20.68 | 40.85
2014-11-04 03:20 | 20.65 | 40.98
2014-11-04 03:25 | 20.62 | 40.88
2014-11-04 03:30 | 20.67 | 40.78
2014-11-04 03:35 | 20.66 | 40.81
2014-11-04 03:40 | 20.6 | 40.78
2014-11-04 03:45 | 20.61 | 40.95
2014-11-04 03:50 | 20.6 | 41
2014-11-04 03:55 | 20.61 | 40.95
2014-11-04 04:00 | 20.66 | 40.76
2014-11-04 04:05 | 20.68 | 40.66
2014-11-04 04:10 | 20.68 | 40.78
2014-11-04 04:15 | 20.69 | 40.68
2014-11-04 04:20 | 20.66 | 40.81
2014-11-04 04:25 | 20.63 | 40.78
2014-11-04 04:30 | 20.65 | 40.85
2014-11-04 04:35 | 20.6 | 40.81
2014-11-04 04:40 | 20.6 | 40.68
2014-11-04 04:45 | 20.62 | 40.73
2014-11-04 04:50 | 20.68 | 40.66
2014-11-04 04:55 | 20.69 | 40.81
2014-11-04 05:00 | 20.58 | 40.95
2014-11-04 05:05 | 20.62 | 40.73
2014-11-04 05:10 | 20.61 | 40.78
2014-11-04 05:15 | 20.57 | 40.61
2014-11-04 05:20 | 20.58 | 40.61
2014-11-04 05:25 | 20.56 | 40.66
2014-11-04 05:30 | 20.55 | 40.78
2014-11-04 05:35 | 20.6 | 40.81
2014-11-04 05:40 | 20.61 | 40.78
2014-11-04 05:45 | 20.62 | 40.73
2014-11-04 05:50 | 20.55 | 40.66
2014-11-04 05:55 | 20.57 | 40.63
2014-11-04 06:00 | 20.63 | 40.61
2014-11-04 06:05 | 20.58 | 40.85
2014-11-04 06:10 | 20.58 | 40.81
2014-11-04 06:15 | 20.58 | 40.73
2014-11-04 06:20 | 20.57 | 40.73
2014-11-04 06:25 | 20.63 | 40.61
2014-11-04 06:30 | 20.62 | 40.81
2014-11-04 06:35 | 20.61 | 40.81
2014-11-04 06:40 | 20.57 | 40.9
2014-11-04 06:45 | 20.53 | 40.83
2014-11-04 06:50 | 20.51 | 40.76
2014-11-04 06:55 | 20.58 | 40.63
2014-11-04 07:00 | 20.63 | 40.63
2014-11-04 07:05 | 20.62 | 40.56
2014-11-04 07:10 | 20.56 | 40.83
2014-11-04 07:15 | 20.62 | 40.63
2014-11-04 07:20 | 20.58 | 40.66
2014-11-04 07:25 | 20.58 | 40.58
2014-11-04 07:30 | 20.57 | 40.61
2014-11-04 07:35 | 20.56 | 40.76
2014-11-04 07:40 | 20.57 | 40.81
2014-11-04 07:45 | 20.57 | 40.88
2014-11-04 07:50 | 20.61 | 40.76
2014-11-04 07:55 | 20.57 | 40.76
2014-11-04 08:00 | 20.53 | 40.81
2014-11-04 08:05 | 20.49 | 40.81
2014-11-04 08:10 | 20.44 | 41.03
2014-11-04 08:15 | 20.5 | 40.85
2014-11-04 08:20 | 20.55 | 40.78
2014-11-04 08:25 | 20.53 | 40.66
2014-11-04 08:30 | 20.6 | 40.66
2014-11-04 08:35 | 20.56 | 40.78
2014-11-04 08:40 | 20.55 | 40.9
2014-11-04 08:45 | 20.52 | 41.03
2014-11-04 08:50 | 20.56 | 40.81
2014-11-04 08:55 | 20.61 | 40.83
2014-11-04 09:00 | 20.6 | 40.95
2014-11-04 09:05 | 20.66 | 40.81
2014-11-04 09:10 | 20.68 | 40.71
2014-11-04 09:15 | 20.65 | 40.61
2014-11-04 09:20 | 20.56 | 40.93
2014-11-04 09:25 | 20.56 | 41.03
2014-11-04 09:30 | 20.57 | 41.15
2014-11-04 09:35 | 20.58 | 40.81
2014-11-04 09:40 | 20.55 | 40.61
2014-11-04 09:45 | 20.4 | 40.9
2014-11-04 09:50 | 20.55 | 41.03
2014-11-04 09:55 | 20.52 | 41.2
2014-11-04 10:00 | 20.56 | 41.12
2014-11-04 10:05 | 20.56 | 40.98
2014-11-04 10:10 | 20.56 | 40.95
2014-11-04 10:15 | 20.56 | 41
2014-11-04 10:20 | 20.65 | 41
2014-11-04 10:25 | 20.56 | 41.12
2014-11-04 10:30 | 20.6 | 41.12
2014-11-04 10:35 | 20.55 | 41.17
2014-11-04 10:40 | 20.56 | 41.25
2014-11-04 10:45 | 20.6 | 41.1
2014-11-04 10:50 | 20.6 | 41
2014-11-04 10:55 | 20.6 | 40.98
2014-11-04 11:00 | 20.53 | 41.25
2014-11-04 11:05 | 20.52 | 41.2
2014-11-04 11:10 | 20.46 | 41.07
2014-11-04 11:15 | 20.44 | 41.03
2014-11-04 11:20 | 20.4 | 41.22
2014-11-04 11:25 | 20.41 | 41.44
2014-11-04 11:30 | 20.41 | 41.64
2014-11-04 11:35 | 20.45 | 41.69
2014-11-04 11:40 | 20.49 | 41.66
2014-11-04 11:45 | 20.44 | 41.69
2014-11-04 11:50 | 20.4 | 41.96
2014-11-04 11:55 | 20.4 | 42.11
2014-11-04 12:00 | 20.39 | 41.91
2014-11-04 12:05 | 20.38 | 41.91
2014-11-04 12:10 | 20.35 | 41.84
2014-11-04 12:15 | 20.35 | 42.15
2014-11-04 12:20 | 20.34 | 41.61
2014-11-04 12:25 | 20.29 | 41.93
2014-11-04 12:30 | 20.29 | 41.91
2014-11-04 12:35 | 20.33 | 41.79
2014-11-04 12:40 | 20.34 | 41.88
2014-11-04 12:45 | 20.34 | 41.93
2014-11-04 12:50 | 20.38 | 41.96
2014-11-04 12:55 | 20.31 | 41.86
2014-11-04 13:00 | 20.33 | 41.86
2014-11-04 13:05 | 20.3 | 41.79
2014-11-04 13:10 | 20.28 | 42.15
2014-11-04 13:15 | 20.29 | 41.88
2014-11-04 13:20 | 20.3 | 42.13
2014-11-04 13:25 | 20.33 | 41.88
2014-11-04 13:30 | 20.4 | 41.98
2014-11-04 13:35 | 20.42 | 41.76
2014-11-04 13:40 | 20.41 | 41.98
2014-11-04 13:45 | 20.4 | 41.86
2014-11-04 13:50 | 20.41 | 42.03
2014-11-04 13:55 | 20.39 | 41.93
2014-11-04 14:00 | 20.42 | 41.96
2014-11-04 14:05 | 20.42 | 41.91
2014-11-04 14:10 | 20.41 | 42.25
2014-11-04 14:15 | 20.42 | 41.86
2014-11-04 14:20 | 20.45 | 42.03
2014-11-04 14:25 | 20.4 | 42.03
2014-11-04 14:30 | 20.39 | 41.88
2014-11-04 14:35 | 20.42 | 41.81
2014-11-04 14:40 | 20.44 | 41.74
2014-11-04 14:45 | 20.4 | 41.84
2014-11-04 14:50 | 20.35 | 41.74
2014-11-04 14:55 | 20.31 | 41.88
2014-11-04 15:00 | 20.41 | 41.88
2014-11-04 15:05 | 20.31 | 42.15
2014-11-04 15:10 | 20.36 | 41.93
2014-11-04 15:15 | 20.3 | 41.81
2014-11-04 15:20 | 20.27 | 41.91
2014-11-04 15:25 | 20.35 | 42.15
2014-11-04 15:30 | 20.42 | 42.18
2014-11-04 15:35 | 20.4 | 42.01
2014-11-04 15:40 | 20.44 | 41.88
2014-11-04 15:45 | 20.52 | 42.06
2014-11-04 15:50 | 20.49 | 42.33
2014-11-04 15:55 | 20.4 | 42.28
2014-11-04 16:00 | 20.39 | 42.18
2014-11-04 16:05 | 20.39 | 42.42
2014-11-04 16:10 | 20.39 | 42.25
2014-11-04 16:15 | 20.42 | 42.08
2014-11-04 16:20 | 20.4 | 42.25
2014-11-04 16:25 | 20.39 | 42.18
2014-11-04 16:30 | 20.4 | 42.08
2014-11-04 16:35 | 20.39 | 42.06
2014-11-04 16:40 | 20.4 | 42.01
2014-11-04 16:45 | 20.38 | 42.18
2014-11-04 16:50 | 20.39 | 42.2
2014-11-04 16:55 | 20.41 | 42.03
2014-11-04 17:00 | 20.33 | 42.25
41.6641
0.75659