## Overview
In the last report (S21959, 2015) both the AH2700 and the UB were used.
Uncertainty has a formula for series resistance impact on loss angle that should be done in Python.
The final report used UB results for capacitance, but AH2700A results for tan delta up to step 10 on the 100 pF dial.
The two-terminal nature of the box means we stick with about a 1 pF uncertainty which is not challenging for the UB.
There is a 'coax zero', a 'copper jig zero' and a 'box zero' to consider.
The 'box zero' is what should be subtracted off the change values.
The 'copper jig zero' is what is subtracted to give the quoted 'box zero'
The 'box zero' readings will depend on range for the UB while the quoted 'box zero' - 'jig zero' is done on range 6.
It seems sensible to start with the usual UB style template but to also have an AH2700 readings sheet that can be read in.
Perhaps we should start with an AH2700A class structured like the UB class and do a parallel calibration.
Relying on pasting the different result calculations would be OK for now.
Ultimately it can be put together into a generalised UUT class that pulls in both the UB and AH2700 classes.
Expect that AH2700A results will be computer collected so that they come with a standard deviation.
It is good practice to check the coax zero on all ranges and results have this coax zero subtracted
This is only essential if the measured component is coaxially connected (e.g. the cross check with GR1000 set)
When only the box zero is being subtracted we can use dummy values for the coax zero.
The box zero is measured with the copper zero subtracted.

## pyUBreadings
Template for capacitors
Capacitors |  |  |  | Client | [5 empty] | Enter, say, 5 ppm as 5e-6
Date |  |  |  | Job number | [5 empty] | 1 sigma uncertainties from specifications
|  |  |  | S Number | [5 empty] | specs say -140 ppm for capacitance, assume 100 ppm for G
|  |  | Data from lab book |  |  |  | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff C | u | df | Temp coeff G | u | df | Other | u | df
| Hz |  |  |  | Hz | °C
coax zero | 1591.5 | 3Y | 0 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
box_zero | 1591.5 | 3Y | 49 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A1 | 1591.5 | 3Y | 100013 | 40 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A2 | 1591.5 | 3Y | 200413 | 90 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A3 | 1591.5 | 3Y | 300376 | 119 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A4 | 1591.5 | 3Y | 400767 | 209 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A5 | 1591.5 | 3Y | 500576 | 346 | 1591.5 | 19.65 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A6 | 1591.5 | 3Y | 600560 | 84 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A7 | 1591.5 | 3Y | 700940 | 522 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A8 | 1591.5 | 3Y | 800803 | 736 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A9 | 1591.5 | 3Y | 901295 | 1013 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
A10 | 1591.5 | 3Y | 1001258 | 1327 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
coax zero | 1591.5 | 4Y | 0 | 0 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
box_zero | 1591.5 | 4Y | 500 | 10 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B1 | 1591.5 | 4Y | 100618 | 47 | 1591.5 | 19.58 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B2 | 1591.5 | 4Y | 200526 | 67 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B3 | 1591.5 | 4Y | 300641 | 70 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B4 | 1591.5 | 4Y | 400667 | 64 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B5 | 1591.5 | 4Y | 500749 | 57 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B6 | 1591.5 | 4Y | 600867 | -61 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B7 | 1591.5 | 4Y | 700768 | 99 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B8 | 1591.5 | 4Y | 800879 | 189 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B9 | 1591.5 | 4Y | 900901 | 329 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
B10 | 1591.5 | 4Y | 1001011 | 514 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
box_zero | 1591.5 | 5Y | 5001 | 123 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C1 | 1591.5 | 5Y | 104502 | 163 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C2 | 1591.5 | 5Y | 204731 | 157 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C3 | 1591.5 | 5Y | 304202 | 160 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C4 | 1591.5 | 5Y | 404451 | 146 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C5 | 1591.5 | 5Y | 504786 | 116 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C6 | 1591.5 | 5Y | 604205 | 67 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C7 | 1591.5 | 5Y | 704439 | 137 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C8 | 1591.5 | 5Y | 803838 | 238 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C9 | 1591.5 | 5Y | 904057 | 329 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
C10 | 1591.5 | 5Y | 1003463 | 493 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
box_zero | 1591.5 | 6Y | 50024 | 1095 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D1 | 1591.5 | 6Y | 146641 | 1150 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D2 | 1591.5 | 6Y | 245729 | 1138 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D3 | 1591.5 | 6Y | 342214 | 1170 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D4 | 1591.5 | 6Y | 442013 | 1123 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D5 | 1591.5 | 6Y | 538500 | 1140 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D6 | 1591.5 | 6Y | 641458 | 1058 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D7 | 1591.5 | 6Y | 740537 | 1164 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D8 | 1591.5 | 6Y | 837056 | 1308 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D9 | 1591.5 | 6Y | 936847 | 1426 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
D10 | 1591.5 | 6Y | 1033147 | 1534 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
coax zero | 1591.5 | 6Y | 0 | 17 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
GR1000A | 1591.5 | 6Y | 999924 | -113 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
GR1000B | 1591.5 | 6Y | 999907 | -83 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E1 | 1591.5 | 6Y | 60487 | 1341 | 1591.7 | 19.67 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E2 | 1591.5 | 6Y | 71300 | 1391 | 1591.7 | 19.65 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E3 | 1591.5 | 6Y | 81517 | 1332 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E4 | 1591.5 | 6Y | 91813 | 1324 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E5 | 1591.5 | 6Y | 101710 | 1334 | 1591.7 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E6 | 1591.5 | 6Y | 111557 | 1338 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E7 | 1591.5 | 6Y | 120990 | 1319 | 1591.7 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E8 | 1591.5 | 6Y | 130672 | 1315 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E9 | 1591.5 | 6Y | 140694 | 1310 | 1591.7 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
E10 | 1591.5 | 6Y | 150067 | 1304 | 1591.7 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10
CuZero | 1591.5 | 6Y | 888 | 15 | 1591.6 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | td | utd | ktd | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for G | G | uG | k | multiplier for C | C | uC | k | nom Freq
coax zero | 1591.5 | 3Y | 0 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 3.2278e-12 | 2.0895 | 0 | 3.2274e-08 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-05 | 2.0895 | 1000000 | 0 | 3.2278e-06 | 2.0895 | 1591.5
box_zero | 1591.5 | 3Y | 49 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 0 | 0 | 0 | 2.9379e-15 | 2.4469 | 0 | 0 | 0 | 1.2038e-15 | 1.5931e-11 | 0 | 0.00018442 | box_zero1591.5 | 1000 | 0 | 2.9379e-12 | 2.4469 | 1000000 | 0 | 0 | 0 | 1591.5
A1 | 1591.5 | 3Y | 100013 | 40 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.9974e-08 | 1.4252e-11 | 2.0519 | 3.9998e-08 | 5.6572e-08 | 2.0154 | 4.001e-05 | 5.6589e-05 | 2.0154 | 2.2006e-12 | 2.7008e-08 | 6.4765 | 2.0946 | A11591.5 | 1000 | 3.9998e-05 | 5.6572e-05 | 2.0154 | 1000000 | 0.099974 | 1.4252e-05 | 2.0519 | 1591.5
A2 | 1591.5 | 3Y | 200413 | 90 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 2.0038e-07 | 2.7687e-11 | 2.0673 | 8.9997e-08 | 9.2767e-08 | 2.0596 | 4.4914e-05 | 4.6296e-05 | 2.0596 | 4.4096e-12 | 5.4118e-08 | 6.2788 | 1.7142 | A21591.5 | 1000 | 8.9997e-05 | 9.2767e-05 | 2.0596 | 1000000 | 0.20038 | 2.7687e-05 | 2.0673 | 1591.5
A3 | 1591.5 | 3Y | 300376 | 119 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 3.0036e-07 | 4.2323e-11 | 2.0745 | 1.19e-07 | 1.4934e-07 | 2.0927 | 3.9619e-05 | 4.9721e-05 | 2.0927 | 6.609e-12 | 8.1109e-08 | 6.4038 | 1.8412 | A31591.5 | 1000 | 0.000119 | 0.00014934 | 2.0927 | 1000000 | 0.30036 | 4.2323e-05 | 2.0745 | 1591.5
A4 | 1591.5 | 3Y | 400767 | 209 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.0076e-07 | 5.5875e-11 | 2.0744 | 2.0899e-07 | 2.2923e-07 | 2.1357 | 5.2151e-05 | 5.7201e-05 | 2.1357 | 8.8178e-12 | 1.0822e-07 | 6.3365 | 2.1182 | A41591.5 | 1000 | 0.00020899 | 0.00022923 | 2.1357 | 1000000 | 0.40076 | 5.5875e-05 | 2.0744 | 1591.5
A5 | 1591.5 | 3Y | 500576 | 346 | 1591.5 | 19.65 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 5.0057e-07 | 6.9133e-11 | 2.0737 | 3.4599e-07 | 3.3355e-07 | 2.1832 | 6.912e-05 | 6.6636e-05 | 2.1832 | 1.1014e-11 | 1.3517e-07 | 6.277 | 2.4677 | A51591.5 | 1000 | 0.00034599 | 0.00033355 | 2.1832 | 1000000 | 0.50057 | 6.9133e-05 | 2.0737 | 1591.5
A6 | 1591.5 | 3Y | 600560 | 84 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 6.0057e-07 | 8.3613e-11 | 2.0753 | 8.3997e-08 | 4.6353e-07 | 2.2271 | 1.3987e-05 | 7.7184e-05 | 2.2271 | 1.3214e-11 | 1.6216e-07 | 6.3278 | 2.8584 | A61591.5 | 1000 | 8.3997e-05 | 0.00046353 | 2.2271 | 1000000 | 0.60057 | 8.3613e-05 | 2.0753 | 1591.5
A7 | 1591.5 | 3Y | 700940 | 522 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 7.0096e-07 | 9.756e-11 | 2.0754 | 5.2198e-07 | 6.1965e-07 | 2.2643 | 7.4469e-05 | 8.8403e-05 | 2.2643 | 1.5422e-11 | 1.8927e-07 | 6.3259 | 3.2739 | A71591.5 | 1000 | 0.00052198 | 0.00061965 | 2.2643 | 1000000 | 0.70096 | 9.756e-05 | 2.0754 | 1591.5
A8 | 1591.5 | 3Y | 800803 | 736 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.0083e-07 | 1.1144e-10 | 2.0756 | 7.3597e-07 | 8.0026e-07 | 2.2944 | 9.1904e-05 | 9.9932e-05 | 2.2944 | 1.7619e-11 | 2.1623e-07 | 6.3247 | 3.7009 | A81591.5 | 1000 | 0.00073597 | 0.00080026 | 2.2944 | 1000000 | 0.80083 | 0.00011144 | 2.0756 | 1591.5
A9 | 1591.5 | 3Y | 901295 | 1013 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.0133e-07 | 1.265e-10 | 2.0769 | 1.013e-06 | 1.0073e-06 | 2.3187 | 0.00011239 | 0.00011176 | 2.3187 | 1.9831e-11 | 2.4337e-07 | 6.379 | 4.139 | A91591.5 | 1000 | 0.001013 | 0.0010073 | 2.3187 | 1000000 | 0.90133 | 0.0001265 | 2.0769 | 1591.5
A10 | 1591.5 | 3Y | 1001258 | 1327 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0013e-06 | 1.393e-10 | 2.0757 | 1.3269e-06 | 1.2382e-06 | 2.3381 | 0.00013253 | 0.00012367 | 2.3381 | 2.203e-11 | 2.7036e-07 | 6.3233 | 4.58 | A101591.5 | 1000 | 0.0013269 | 0.0012382 | 2.3381 | 1000000 | 1.0013 | 0.0001393 | 2.0757 | 1591.5
coax zero | 1591.5 | 4Y | 0 | 0 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 3.2278e-13 | 2.0895 | 0 | 3.2274e-09 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-06 | 2.0895 | 1000000000 | 0 | 0.00032278 | 2.0895 | 1591.5
box_zero | 1591.5 | 4Y | 500 | 10 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 0 | 0 | 0 | 3.059e-15 | 2.4469 | 0 | 0 | 0 | 1.2258e-15 | 1.6201e-11 | 0 | 0.00018882 | box_zero1591.5 | 1000 | 0 | 3.059e-12 | 2.4469 | 1000000000 | 0 | 0 | 0 | 1591.5
B1 | 1591.5 | 4Y | 100618 | 47 | 1591.5 | 19.58 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0013e-08 | 1.5272e-12 | 2.0632 | 3.6999e-09 | 5.5802e-09 | 2.0188 | 3.6953e-05 | 5.5732e-05 | 2.0188 | 2.2151e-13 | 2.7196e-09 | 6.8947 | 2.0519 | B11591.5 | 1000 | 3.6999e-06 | 5.5802e-06 | 2.0188 | 1000000000 | 10.0128 | 0.0015272 | 2.0632 | 1591.5
B2 | 1591.5 | 4Y | 200526 | 67 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 2.0005e-08 | 2.914e-12 | 2.0754 | 5.6998e-09 | 8.4289e-09 | 2.0907 | 2.8493e-05 | 4.2136e-05 | 2.0907 | 4.4133e-13 | 5.4173e-09 | 6.603 | 1.5559 | B21591.5 | 1000 | 5.6998e-06 | 8.4289e-06 | 2.0907 | 1000000000 | 20.0045 | 0.002914 | 2.0754 | 1591.5
B3 | 1591.5 | 4Y | 300641 | 70 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 3.0017e-08 | 4.3474e-12 | 2.0786 | 5.9998e-09 | 1.1882e-08 | 2.1432 | 1.9989e-05 | 3.9587e-05 | 2.1432 | 6.616e-13 | 8.1205e-09 | 6.5711 | 1.4632 | B31591.5 | 1000 | 5.9998e-06 | 1.1882e-05 | 2.1432 | 1000000000 | 30.017 | 0.0043474 | 2.0786 | 1591.5
B4 | 1591.5 | 4Y | 400667 | 64 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.0021e-08 | 5.784e-12 | 2.0798 | 5.3998e-09 | 1.5527e-08 | 2.1698 | 1.3493e-05 | 3.88e-05 | 2.1698 | 8.8168e-13 | 1.0821e-08 | 6.5602 | 1.4349 | B41591.5 | 1000 | 5.3998e-06 | 1.5527e-05 | 2.1698 | 1000000000 | 40.0205 | 0.005784 | 2.0798 | 1591.5
B5 | 1591.5 | 4Y | 500749 | 57 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 5.003e-08 | 7.0958e-12 | 2.0776 | 4.6998e-09 | 1.9262e-08 | 2.1823 | 9.3944e-06 | 3.8503e-05 | 2.1823 | 1.1019e-12 | 1.3524e-08 | 6.4397 | 1.4243 | B51591.5 | 1000 | 4.6998e-06 | 1.9262e-05 | 2.1823 | 1000000000 | 50.0297 | 0.0070958 | 2.0776 | 1591.5
B6 | 1591.5 | 4Y | 600867 | -61 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 6.0042e-08 | 8.6638e-12 | 2.0806 | -7.0997e-09 | 2.3056e-08 | 2.1873 | -1.1825e-05 | 3.8401e-05 | 2.1873 | 1.3222e-12 | 1.6227e-08 | 6.5528 | 1.4208 | B61591.5 | 1000 | -7.0997e-06 | 2.3056e-05 | 2.1873 | 1000000000 | 60.0424 | 0.0086638 | 2.0806 | 1591.5
B7 | 1591.5 | 4Y | 700768 | 99 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 7.0033e-08 | 1.0012e-11 | 2.0794 | 8.8997e-09 | 2.689e-08 | 2.188 | 1.2708e-05 | 3.8397e-05 | 2.188 | 1.542e-12 | 1.8925e-08 | 6.4929 | 1.4209 | B71591.5 | 1000 | 8.8997e-06 | 2.689e-05 | 2.188 | 1000000000 | 70.0335 | 0.010012 | 2.0794 | 1591.5
B8 | 1591.5 | 4Y | 800879 | 189 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.0046e-08 | 1.1339e-11 | 2.0782 | 1.7899e-08 | 3.0781e-08 | 2.1861 | 2.2362e-05 | 3.8456e-05 | 2.1861 | 1.7622e-12 | 2.1628e-08 | 6.4346 | 1.4232 | B81591.5 | 1000 | 1.7899e-05 | 3.0781e-05 | 2.1861 | 1000000000 | 80.0456 | 0.011339 | 2.0782 | 1591.5
B9 | 1591.5 | 4Y | 900901 | 329 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.0049e-08 | 1.2533e-11 | 2.0757 | 3.1899e-08 | 3.4722e-08 | 2.1826 | 3.5425e-05 | 3.8561e-05 | 2.1826 | 1.9823e-12 | 2.4329e-08 | 6.3226 | 1.4272 | B91591.5 | 1000 | 3.1899e-05 | 3.4722e-05 | 2.1826 | 1000000000 | 90.0487 | 0.012533 | 2.0757 | 1591.5
B10 | 1591.5 | 4Y | 1001011 | 514 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0006e-07 | 1.4046e-11 | 2.077 | 5.0398e-08 | 3.8726e-08 | 2.178 | 5.0369e-05 | 3.8703e-05 | 2.178 | 2.2026e-12 | 2.7032e-08 | 6.3772 | 1.4326 | B101591.5 | 1000 | 5.0398e-05 | 3.8726e-05 | 2.178 | 1000000000 | 100.0607 | 0.014046 | 2.077 | 1591.5
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 3.2278e-14 | 2.0895 | 0 | 3.2274e-10 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-07 | 2.0895 | 1000000000 | 0 | 3.2278e-05 | 2.0895 | 1591.5
box_zero | 1591.5 | 5Y | 5001 | 123 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 0 | 0 | 0 | 3.0603e-15 | 2.4469 | 0 | 0 | 0 | 1.226e-15 | 1.6203e-11 | 0 | 0.00018886 | box_zero1591.5 | 1000 | 0 | 3.0603e-12 | 2.4469 | 1000000000 | 0 | 0 | 0 | 1591.5
C1 | 1591.5 | 5Y | 104502 | 163 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.9511e-10 | 1.4883e-13 | 2.0582 | 3.9998e-10 | 5.5656e-10 | 2.0185 | 4.0196e-05 | 5.5932e-05 | 2.0185 | 2.3118e-14 | 2.8487e-10 | 6.4379 | 1.9537 | C11591.5 | 1000 | 3.9998e-07 | 5.5656e-07 | 2.0185 | 1000000000 | 0.99511 | 0.00014883 | 2.0582 | 1591.5
C2 | 1591.5 | 5Y | 204731 | 157 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.9975e-09 | 2.8792e-13 | 2.073 | 3.3999e-10 | 8.4116e-10 | 2.0912 | 1.7021e-05 | 4.2112e-05 | 2.0912 | 4.5171e-14 | 5.5551e-10 | 6.3741 | 1.5142 | C21591.5 | 1000 | 3.3999e-07 | 8.4116e-07 | 2.0912 | 1000000000 | 1.9975 | 0.00028792 | 2.073 | 1591.5
C3 | 1591.5 | 5Y | 304202 | 160 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 2.9923e-09 | 4.2654e-13 | 2.0757 | 3.6999e-10 | 1.1824e-09 | 2.1455 | 1.2365e-05 | 3.9517e-05 | 2.1455 | 6.7057e-14 | 8.241e-10 | 6.3609 | 1.4348 | C31591.5 | 1000 | 3.6999e-07 | 1.1824e-06 | 2.1455 | 1000000000 | 2.9923 | 0.00042654 | 2.0757 | 1591.5
C4 | 1591.5 | 5Y | 404451 | 146 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 3.9949e-09 | 5.6671e-13 | 2.0766 | 2.2999e-10 | 1.5445e-09 | 2.1754 | 5.7573e-06 | 3.8664e-05 | 2.1754 | 8.9113e-14 | 1.0948e-09 | 6.3594 | 1.4108 | C41591.5 | 1000 | 2.2999e-07 | 1.5445e-06 | 2.1754 | 1000000000 | 3.9949 | 0.00056671 | 2.0766 | 1591.5
C5 | 1591.5 | 5Y | 504786 | 116 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.9983e-09 | 7.0718e-13 | 2.077 | -6.9997e-11 | 1.9139e-09 | 2.1922 | -1.4005e-06 | 3.8292e-05 | 2.1922 | 1.1119e-13 | 1.3657e-09 | 6.3601 | 1.4014 | C51591.5 | 1000 | -6.9997e-08 | 1.9139e-06 | 2.1922 | 1000000000 | 4.9983 | 0.00070718 | 2.077 | 1591.5
C6 | 1591.5 | 5Y | 604205 | 67 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 5.9926e-09 | 8.4646e-13 | 2.0772 | -5.5998e-10 | 2.2831e-09 | 2.2022 | -9.3448e-06 | 3.8099e-05 | 2.2022 | 1.3306e-13 | 1.6342e-09 | 6.3613 | 1.3971 | C61591.5 | 1000 | -5.5998e-07 | 2.2831e-06 | 2.2022 | 1000000000 | 5.9926 | 0.00084646 | 2.0772 | 1591.5
C7 | 1591.5 | 5Y | 704439 | 137 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 6.995e-09 | 9.9568e-13 | 2.0786 | 1.3999e-10 | 2.657e-09 | 2.2085 | 2.0014e-06 | 3.7985e-05 | 2.2085 | 1.5512e-13 | 1.9048e-09 | 6.4189 | 1.3949 | C71591.5 | 1000 | 1.3999e-07 | 2.657e-06 | 2.2085 | 1000000000 | 6.995 | 0.00099568 | 2.0786 | 1591.5
C8 | 1591.5 | 5Y | 803838 | 238 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 7.9891e-09 | 1.1464e-12 | 2.08 | 1.15e-09 | 3.0289e-09 | 2.2128 | 1.4394e-05 | 3.7914e-05 | 2.2128 | 1.7699e-13 | 2.1732e-09 | 6.4774 | 1.3937 | C81591.5 | 1000 | 1.15e-06 | 3.0289e-06 | 2.2128 | 1000000000 | 7.9891 | 0.0011464 | 2.08 | 1591.5
C9 | 1591.5 | 5Y | 904057 | 329 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.9914e-09 | 1.2668e-12 | 2.0774 | 2.0599e-09 | 3.4045e-09 | 2.2157 | 2.2911e-05 | 3.7865e-05 | 2.2157 | 1.9904e-13 | 2.4438e-09 | 6.3646 | 1.3931 | C91591.5 | 1000 | 2.0599e-06 | 3.4045e-06 | 2.2157 | 1000000000 | 8.9914 | 0.0012668 | 2.0774 | 1591.5
C10 | 1591.5 | 5Y | 1003463 | 493 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.9856e-09 | 1.4186e-12 | 2.0787 | 3.6999e-09 | 3.7776e-09 | 2.2178 | 3.7053e-05 | 3.7832e-05 | 2.2178 | 2.2091e-13 | 2.7122e-09 | 6.4218 | 1.3928 | C101591.5 | 1000 | 3.6999e-06 | 3.7776e-06 | 2.2178 | 1000000000 | 9.9856 | 0.0014186 | 2.0787 | 1591.5
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 3.2278e-14 | 2.0895 | 0 | 3.2274e-10 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-07 | 2.0895 | 1000000000000 | 0 | 0.032278 | 2.0895 | 1591.5
main_zero | 1591.5 | 6Y | 50024 | 1095 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.9144e-11 | 1.5423e-12 | 2.5705 | 1.08e-09 | 4.7013e-11 | 2.0115 | 0.0021978 | 0.00011021 | 2.0182 | 1.2264e-15 | 1.6208e-11 | 1257.6479 | 2.9006 | main_zero1591.5 | 1000 | 1.08e-06 | 4.7013e-08 | 2.0115 | 1000000000000 | 49.1438 | 1.5423 | 2.5705 | 1591.5
D1 | 1591.5 | 6Y | 146641 | 1150 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.6632e-11 | 2.0514e-14 | 2.0807 | 5.5001e-11 | 5.5027e-11 | 2.0171 | 5.692e-05 | 5.6946e-05 | 2.0171 | 3.3523e-15 | 4.2298e-11 | 6.1195 | 1.3009 | D11591.5 | 1000 | 5.5001e-08 | 5.5027e-08 | 2.0171 | 1000000000000 | 96.6323 | 0.020514 | 2.0807 | 1591.5
D2 | 1591.5 | 6Y | 245729 | 1138 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.9574e-10 | 3.3875e-14 | 2.0869 | 4.3001e-11 | 8.2816e-11 | 2.0883 | 2.197e-05 | 4.2311e-05 | 2.0883 | 5.5326e-15 | 6.9055e-11 | 6.1228 | 1.1993 | D21591.5 | 1000 | 4.3001e-08 | 8.2816e-08 | 2.0883 | 1000000000000 | 195.736 | 0.033875 | 2.0869 | 1591.5
D3 | 1591.5 | 6Y | 342214 | 1170 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 2.9224e-10 | 4.7872e-14 | 2.0888 | 7.5002e-11 | 1.1576e-10 | 2.1426 | 2.5666e-05 | 3.9614e-05 | 2.1426 | 7.6556e-15 | 9.5109e-11 | 6.2533 | 1.2171 | D31591.5 | 1000 | 7.5002e-08 | 1.1576e-07 | 2.1426 | 1000000000000 | 292.2363 | 0.047872 | 2.0888 | 1591.5
D4 | 1591.5 | 6Y | 442013 | 1123 | 1591.5 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 3.9205e-10 | 6.0946e-14 | 2.085 | 2.8001e-11 | 1.5173e-10 | 2.1738 | 7.1423e-06 | 3.8703e-05 | 2.1738 | 9.8515e-15 | 1.2206e-10 | 6.1865 | 1.2431 | D41591.5 | 1000 | 2.8001e-08 | 1.5173e-07 | 2.1738 | 1000000000000 | 392.0512 | 0.060946 | 2.085 | 1591.5
D5 | 1591.5 | 6Y | 538500 | 1140 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.8855e-10 | 7.5221e-14 | 2.0858 | 4.5001e-11 | 1.8721e-10 | 2.1909 | 9.2114e-06 | 3.832e-05 | 2.1909 | 1.1975e-14 | 1.4811e-10 | 6.2817 | 1.264 | D51591.5 | 1000 | 4.5001e-08 | 1.8721e-07 | 2.1909 | 1000000000000 | 488.5535 | 0.075221 | 2.0858 | 1591.5
D6 | 1591.5 | 6Y | 641458 | 1058 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 5.9153e-10 | 8.9862e-14 | 2.0851 | -3.7001e-11 | 2.2541e-10 | 2.2017 | -6.2553e-06 | 3.8108e-05 | 2.2017 | 1.424e-14 | 1.7592e-10 | 6.3105 | 1.2814 | D61591.5 | 1000 | -3.7001e-08 | 2.2541e-07 | 2.2017 | 1000000000000 | 591.5278 | 0.089862 | 2.0851 | 1591.5
D7 | 1591.5 | 6Y | 740537 | 1164 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 6.9062e-10 | 1.0397e-13 | 2.0846 | 6.9002e-11 | 2.6236e-10 | 2.2083 | 9.9915e-06 | 3.799e-05 | 2.2083 | 1.642e-14 | 2.0267e-10 | 6.3317 | 1.2945 | D71591.5 | 1000 | 6.9002e-08 | 2.6236e-07 | 2.2083 | 1000000000000 | 690.6225 | 0.10397 | 2.0846 | 1591.5
D8 | 1591.5 | 6Y | 837056 | 1308 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 7.8716e-10 | 1.1558e-13 | 2.0813 | 2.1301e-10 | 2.9845e-10 | 2.2126 | 2.7061e-05 | 3.7916e-05 | 2.2126 | 1.8544e-14 | 2.2873e-10 | 6.2326 | 1.3048 | D81591.5 | 1000 | 2.1301e-07 | 2.9845e-07 | 2.2126 | 1000000000000 | 787.1568 | 0.11558 | 2.0813 | 1591.5
D9 | 1591.5 | 6Y | 936847 | 1426 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.8696e-10 | 1.2955e-13 | 2.0809 | 3.3101e-10 | 3.3583e-10 | 2.2158 | 3.732e-05 | 3.7864e-05 | 2.2158 | 2.074e-14 | 2.5568e-10 | 6.2465 | 1.3135 | D91591.5 | 1000 | 3.3101e-07 | 3.3583e-07 | 2.2158 | 1000000000000 | 886.9637 | 0.12955 | 2.0809 | 1591.5
D10 | 1591.5 | 6Y | 1033147 | 1534 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.8328e-10 | 1.4304e-13 | 2.0806 | 4.3901e-10 | 3.7195e-10 | 2.218 | 4.4649e-05 | 3.7829e-05 | 2.218 | 2.2859e-14 | 2.8169e-10 | 6.2576 | 1.3204 | D101591.5 | 1000 | 4.3901e-07 | 3.7195e-07 | 2.218 | 1000000000000 | 983.2789 | 0.14304 | 2.0806 | 1591.5
coax zero | 1591.5 | 6Y | 0 | 17 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 0 | 3.2281e-15 | 2.0895 | 1.7e-11 | 3.2277e-11 | 2.0895 | 1.2568e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 1.7e-08 | 3.2277e-08 | 2.0895 | 1000000000000 | 0 | 0.0032281 | 2.0895 | 1591.5
GR1000A | 1591.5 | 6Y | 999924 | -113 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0001e-09 | 1.391e-13 | 2.0758 | -1.13e-10 | 3.7747e-10 | 2.2232 | 2.2127e-14 | 2.727e-10 |  |  |  | #DIV/0! | #DIV/0! | GR1000A1591.5 | 1000 | -1.13e-07 | 3.7747e-07 | 2.2232 | 1000000000000 | 1000.0826 | 0.1391 | 2.0758 | 1591.5
GR1000B | 1591.5 | 6Y | 999907 | -83 | 1591.4 | 19.64 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0001e-09 | 1.391e-13 | 2.0758 | -8.3002e-11 | 3.7746e-10 | 2.2232 | 2.2127e-14 | 2.7269e-10 |  |  |  | #DIV/0! | #DIV/0! | GR1000B1591.5 | 1000 | -8.3002e-08 | 3.7746e-07 | 2.2232 | 1000000000000 | 1000.0656 | 0.1391 | 2.0758 | 1591.5
E1 | 1591.5 | 6Y | 60487 | 1341 | 1591.7 | 19.67 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0465e-11 | 1.2853e-12 | 2.5704 | 2.4601e-10 | 4.4313e-11 | 2.0219 | 0.0023509 | 0.00048048 | 2.0217 | 1.4566e-15 | 1.9036e-11 | 882.3969 | 2.3279 | E11591.5 | 1000 | 2.4601e-07 | 4.4313e-08 | 2.0219 | 1000000000000 | 10.4647 | 1.2853 | 2.5704 | 1591.5
E2 | 1591.5 | 6Y | 71300 | 1391 | 1591.7 | 19.65 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 2.1279e-11 | 1.2853e-12 | 2.5704 | 2.9601e-10 | 4.4711e-11 | 2.0195 | 0.0013911 | 0.00021955 | 2.0131 | 1.6945e-15 | 2.1956e-11 | 758.4985 | 2.0364 | E21591.5 | 1000 | 2.9601e-07 | 4.4711e-08 | 2.0195 | 1000000000000 | 21.2794 | 1.2853 | 2.5704 | 1591.5
E3 | 1591.5 | 6Y | 81517 | 1332 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 3.1498e-11 | 1.2853e-12 | 2.5704 | 2.3701e-10 | 4.5341e-11 | 2.0165 | 0.00075247 | 0.00014576 | 2.0137 | 1.9193e-15 | 2.4715e-11 | 669.6534 | 1.8345 | E31591.5 | 1000 | 2.3701e-07 | 4.5341e-08 | 2.0165 | 1000000000000 | 31.498 | 1.2853 | 2.5704 | 1591.5
E4 | 1591.5 | 6Y | 91813 | 1324 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 4.1796e-11 | 1.2853e-12 | 2.5703 | 2.2901e-10 | 4.6228e-11 | 2.0133 | 0.00054793 | 0.00011131 | 2.0119 | 2.1459e-15 | 2.7496e-11 | 598.9535 | 1.6813 | E41591.5 | 1000 | 2.2901e-07 | 4.6228e-08 | 2.0133 | 1000000000000 | 41.7956 | 1.2853 | 2.5703 | 1591.5
E5 | 1591.5 | 6Y | 101710 | 1334 | 1591.7 | 19.63 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 5.1694e-11 | 1.2853e-12 | 2.5703 | 2.3901e-10 | 4.7318e-11 | 2.011 | 0.00046236 | 9.1936e-05 | 2.01 | 2.3636e-15 | 3.0169e-11 | 543.7685 | 1.5685 | E51591.5 | 1000 | 2.3901e-07 | 4.7318e-08 | 2.011 | 1000000000000 | 51.6942 | 1.2853 | 2.5703 | 1591.5
E6 | 1591.5 | 6Y | 111557 | 1338 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 6.1543e-11 | 1.2853e-12 | 2.5702 | 2.4301e-10 | 4.8633e-11 | 2.0097 | 0.00039487 | 7.9262e-05 | 2.0091 | 2.5803e-15 | 3.2828e-11 | 498.1062 | 1.4814 | E61591.5 | 1000 | 2.4301e-07 | 4.8633e-08 | 2.0097 | 1000000000000 | 61.5428 | 1.2853 | 2.5702 | 1591.5
E7 | 1591.5 | 6Y | 120990 | 1319 | 1591.7 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 7.0977e-11 | 1.2853e-12 | 2.5702 | 2.2401e-10 | 5.0099e-11 | 2.0098 | 0.00031561 | 7.0714e-05 | 2.0094 | 2.7879e-15 | 3.5376e-11 | 461.0198 | 1.4162 | E71591.5 | 1000 | 2.2401e-07 | 5.0099e-08 | 2.0098 | 1000000000000 | 70.9773 | 1.2853 | 2.5702 | 1591.5
E8 | 1591.5 | 6Y | 130672 | 1315 | 1591.7 | 19.62 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.0661e-11 | 1.2853e-12 | 2.5701 | 2.2001e-10 | 5.1806e-11 | 2.0114 | 0.00027276 | 6.431e-05 | 2.0111 | 3.0009e-15 | 3.7991e-11 | 428.2898 | 1.3637 | E81591.5 | 1000 | 2.2001e-07 | 5.1806e-08 | 2.0114 | 1000000000000 | 80.6608 | 1.2853 | 2.5701 | 1591.5
E9 | 1591.5 | 6Y | 140694 | 1310 | 1591.7 | 19.6 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 9.0684e-11 | 1.2852e-12 | 2.57 | 2.1501e-10 | 5.3775e-11 | 2.0146 | 0.0002371 | 5.9353e-05 | 2.0144 | 3.2214e-15 | 4.0697e-11 | 398.9693 | 1.3213 | E91591.5 | 1000 | 2.1501e-07 | 5.3775e-08 | 2.0146 | 1000000000000 | 90.6844 | 1.2852 | 2.57 | 1591.5
E10 | 1591.5 | 6Y | 150067 | 1304 | 1591.7 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 1.0006e-10 | 1.2852e-12 | 2.57 | 2.0901e-10 | 5.5788e-11 | 2.0188 | 0.00020889 | 5.5792e-05 | 2.0186 | 3.4276e-15 | 4.3229e-11 | 374.9622 | 1.2905 | E101591.5 | 1000 | 2.0901e-07 | 5.5788e-08 | 2.0188 | 1000000000000 | 100.0589 | 1.2852 | 2.57 | 1591.5
CuZero | 1591.5 | 6Y | 888 | 15 | 1591.6 | 19.61 | 0.2 | 10 | 0 | 0.0001 | 10 | 0 | 7e-05 | 10 | 8.8814e-13 | 3.2298e-15 | 2.089 | 1.5e-11 | 3.2276e-11 | 2.0894 | 1.452e-16 | 2.9399e-12 |  |  |  | #DIV/0! | #DIV/0! | CuZero1591.5 | 1000 | 1.5e-08 | 3.2276e-08 | 2.0894 | 1000000000000 | 0.88814 | 0.0032298 | 2.089 | 1591.5

## AH2700_data
Capacitor description: | AH2700A Zero with AH2 Leads
2020-10-12 | 15:58
-2.1e-07 | 3.05e-05 | 1591.5 | 15
-2e-07 | 2.9e-05 | 1591.5 | 15
-2e-07 | 2.45e-05 | 1591.5 | 15
-2.1e-07 | 2.75e-05 | 1591.5 | 15
-2e-07 | 2.6e-05 | 1591.5 | 15
1.1e-07 | 2.9e-05 | 1591.5 | 15
-1e-07 | 2.9e-05 | 1591.5 | 15
1e-07 | 2.75e-05 | 1591.5 | 15
-2e-07 | 2.9e-05 | 1591.5 | 15
1e-07 | 2.75e-05 | 1591.5 | 15
-3.5e-07 | 2.3e-05 | 1591.5 | 15
-3.5e-07 | 2.75e-05 | 1591.5 | 15
-1e-07 | 2.9e-05 | 1591.5 | 15
-1e-07 | 2.75e-05 | 1591.5 | 15
-1e-07 | 2.45e-05 | 1591.5 | 15
1e-07 | 2.6e-05 | 1591.5 | 15
-2.1e-07 | 2.75e-05 | 1591.5 | 15
-2.1e-07 | 2.9e-05 | 1591.5 | 15
-2.1e-07 | 2.6e-05 | 1591.5 | 15
-1e-07 | 2.45e-05 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
-1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06
end | of | the | test
Capacitor description: | 1AH11A measure
2020-10-12 | 16:07
10 | -0.0001232 | 1591.5 | 15
10 | -0.000119 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001262 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001202 | 1591.5 | 15
10 | -0.0001202 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001262 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001187 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10 | 1.8e-07 | -0.00012269 | 2.05e-06
end | of | the | test
Capacitor description: | 1AH11B measure
2020-10-12 | 16:10
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -5.86e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -7.06e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10 | 2.65e-07 | -6.61e-05 | 2.22e-06
end | of | the | test
Capacitor description: | 1AH11C measure
2020-10-12 | 16:13
Capacitor description: | 1AH11C measure
2020-10-12 | 16:14
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001634 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06
end | of | the | test
Capacitor description: | 1AH11D measure
2020-10-12 | 16:17
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002085 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002077 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002085 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06
end | of | the | test
Capacitor description: | GR1000A measure
2020-10-12 | 16:22
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17947 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17939 | 1591.5 | 1.5
1000.0317 | 0.1794 | 1591.5 | 1.5
1000.0317 | 0.17937 | 1591.5 | 1.5
1000.0317 | 0.17925 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17935 | 1591.5 | 1.5
1000.0317 | 0.17938 | 1591.5 | 1.5
1000.0317 | 0.1794 | 1591.5 | 1.5
1000.0317 | 0.17949 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17937 | 1591.5 | 1.5
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17935 | 1591.5 | 1.5
1000.0317 | 0.17933 | 1591.5 | 1.5
1000.0317 | 0.17932 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05
end | of | the | test
Capacitor description: | GR1000B measure
2020-10-12 | 16:28
1000.0164 | 0.18445 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.18445 | 1591.5 | 1.5
1000.0164 | 0.18455 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.18443 | 1591.5 | 1.5
1000.0164 | 0.18449 | 1591.5 | 1.5
1000.0164 | 0.18443 | 1591.5 | 1.5
1000.0164 | 0.1844 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.1844 | 1591.5 | 1.5
1000.0164 | 0.18451 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.18454 | 1591.5 | 1.5
1000.0164 | 0.18435 | 1591.5 | 1.5
1000.0164 | 0.18448 | 1591.5 | 1.5
1000.0164 | 0.18435 | 1591.5 | 1.5
1000.0164 | 0.18438 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05
end | of | the | test
Capacitor description: | UB1000 measure 20.07deg, 1019.93 mbar, 57.4%RH
2020-10-12 | 16:33
1000.1471 | 0.14212 | 1591.5 | 1.5
1000.1471 | 0.14209 | 1591.5 | 1.5
1000.1471 | 0.14217 | 1591.5 | 1.5
1000.1471 | 0.14217 | 1591.5 | 1.5
1000.1471 | 0.14208 | 1591.5 | 1.5
1000.1471 | 0.14199 | 1591.5 | 1.5
1000.1471 | 0.14211 | 1591.5 | 1.5
1000.1471 | 0.14202 | 1591.5 | 1.5
1000.1471 | 0.14205 | 1591.5 | 1.5
1000.1471 | 0.14189 | 1591.5 | 1.5
1000.1471 | 0.14199 | 1591.5 | 1.5
1000.1471 | 0.14195 | 1591.5 | 1.5
1000.1471 | 0.142 | 1591.5 | 1.5
1000.1471 | 0.14193 | 1591.5 | 1.5
1000.1471 | 0.14197 | 1591.5 | 1.5
1000.1471 | 0.14194 | 1591.5 | 1.5
1000.1471 | 0.14185 | 1591.5 | 1.5
1000.1471 | 0.14192 | 1591.5 | 1.5
1000.1471 | 0.1419 | 1591.5 | 1.5
1000.1471 | 0.14184 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1471 | 5.89e-06 | 0.142 | 9.76e-05
end | of | the | test
Capacitor description: | Zero point of GR Decade Capacitor
2020-10-12 | 16:40
49.221 | 1.3388 | 1591.5 | 15
49.221 | 1.3388 | 1591.5 | 15
49.2213 | 1.3365 | 1591.5 | 15
49.2218 | 1.3297 | 1591.5 | 15
49.221 | 1.3394 | 1591.5 | 15
49.222 | 1.3274 | 1591.5 | 15
49.222 | 1.3271 | 1591.5 | 15
49.222 | 1.3268 | 1591.5 | 15
49.222 | 1.3261 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.2219 | 1.3288 | 1591.5 | 15
49.2219 | 1.3291 | 1591.5 | 15
49.2219 | 1.3293 | 1591.5 | 15
49.222 | 1.3281 | 1591.5 | 15
49.2215 | 1.3358 | 1591.5 | 15
49.2207 | 1.3411 | 1591.5 | 15
49.2212 | 1.3381 | 1591.5 | 15
49.2214 | 1.3351 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
49.2216 | 0.00043741 | 1.3317 | 0.0053142
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 10 measure
2020-10-12 | 16:44
59.6543 | 1.3374 | 1591.5 | 15
59.6543 | 1.3374 | 1591.5 | 15
59.6544 | 1.3364 | 1591.5 | 15
59.6544 | 1.3364 | 1591.5 | 15
59.6542 | 1.3395 | 1591.5 | 15
59.6542 | 1.3389 | 1591.5 | 15
59.6544 | 1.337 | 1591.5 | 15
59.6542 | 1.3393 | 1591.5 | 15
59.6541 | 1.3409 | 1591.5 | 15
59.6541 | 1.3412 | 1591.5 | 15
59.6543 | 1.3383 | 1591.5 | 15
59.6543 | 1.3382 | 1591.5 | 15
59.6544 | 1.3357 | 1591.5 | 15
59.6544 | 1.3357 | 1591.5 | 15
59.6545 | 1.3349 | 1591.5 | 15
59.6545 | 1.3353 | 1591.5 | 15
59.6545 | 1.3345 | 1591.5 | 15
59.6544 | 1.3371 | 1591.5 | 15
59.6544 | 1.3365 | 1591.5 | 15
59.6539 | 1.3432 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
59.6543 | 0.00014858 | 1.3377 | 0.0022088
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 20 measure
2020-10-12 | 16:55
70.2236 | 1.3732 | 1591.5 | 15
70.2235 | 1.3685 | 1591.5 | 15
70.2235 | 1.3704 | 1591.5 | 15
70.2233 | 1.3735 | 1591.5 | 15
70.2235 | 1.3686 | 1591.5 | 15
70.2234 | 1.3679 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2234 | 1.3677 | 1591.5 | 15
70.2234 | 1.3662 | 1591.5 | 15
70.2234 | 1.3667 | 1591.5 | 15
70.2234 | 1.3664 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2234 | 1.368 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2233 | 1.3679 | 1591.5 | 15
70.2234 | 1.3669 | 1591.5 | 15
70.2234 | 1.3713 | 1591.5 | 15
70.2234 | 1.3696 | 1591.5 | 15
70.2235 | 1.3698 | 1591.5 | 15
70.2234 | 1.369 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
70.2234 | 6.02e-05 | 1.3687 | 0.0020591
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 30 measure
2020-10-12 | 16:58
80.8297 | 1.3379 | 1591.5 | 15
80.8297 | 1.338 | 1591.5 | 15
80.8297 | 1.3374 | 1591.5 | 15
80.8298 | 1.3374 | 1591.5 | 15
80.8284 | 1.351 | 1591.5 | 15
80.8284 | 1.3512 | 1591.5 | 15
80.8284 | 1.3513 | 1591.5 | 15
80.8284 | 1.3512 | 1591.5 | 15
80.8284 | 1.351 | 1591.5 | 15
80.8285 | 1.351 | 1591.5 | 15
80.8285 | 1.3509 | 1591.5 | 15
80.8285 | 1.3511 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3502 | 1591.5 | 15
80.8287 | 1.3502 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3505 | 1591.5 | 15
80.8287 | 1.3505 | 1591.5 | 15
80.8288 | 1.3499 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
80.8288 | 0.00048247 | 1.3481 | 0.0052107
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 40 measure
2020-10-12 | 17:01
90.9233 | 1.3518 | 1591.5 | 15
90.9233 | 1.3519 | 1591.5 | 15
90.9233 | 1.3519 | 1591.5 | 15
90.9237 | 1.3498 | 1591.5 | 15
90.9237 | 1.3495 | 1591.5 | 15
90.9237 | 1.3498 | 1591.5 | 15
90.9238 | 1.3494 | 1591.5 | 15
90.9238 | 1.3495 | 1591.5 | 15
90.9238 | 1.3499 | 1591.5 | 15
90.9238 | 1.3495 | 1591.5 | 15
90.9239 | 1.3492 | 1591.5 | 15
90.9239 | 1.3495 | 1591.5 | 15
90.9239 | 1.3497 | 1591.5 | 15
90.9239 | 1.3498 | 1591.5 | 15
90.9239 | 1.3497 | 1591.5 | 15
90.9239 | 1.3498 | 1591.5 | 15
90.9239 | 1.3502 | 1591.5 | 15
90.9239 | 1.3501 | 1591.5 | 15
90.924 | 1.3498 | 1591.5 | 15
90.924 | 1.3497 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
90.9238 | 0.00021909 | 1.35 | 0.00079946
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 50 measure
2020-10-12 | 17:04
100.4765 | 1.3524 | 1591.5 | 15
100.4766 | 1.3524 | 1591.5 | 15
100.4766 | 1.3529 | 1591.5 | 15
100.4773 | 1.3434 | 1591.5 | 15
100.4774 | 1.343 | 1591.5 | 15
100.4774 | 1.3424 | 1591.5 | 15
100.4774 | 1.3426 | 1591.5 | 15
100.4775 | 1.3425 | 1591.5 | 15
100.4775 | 1.343 | 1591.5 | 15
100.4776 | 1.3409 | 1591.5 | 15
100.4776 | 1.3407 | 1591.5 | 15
100.4775 | 1.3439 | 1591.5 | 15
100.4773 | 1.3471 | 1591.5 | 15
100.4775 | 1.3443 | 1591.5 | 15
100.4776 | 1.3431 | 1591.5 | 15
100.4776 | 1.3424 | 1591.5 | 15
100.4777 | 1.3413 | 1591.5 | 15
100.4777 | 1.3413 | 1591.5 | 15
100.4777 | 1.341 | 1591.5 | 15
100.4777 | 1.3409 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100.4774 | 0.00036577 | 1.3441 | 0.0038393
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 60 measure
2020-10-12 | 17:06
110.69 | 1.346 | 1591.5 | 15
110.6902 | 1.3465 | 1591.5 | 15
110.6904 | 1.3467 | 1591.5 | 15
110.6906 | 1.3468 | 1591.5 | 15
110.6907 | 1.3476 | 1591.5 | 15
110.6909 | 1.3457 | 1591.5 | 15
110.691 | 1.345 | 1591.5 | 15
110.6911 | 1.3443 | 1591.5 | 15
110.6912 | 1.345 | 1591.5 | 15
110.6913 | 1.345 | 1591.5 | 15
110.6913 | 1.3452 | 1591.5 | 15
110.6914 | 1.3452 | 1591.5 | 15
110.6915 | 1.3451 | 1591.5 | 15
110.6915 | 1.3453 | 1591.5 | 15
110.6915 | 1.3454 | 1591.5 | 15
110.6916 | 1.3452 | 1591.5 | 15
110.6916 | 1.345 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
110.6912 | 0.00050256 | 1.3455 | 0.00079577
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 70 measure
2020-10-12 | 17:11
120.5074 | 1.3464 | 1591.5 | 15
120.5075 | 1.3463 | 1591.5 | 15
120.5076 | 1.346 | 1591.5 | 15
120.5077 | 1.346 | 1591.5 | 15
120.5078 | 1.3461 | 1591.5 | 15
120.5079 | 1.3461 | 1591.5 | 15
120.5079 | 1.3462 | 1591.5 | 15
120.508 | 1.3462 | 1591.5 | 15
120.5081 | 1.3462 | 1591.5 | 15
120.5081 | 1.3461 | 1591.5 | 15
120.5082 | 1.3461 | 1591.5 | 15
120.5082 | 1.3462 | 1591.5 | 15
120.5082 | 1.3464 | 1591.5 | 15
120.5083 | 1.3464 | 1591.5 | 15
120.5083 | 1.3464 | 1591.5 | 15
120.5083 | 1.3467 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3464 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
120.5081 | 0.00029987 | 1.3463 | 0.00018653
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 80 measure
2020-10-12 | 17:23
129.9739 | 1.3534 | 1591.5 | 15
129.9739 | 1.3539 | 1591.5 | 15
129.9738 | 1.3542 | 1591.5 | 15
129.9739 | 1.3535 | 1591.5 | 15
129.974 | 1.3516 | 1591.5 | 15
129.974 | 1.3516 | 1591.5 | 15
129.974 | 1.3522 | 1591.5 | 15
129.974 | 1.3527 | 1591.5 | 15
129.9739 | 1.3529 | 1591.5 | 15
129.9739 | 1.3529 | 1591.5 | 15
129.9739 | 1.3532 | 1591.5 | 15
129.974 | 1.3517 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.3524 | 1591.5 | 15
129.974 | 1.3513 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.3523 | 1591.5 | 15
129.9741 | 1.3512 | 1591.5 | 15
129.974 | 1.3523 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
129.974 | 5.59e-05 | 1.3525 | 0.00082738
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 90 measure
2020-10-12 | 17:25
139.7372 | 1.353 | 1591.5 | 7.5
139.7376 | 1.3527 | 1591.5 | 7.5
139.7378 | 1.3535 | 1591.5 | 7.5
139.7381 | 1.3514 | 1591.5 | 7.5
139.7382 | 1.3511 | 1591.5 | 7.5
139.7384 | 1.3512 | 1591.5 | 7.5
139.7384 | 1.3522 | 1591.5 | 7.5
139.7386 | 1.3513 | 1591.5 | 7.5
139.7387 | 1.3515 | 1591.5 | 7.5
139.7387 | 1.3521 | 1591.5 | 7.5
139.7387 | 1.3529 | 1591.5 | 7.5
139.7388 | 1.3529 | 1591.5 | 7.5
139.7389 | 1.3528 | 1591.5 | 7.5
139.739 | 1.3515 | 1591.5 | 7.5
139.739 | 1.3524 | 1591.5 | 7.5
139.739 | 1.3521 | 1591.5 | 7.5
139.7391 | 1.3515 | 1591.5 | 7.5
139.7391 | 1.3523 | 1591.5 | 7.5
139.7392 | 1.3522 | 1591.5 | 7.5
139.7392 | 1.3516 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
139.7386 | 0.0005509 | 1.3521 | 0.0006845
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail '100' measure
2020-10-12 | 17:28
149.5761 | 1.356 | 1591.5 | 7.5
149.5763 | 1.3542 | 1591.5 | 7.5
149.5764 | 1.3538 | 1591.5 | 7.5
149.5764 | 1.3558 | 1591.5 | 7.5
149.5765 | 1.3557 | 1591.5 | 7.5
149.5766 | 1.3552 | 1591.5 | 7.5
149.5765 | 1.3565 | 1591.5 | 7.5
149.5766 | 1.3566 | 1591.5 | 7.5
149.5767 | 1.3559 | 1591.5 | 7.5
149.5768 | 1.3545 | 1591.5 | 7.5
149.5768 | 1.355 | 1591.5 | 7.5
149.5769 | 1.3549 | 1591.5 | 7.5
149.5769 | 1.355 | 1591.5 | 7.5
149.5768 | 1.3557 | 1591.5 | 7.5
149.5769 | 1.3561 | 1591.5 | 7.5
149.5769 | 1.3565 | 1591.5 | 7.5
149.5769 | 1.3562 | 1591.5 | 7.5
149.577 | 1.3555 | 1591.5 | 7.5
149.577 | 1.3556 | 1591.5 | 7.5
149.577 | 1.3556 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
149.5767 | 0.00025146 | 1.3555 | 0.00075452
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 1
2020-10-12 | 17:32
145.8881 | 1.6341 | 1591.5 | 7.5
145.8884 | 1.6282 | 1591.5 | 7.5
145.8882 | 1.6363 | 1591.5 | 7.5
145.8884 | 1.629 | 1591.5 | 7.5
145.8881 | 1.6392 | 1591.5 | 7.5
145.8886 | 1.623 | 1591.5 | 7.5
145.8883 | 1.6321 | 1591.5 | 7.5
145.8884 | 1.6312 | 1591.5 | 7.5
145.888 | 1.6409 | 1591.5 | 7.5
145.8885 | 1.6257 | 1591.5 | 7.5
145.8891 | 1.6075 | 1591.5 | 7.5
145.889 | 1.6085 | 1591.5 | 7.5
145.889 | 1.6094 | 1591.5 | 7.5
145.889 | 1.6081 | 1591.5 | 7.5
145.8893 | 1.5986 | 1591.5 | 7.5
145.8894 | 1.5967 | 1591.5 | 7.5
145.8892 | 1.6014 | 1591.5 | 7.5
145.8891 | 1.606 | 1591.5 | 7.5
145.8893 | 1.5992 | 1591.5 | 7.5
145.8894 | 1.5961 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
145.8887 | 0.00046896 | 1.6176 | 0.015335
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 2
2020-10-12 | 17:34
244.935 | 1.4324 | 1591.5 | 7.5
244.935 | 1.4322 | 1591.5 | 7.5
244.935 | 1.432 | 1591.5 | 7.5
244.9351 | 1.4313 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4328 | 1591.5 | 7.5
244.9348 | 1.4348 | 1591.5 | 7.5
244.9348 | 1.4349 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4347 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4355 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4352 | 1591.5 | 7.5
244.9348 | 1.435 | 1591.5 | 7.5
244.9348 | 1.4356 | 1591.5 | 7.5
244.9348 | 1.435 | 1591.5 | 7.5
244.9348 | 1.4356 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
244.9349 | 0.00010611 | 1.4339 | 0.0016017
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 3
2020-10-12 | 17:37
341.4798 | 1.517 | 1591.5 | 3
341.4799 | 1.516 | 1591.5 | 3
341.4801 | 1.5131 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.4798 | 1.5161 | 1591.5 | 3
341.48 | 1.5135 | 1591.5 | 3
341.48 | 1.5134 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.4798 | 1.5159 | 1591.5 | 3
341.48 | 1.5132 | 1591.5 | 3
341.48 | 1.5134 | 1591.5 | 3
341.48 | 1.5131 | 1591.5 | 3
341.48 | 1.5138 | 1591.5 | 3
341.4799 | 1.515 | 1591.5 | 3
341.4789 | 1.523 | 1591.5 | 3
341.479 | 1.5229 | 1591.5 | 3
341.4796 | 1.5184 | 1591.5 | 3
341.4797 | 1.5176 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
341.4798 | 0.00032198 | 1.5154 | 0.0029831
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 4
2020-10-12 | 17:39
441.2625 | 1.5404 | 1591.5 | 3
441.2625 | 1.5404 | 1591.5 | 3
441.2627 | 1.5386 | 1591.5 | 3
441.2626 | 1.5398 | 1591.5 | 3
441.2626 | 1.5397 | 1591.5 | 3
441.2626 | 1.5395 | 1591.5 | 3
441.2626 | 1.5394 | 1591.5 | 3
441.2627 | 1.5391 | 1591.5 | 3
441.2627 | 1.5389 | 1591.5 | 3
441.2626 | 1.5395 | 1591.5 | 3
441.2626 | 1.5397 | 1591.5 | 3
441.2623 | 1.5422 | 1591.5 | 3
441.2625 | 1.541 | 1591.5 | 3
441.2623 | 1.5421 | 1591.5 | 3
441.2626 | 1.5399 | 1591.5 | 3
441.2622 | 1.5429 | 1591.5 | 3
441.2623 | 1.5423 | 1591.5 | 3
441.2623 | 1.5425 | 1591.5 | 3
441.2624 | 1.5412 | 1591.5 | 3
441.2624 | 1.5413 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
441.2625 | 0.0001467 | 1.5405 | 0.0012842
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 5
2020-10-12 | 17:41
537.7662 | 1.6209 | 1591.5 | 3
537.766 | 1.622 | 1591.5 | 3
537.7661 | 1.6214 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.621 | 1591.5 | 3
537.7662 | 1.6204 | 1591.5 | 3
537.7662 | 1.6202 | 1591.5 | 3
537.7662 | 1.6199 | 1591.5 | 3
537.7661 | 1.6211 | 1591.5 | 3
537.7661 | 1.6209 | 1591.5 | 3
537.7661 | 1.6211 | 1591.5 | 3
537.7661 | 1.6214 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.766 | 1.6217 | 1591.5 | 3
537.7661 | 1.6215 | 1591.5 | 3
537.7663 | 1.6191 | 1591.5 | 3
537.7662 | 1.6202 | 1591.5 | 3
537.7662 | 1.6203 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
537.7661 | 6.53e-05 | 1.6209 | 0.0006773
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 6
2020-10-12 | 17:43
640.7043 | 1.6659 | 1591.5 | 3
640.7043 | 1.6662 | 1591.5 | 3
640.7042 | 1.6663 | 1591.5 | 3
640.7041 | 1.6672 | 1591.5 | 3
640.7042 | 1.6661 | 1591.5 | 3
640.7042 | 1.6664 | 1591.5 | 3
640.7042 | 1.6659 | 1591.5 | 3
640.7042 | 1.6662 | 1591.5 | 3
640.7042 | 1.666 | 1591.5 | 3
640.7042 | 1.6662 | 1591.5 | 3
640.7041 | 1.6665 | 1591.5 | 3
640.7041 | 1.6666 | 1591.5 | 3
640.7041 | 1.6667 | 1591.5 | 3
640.7041 | 1.6666 | 1591.5 | 3
640.7041 | 1.6665 | 1591.5 | 3
640.704 | 1.6665 | 1591.5 | 3
640.704 | 1.6666 | 1591.5 | 3
640.704 | 1.6668 | 1591.5 | 3
640.7039 | 1.6674 | 1591.5 | 3
640.7039 | 1.6672 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
640.7041 | 0.00010485 | 1.6665 | 0.00040726
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 7
2020-10-12 | 17:45
739.8079 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6858 | 1591.5 | 1.5
739.808 | 1.686 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6862 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.8081 | 1.6851 | 1591.5 | 1.5
739.8081 | 1.6852 | 1591.5 | 1.5
739.8081 | 1.6849 | 1591.5 | 1.5
739.8081 | 1.6853 | 1591.5 | 1.5
739.8081 | 1.6852 | 1591.5 | 1.5
739.8074 | 1.6895 | 1591.5 | 1.5
739.8074 | 1.6893 | 1591.5 | 1.5
739.8075 | 1.6893 | 1591.5 | 1.5
739.8074 | 1.6894 | 1591.5 | 1.5
739.8076 | 1.6888 | 1591.5 | 1.5
739.8077 | 1.688 | 1591.5 | 1.5
739.8077 | 1.6881 | 1591.5 | 1.5
739.8076 | 1.6888 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
739.8078 | 0.00025568 | 1.687 | 0.0016552
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 8
2020-10-12 | 17:47
836.3519 | 1.7676 | 1591.5 | 1.5
836.3519 | 1.7677 | 1591.5 | 1.5
836.352 | 1.7677 | 1591.5 | 1.5
836.352 | 1.7674 | 1591.5 | 1.5
836.3521 | 1.767 | 1591.5 | 1.5
836.3521 | 1.7675 | 1591.5 | 1.5
836.3521 | 1.7673 | 1591.5 | 1.5
836.3521 | 1.7675 | 1591.5 | 1.5
836.352 | 1.768 | 1591.5 | 1.5
836.352 | 1.7682 | 1591.5 | 1.5
836.3521 | 1.7678 | 1591.5 | 1.5
836.352 | 1.7682 | 1591.5 | 1.5
836.352 | 1.7684 | 1591.5 | 1.5
836.3522 | 1.7676 | 1591.5 | 1.5
836.3521 | 1.7677 | 1591.5 | 1.5
836.3521 | 1.7677 | 1591.5 | 1.5
836.3522 | 1.7674 | 1591.5 | 1.5
836.3522 | 1.7674 | 1591.5 | 1.5
836.3523 | 1.7671 | 1591.5 | 1.5
836.3523 | 1.7671 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
836.3521 | 0.00010751 | 1.7676 | 0.00037086
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 9
2020-10-12 | 17:50
936.1407 | 1.7891 | 1591.5 | 1.5
936.1407 | 1.7896 | 1591.5 | 1.5
936.1409 | 1.7891 | 1591.5 | 1.5
936.141 | 1.7888 | 1591.5 | 1.5
936.141 | 1.7888 | 1591.5 | 1.5
936.1411 | 1.7889 | 1591.5 | 1.5
936.1416 | 1.7853 | 1591.5 | 1.5
936.1412 | 1.7889 | 1591.5 | 1.5
936.1412 | 1.7886 | 1591.5 | 1.5
936.1413 | 1.7885 | 1591.5 | 1.5
936.1413 | 1.7885 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7883 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7888 | 1591.5 | 1.5
936.1415 | 1.7886 | 1591.5 | 1.5
936.1414 | 1.789 | 1591.5 | 1.5
936.1415 | 1.7889 | 1591.5 | 1.5
936.1414 | 1.7894 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
936.1413 | 0.00025167 | 1.7886 | 0.00083413
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail X
2020-10-12 | 17:52
1032.4676 | 1.8824 | 1591.5 | 1.5
1032.4676 | 1.8825 | 1591.5 | 1.5
1032.4678 | 1.8811 | 1591.5 | 1.5
1032.4677 | 1.881 | 1591.5 | 1.5
1032.4677 | 1.8814 | 1591.5 | 1.5
1032.4678 | 1.8803 | 1591.5 | 1.5
1032.4677 | 1.8806 | 1591.5 | 1.5
1032.4677 | 1.8807 | 1591.5 | 1.5
1032.4678 | 1.8809 | 1591.5 | 1.5
1032.4678 | 1.8808 | 1591.5 | 1.5
1032.4678 | 1.8808 | 1591.5 | 1.5
1032.4678 | 1.8805 | 1591.5 | 1.5
1032.4678 | 1.8812 | 1591.5 | 1.5
1032.4677 | 1.8814 | 1591.5 | 1.5
1032.4677 | 1.881 | 1591.5 | 1.5
1032.4678 | 1.8809 | 1591.5 | 1.5
1032.4677 | 1.8803 | 1591.5 | 1.5
1032.4681 | 1.8779 | 1591.5 | 1.5
1032.4681 | 1.8773 | 1591.5 | 1.5
1032.4681 | 1.8769 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1032.4678 | 0.00013984 | 1.8805 | 0.0014275
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 1
2020-10-12 | 17:55
1044.2428 | 1.8791 | 1591.5 | 1.5
1044.2429 | 1.8798 | 1591.5 | 1.5
1044.2429 | 1.8793 | 1591.5 | 1.5
1044.2429 | 1.8797 | 1591.5 | 1.5
1044.243 | 1.8804 | 1591.5 | 1.5
1044.243 | 1.8802 | 1591.5 | 1.5
1044.2431 | 1.8804 | 1591.5 | 1.5
1044.243 | 1.8799 | 1591.5 | 1.5
1044.2431 | 1.8802 | 1591.5 | 1.5
1044.2431 | 1.8804 | 1591.5 | 1.5
1044.2432 | 1.8804 | 1591.5 | 1.5
1044.2432 | 1.8803 | 1591.5 | 1.5
1044.2431 | 1.8795 | 1591.5 | 1.5
1044.2432 | 1.8801 | 1591.5 | 1.5
1044.2432 | 1.8799 | 1591.5 | 1.5
1044.2432 | 1.88 | 1591.5 | 1.5
1044.2434 | 1.8814 | 1591.5 | 1.5
1044.2434 | 1.8812 | 1591.5 | 1.5
1044.2435 | 1.8826 | 1591.5 | 1.5
1044.2435 | 1.882 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1044.2431 | 0.00020482 | 1.8803 | 0.00085081
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 2
2020-10-12 | 17:57
2046.6228 | 2.202 | 1591.5 | 0.75
2046.6228 | 2.202 | 1591.5 | 0.75
2046.6229 | 2.202 | 1591.5 | 0.75
2046.6229 | 2.2016 | 1591.5 | 0.75
2046.6229 | 2.2018 | 1591.5 | 0.75
2046.6229 | 2.2023 | 1591.5 | 0.75
2046.6229 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.2019 | 1591.5 | 0.75
2046.6229 | 2.2019 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.6229 | 2.2018 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.623 | 2.2019 | 1591.5 | 0.75
2046.623 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.202 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.623 | 2.202 | 1591.5 | 0.75
2046.6231 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.2018 | 1591.5 | 0.75
2046.6231 | 2.2017 | 1591.5 | 0.75
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
2046.623 | 7.55e-05 | 2.2018 | 0.00016795
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 3
2020-10-12 | 17:59
3041.324 | 2.7767 | 1591.5 | 0.25
3041.3239 | 2.775 | 1591.5 | 0.25
3041.3241 | 2.7756 | 1591.5 | 0.25
3041.3241 | 2.7771 | 1591.5 | 0.25
3041.3241 | 2.7756 | 1591.5 | 0.25
3041.3241 | 2.7759 | 1591.5 | 0.25
3041.3242 | 2.775 | 1591.5 | 0.25
3041.3243 | 2.7754 | 1591.5 | 0.25
3041.3242 | 2.7756 | 1591.5 | 0.25
3041.3242 | 2.7756 | 1591.5 | 0.25
3041.3243 | 2.7758 | 1591.5 | 0.25
3041.3243 | 2.7756 | 1591.5 | 0.25
3041.3243 | 2.7743 | 1591.5 | 0.25
3041.3245 | 2.7765 | 1591.5 | 0.25
3041.3244 | 2.7756 | 1591.5 | 0.25
3041.3245 | 2.7763 | 1591.5 | 0.25
3041.3246 | 2.7766 | 1591.5 | 0.25
3041.3245 | 2.7753 | 1591.5 | 0.25
3041.3246 | 2.7759 | 1591.5 | 0.25
3041.3246 | 2.7772 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
3041.3243 | 0.0001986 | 2.7758 | 0.00070788
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 4
2020-10-12 | 18:01
4043.9172 | 3.217 | 1591.5 | 0.25
4043.9172 | 3.2155 | 1591.5 | 0.25
4043.9173 | 3.2154 | 1591.5 | 0.25
4043.9173 | 3.216 | 1591.5 | 0.25
4043.9172 | 3.2173 | 1591.5 | 0.25
4043.9172 | 3.2165 | 1591.5 | 0.25
4043.9174 | 3.2151 | 1591.5 | 0.25
4043.9173 | 3.2155 | 1591.5 | 0.25
4043.9175 | 3.2167 | 1591.5 | 0.25
4043.9175 | 3.2156 | 1591.5 | 0.25
4043.9176 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2163 | 1591.5 | 0.25
4043.9175 | 3.215 | 1591.5 | 0.25
4043.9177 | 3.2162 | 1591.5 | 0.25
4043.9177 | 3.2167 | 1591.5 | 0.25
4043.9175 | 3.2159 | 1591.5 | 0.25
4043.9177 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2158 | 1591.5 | 0.25
4043.9177 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2156 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
4043.9175 | 0.00020302 | 3.2162 | 0.00073231
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 5
2020-10-12 | 18:03
5047.278 | 3.7849 | 1591.5 | 0.25
5047.2783 | 3.7828 | 1591.5 | 0.25
5047.2782 | 3.7828 | 1591.5 | 0.25
5047.2783 | 3.786 | 1591.5 | 0.25
5047.2785 | 3.7852 | 1591.5 | 0.25
5047.2784 | 3.7849 | 1591.5 | 0.25
5047.2787 | 3.783 | 1591.5 | 0.25
5047.2787 | 3.785 | 1591.5 | 0.25
5047.2787 | 3.7848 | 1591.5 | 0.25
5047.2788 | 3.7848 | 1591.5 | 0.25
5047.2788 | 3.7841 | 1591.5 | 0.25
5047.2787 | 3.7827 | 1591.5 | 0.25
5047.2789 | 3.7848 | 1591.5 | 0.25
5047.2787 | 3.7848 | 1591.5 | 0.25
5047.279 | 3.7856 | 1591.5 | 0.25
5047.2789 | 3.7844 | 1591.5 | 0.25
5047.2792 | 3.7849 | 1591.5 | 0.25
5047.2792 | 3.7858 | 1591.5 | 0.25
5047.2793 | 3.7845 | 1591.5 | 0.25
5047.2791 | 3.7856 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
5047.2787 | 0.00034511 | 3.7846 | 0.00098189
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 6
2020-10-12 | 18:05
6041.457 | 4.3639 | 1591.5 | 0.25
6041.457 | 4.3627 | 1591.5 | 0.25
6041.4575 | 4.3638 | 1591.5 | 0.25
6041.4574 | 4.3633 | 1591.5 | 0.25
6041.4573 | 4.3641 | 1591.5 | 0.25
6041.4575 | 4.3642 | 1591.5 | 0.25
6041.4573 | 4.3651 | 1591.5 | 0.25
6041.4577 | 4.3633 | 1591.5 | 0.25
6041.4576 | 4.3622 | 1591.5 | 0.25
6041.4577 | 4.3636 | 1591.5 | 0.25
6041.4578 | 4.3638 | 1591.5 | 0.25
6041.4577 | 4.3653 | 1591.5 | 0.25
6041.4578 | 4.3661 | 1591.5 | 0.25
6041.4579 | 4.3618 | 1591.5 | 0.25
6041.458 | 4.3634 | 1591.5 | 0.25
6041.4578 | 4.3644 | 1591.5 | 0.25
6041.458 | 4.3636 | 1591.5 | 0.25
6041.4583 | 4.3628 | 1591.5 | 0.25
6041.4582 | 4.3627 | 1591.5 | 0.25
6041.4583 | 4.3618 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
6041.4577 | 0.00036621 | 4.3636 | 0.0010893
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 7
2020-10-12 | 18:08
7043.973 | 4.8055 | 1591.5 | 0.25
7043.9727 | 4.806 | 1591.5 | 0.25
7043.973 | 4.8041 | 1591.5 | 0.25
7043.973 | 4.8091 | 1591.5 | 0.25
7043.9729 | 4.8074 | 1591.5 | 0.25
7043.973 | 4.8036 | 1591.5 | 0.25
7043.9733 | 4.8065 | 1591.5 | 0.25
7043.973 | 4.8035 | 1591.5 | 0.25
7043.9731 | 4.8027 | 1591.5 | 0.25
7043.9735 | 4.8041 | 1591.5 | 0.25
7043.9733 | 4.8057 | 1591.5 | 0.25
7043.9735 | 4.8035 | 1591.5 | 0.25
7043.9733 | 4.8057 | 1591.5 | 0.25
7043.9735 | 4.8037 | 1591.5 | 0.25
7043.9735 | 4.8045 | 1591.5 | 0.25
7043.9738 | 4.805 | 1591.5 | 0.25
7043.9738 | 4.8067 | 1591.5 | 0.25
7043.9738 | 4.8046 | 1591.5 | 0.25
7043.9739 | 4.8079 | 1591.5 | 0.25
7043.9741 | 4.8051 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
7043.9733 | 0.00037834 | 4.8052 | 0.0016222
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 8
2020-10-12 | 18:10
8037.9912 | 5.5 | 1591.5 | 0.25
8037.991 | 5.506 | 1591.5 | 0.25
8037.9916 | 5.49 | 1591.5 | 0.25
8037.9915 | 5.494 | 1591.5 | 0.25
8037.9916 | 5.497 | 1591.5 | 0.25
8037.9911 | 5.504 | 1591.5 | 0.25
8037.9915 | 5.502 | 1591.5 | 0.25
8037.9918 | 5.502 | 1591.5 | 0.25
8037.9909 | 5.512 | 1591.5 | 0.25
8037.9916 | 5.505 | 1591.5 | 0.25
8037.992 | 5.501 | 1591.5 | 0.25
8037.9916 | 5.519 | 1591.5 | 0.25
8037.9922 | 5.505 | 1591.5 | 0.25
8037.9919 | 5.498 | 1591.5 | 0.25
8037.9918 | 5.51 | 1591.5 | 0.25
8037.9923 | 5.504 | 1591.5 | 0.25
8037.992 | 5.514 | 1591.5 | 0.25
8037.9911 | 5.525 | 1591.5 | 0.25
8037.9921 | 5.511 | 1591.5 | 0.25
8037.9922 | 5.51 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
8037.9916 | 0.00041509 | 5.5054 | 0.0081515
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 9
2020-10-12 | 18:12
9040.2661 | 5.957 | 1591.5 | 0.1
9040.2664 | 5.948 | 1591.5 | 0.1
9040.2657 | 5.947 | 1591.5 | 0.1
9040.2659 | 5.947 | 1591.5 | 0.1
9040.2662 | 5.954 | 1591.5 | 0.1
9040.266 | 5.95 | 1591.5 | 0.1
9040.2663 | 5.945 | 1591.5 | 0.1
9040.2665 | 5.95 | 1591.5 | 0.1
9040.2656 | 5.954 | 1591.5 | 0.1
9040.2672 | 5.952 | 1591.5 | 0.1
9040.2667 | 5.952 | 1591.5 | 0.1
9040.2667 | 5.942 | 1591.5 | 0.1
9040.2669 | 5.949 | 1591.5 | 0.1
9040.2666 | 5.94 | 1591.5 | 0.1
9040.2667 | 5.95 | 1591.5 | 0.1
9040.2665 | 5.952 | 1591.5 | 0.1
9040.2662 | 5.949 | 1591.5 | 0.1
9040.2667 | 5.946 | 1591.5 | 0.1
9040.2668 | 5.949 | 1591.5 | 0.1
9040.2672 | 5.956 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
9040.2665 | 0.00043352 | 5.9494 | 0.0042009
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail X
2020-10-12 | 18:14
10034.2262 | 6.724 | 1591.5 | 0.1
10034.2263 | 6.714 | 1591.5 | 0.1
10034.2256 | 6.722 | 1591.5 | 0.1
10034.2259 | 6.714 | 1591.5 | 0.1
10034.2264 | 6.716 | 1591.5 | 0.1
10034.2269 | 6.707 | 1591.5 | 0.1
10034.2267 | 6.717 | 1591.5 | 0.1
10034.2266 | 6.73 | 1591.5 | 0.1
10034.2257 | 6.717 | 1591.5 | 0.1
10034.2269 | 6.721 | 1591.5 | 0.1
10034.2272 | 6.723 | 1591.5 | 0.1
10034.2272 | 6.716 | 1591.5 | 0.1
10034.2279 | 6.719 | 1591.5 | 0.1
10034.2263 | 6.726 | 1591.5 | 0.1
10034.2275 | 6.723 | 1591.5 | 0.1
10034.2272 | 6.72 | 1591.5 | 0.1
10034.2279 | 6.723 | 1591.5 | 0.1
10034.2283 | 6.708 | 1591.5 | 0.1
10034.2273 | 6.717 | 1591.5 | 0.1
10034.2287 | 6.727 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10034.2269 | 0.00082721 | 6.7192 | 0.0057671
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 1
2020-10-12 | 18:17
10060.9467 | 6.713 | 1591.5 | 0.1
10060.9462 | 6.713 | 1591.5 | 0.1
10060.9466 | 6.713 | 1591.5 | 0.1
10060.9469 | 6.715 | 1591.5 | 0.1
10060.9464 | 6.714 | 1591.5 | 0.1
10060.9477 | 6.718 | 1591.5 | 0.1
10060.9475 | 6.718 | 1591.5 | 0.1
10060.9474 | 6.714 | 1591.5 | 0.1
10060.9479 | 6.72 | 1591.5 | 0.1
10060.9464 | 6.731 | 1591.5 | 0.1
10060.9484 | 6.714 | 1591.5 | 0.1
10060.9474 | 6.714 | 1591.5 | 0.1
10060.9478 | 6.71 | 1591.5 | 0.1
10060.9483 | 6.702 | 1591.5 | 0.1
10060.9473 | 6.719 | 1591.5 | 0.1
10060.9483 | 6.729 | 1591.5 | 0.1
10060.948 | 6.712 | 1591.5 | 0.1
10060.9485 | 6.709 | 1591.5 | 0.1
10060.9485 | 6.72 | 1591.5 | 0.1
10060.9494 | 6.723 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10060.9476 | 0.0008394 | 6.7161 | 0.0064535
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 2
2020-10-12 | 18:19
20052.3176 | 14.119 | 1591.5 | 0.1
20052.3172 | 14.11 | 1591.5 | 0.1
20052.3163 | 14.109 | 1591.5 | 0.1
20052.3184 | 14.111 | 1591.5 | 0.1
20052.3179 | 14.117 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
20052.3178 | 14.113 | 1591.5 | 0.1
20052.3201 | 14.11 | 1591.5 | 0.1
20052.3199 | 14.139 | 1591.5 | 0.1
20052.319 | 14.126 | 1591.5 | 0.1
20052.32 | 14.121 | 1591.5 | 0.1
20052.3183 | 14.108 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
20052.3199 | 14.121 | 1591.5 | 0.1
20052.3209 | 14.114 | 1591.5 | 0.1
20052.3183 | 14.123 | 1591.5 | 0.1
20052.3208 | 14.112 | 1591.5 | 0.1
20052.3215 | 14.137 | 1591.5 | 0.1
20052.3213 | 14.137 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
20052.3191 | 0.0013909 | 14.1209 | 0.0099412
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 3
2020-10-12 | 18:22
30063.9944 | 21.56 | 1591.5 | 0.03
30063.987 | 21.56 | 1591.5 | 0.03
30063.989 | 21.57 | 1591.5 | 0.03
30063.984 | 21.59 | 1591.5 | 0.03
30063.989 | 21.59 | 1591.5 | 0.03
30063.991 | 21.6 | 1591.5 | 0.03
30063.9861 | 21.6 | 1591.5 | 0.03
30063.999 | 21.63 | 1591.5 | 0.03
30063.991 | 21.51 | 1591.5 | 0.03
30063.993 | 21.57 | 1591.5 | 0.03
30063.9966 | 21.64 | 1591.5 | 0.03
30063.999 | 21.69 | 1591.5 | 0.03
30063.993 | 21.63 | 1591.5 | 0.03
30063.9883 | 21.56 | 1591.5 | 0.03
30063.9997 | 21.56 | 1591.5 | 0.03
30063.991 | 21.61 | 1591.5 | 0.03
30063.9972 | 21.59 | 1591.5 | 0.03
30063.988 | 21.74 | 1591.5 | 0.03
30064.008 | 21.52 | 1591.5 | 0.03
30063.997 | 21.6 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
30063.9931 | 0.0056836 | 21.596 | 0.051904
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 4
2020-10-12 | 18:24
40066.92 | 32.15 | 1591.5 | 0.03
40066.918 | 32.12 | 1591.5 | 0.03
40066.93 | 32.1 | 1591.5 | 0.03
40066.9219 | 32.05 | 1591.5 | 0.03
40066.917 | 32.03 | 1591.5 | 0.03
40066.918 | 32 | 1591.5 | 0.03
40066.926 | 32.04 | 1591.5 | 0.03
40066.9222 | 32.17 | 1591.5 | 0.03
40066.917 | 32.15 | 1591.5 | 0.03
40066.928 | 32.04 | 1591.5 | 0.03
40066.923 | 32.07 | 1591.5 | 0.03
40066.941 | 31.94 | 1591.5 | 0.03
40066.927 | 32.1 | 1591.5 | 0.03
40066.917 | 31.98 | 1591.5 | 0.03
40066.941 | 32.12 | 1591.5 | 0.03
40066.9245 | 32.01 | 1591.5 | 0.03
40066.938 | 31.96 | 1591.5 | 0.03
40066.929 | 32 | 1591.5 | 0.03
40066.929 | 32.07 | 1591.5 | 0.03
40066.926 | 31.96 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
40066.9257 | 0.00735 | 32.053 | 0.067164
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 5
2020-10-12 | 18:27
50072.953 | 44.52 | 1591.5 | 0.03
50072.948 | 44.6 | 1591.5 | 0.03
50072.953 | 44.57 | 1591.5 | 0.03
50072.964 | 44.57 | 1591.5 | 0.03
50072.953 | 44.7 | 1591.5 | 0.03
50072.932 | 44.73 | 1591.5 | 0.03
50072.9412 | 44.8 | 1591.5 | 0.03
50072.95 | 44.57 | 1591.5 | 0.03
50072.952 | 44.67 | 1591.5 | 0.03
50072.942 | 44.72 | 1591.5 | 0.03
50072.955 | 44.64 | 1591.5 | 0.03
50072.943 | 44.69 | 1591.5 | 0.03
50072.943 | 44.63 | 1591.5 | 0.03
50072.955 | 44.66 | 1591.5 | 0.03
50072.959 | 44.69 | 1591.5 | 0.03
50072.949 | 44.6 | 1591.5 | 0.03
50072.955 | 44.6 | 1591.5 | 0.03
50072.95 | 44.6 | 1591.5 | 0.03
50072.952 | 44.67 | 1591.5 | 0.03
50072.9596 | 44.72 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
50072.9504 | 0.007224 | 44.6475 | 0.06774
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 6
2020-10-12 | 18:29
60084.285 | 58.6 | 1591.5 | 0.03
60084.301 | 58.7 | 1591.5 | 0.03
60084.301 | 58.4 | 1591.5 | 0.03
60084.313 | 58.7 | 1591.5 | 0.03
60084.294 | 58.5 | 1591.5 | 0.03
60084.3 | 58.5 | 1591.5 | 0.03
60084.287 | 58.6 | 1591.5 | 0.03
60084.291 | 58.5 | 1591.5 | 0.03
60084.294 | 58.5 | 1591.5 | 0.03
60084.29 | 58.4 | 1591.5 | 0.03
60084.297 | 58.6 | 1591.5 | 0.03
60084.3 | 58.5 | 1591.5 | 0.03
60084.303 | 58.4 | 1591.5 | 0.03
60084.306 | 58.5 | 1591.5 | 0.03
60084.305 | 58.5 | 1591.5 | 0.03
60084.278 | 58.5 | 1591.5 | 0.03
60084.318 | 58.6 | 1591.5 | 0.03
60084.291 | 58.5 | 1591.5 | 0.03
60084.289 | 58.3 | 1591.5 | 0.03
60084.303 | 58.4 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
60084.2973 | 0.009408 | 58.51 | 0.099499
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 7
2020-10-12 | 18:35
70075.952 | 73.1 | 1591.5 | 0.01
70075.902 | 73 | 1591.5 | 0.01
70075.916 | 72.9 | 1591.5 | 0.01
70075.941 | 73 | 1591.5 | 0.01
70075.959 | 73.1 | 1591.5 | 0.01
70075.993 | 72.8 | 1591.5 | 0.01
70075.878 | 73.3 | 1591.5 | 0.01
70075.95 | 72.9 | 1591.5 | 0.01
70075.889 | 72.9 | 1591.5 | 0.01
70075.88 | 72.8 | 1591.5 | 0.01
70075.94 | 73.3 | 1591.5 | 0.01
70075.905 | 73.7 | 1591.5 | 0.01
70075.995 | 73.7 | 1591.5 | 0.01
70075.952 | 73 | 1591.5 | 0.01
70075.96 | 72.8 | 1591.5 | 0.01
70075.95 | 73.2 | 1591.5 | 0.01
70075.952 | 73.2 | 1591.5 | 0.01
70075.947 | 73.2 | 1591.5 | 0.01
70075.889 | 72.9 | 1591.5 | 0.01
70075.959 | 72.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
70075.9355 | 0.033996 | 73.07 | 0.27586
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 8
2020-10-12 | 18:37
80087.32 | 91.8 | 1591.5 | 0.01
80087.275 | 91.6 | 1591.5 | 0.01
80087.293 | 91.3 | 1591.5 | 0.01
80087.279 | 91.4 | 1591.5 | 0.01
80087.304 | 91.7 | 1591.5 | 0.01
80087.3 | 91.7 | 1591.5 | 0.01
80087.25 | 91.5 | 1591.5 | 0.01
80087.349 | 91.6 | 1591.5 | 0.01
80087.34 | 91.8 | 1591.5 | 0.01
80087.327 | 90.6 | 1591.5 | 0.01
80087.349 | 92.1 | 1591.5 | 0.01
80087.32 | 91.5 | 1591.5 | 0.01
80087.24 | 92.2 | 1591.5 | 0.01
80087.361 | 91 | 1591.5 | 0.01
80087.261 | 91.4 | 1591.5 | 0.01
80087.36 | 91.6 | 1591.5 | 0.01
80087.354 | 92 | 1591.5 | 0.01
80087.421 | 91.1 | 1591.5 | 0.01
80087.331 | 91.7 | 1591.5 | 0.01
80087.329 | 91.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
80087.3182 | 0.04288 | 91.56 | 0.36387
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 9
2020-10-12 | 18:39
90089.851 | 111.4 | 1591.5 | 0.01
90089.93 | 111.4 | 1591.5 | 0.01
90089.85 | 110.7 | 1591.5 | 0.01
90089.93 | 111.6 | 1591.5 | 0.01
90089.98 | 111.5 | 1591.5 | 0.01
90089.97 | 111 | 1591.5 | 0.01
90089.95 | 110.5 | 1591.5 | 0.01
90089.982 | 110.8 | 1591.5 | 0.01
90089.915 | 111.4 | 1591.5 | 0.01
90089.95 | 110.8 | 1591.5 | 0.01
90089.964 | 110.6 | 1591.5 | 0.01
90089.885 | 110.6 | 1591.5 | 0.01
90089.88 | 111 | 1591.5 | 0.01
90089.95 | 110.9 | 1591.5 | 0.01
90089.87 | 110.8 | 1591.5 | 0.01
90089.885 | 110.6 | 1591.5 | 0.01
90089.99 | 110.8 | 1591.5 | 0.01
90089.984 | 110.9 | 1591.5 | 0.01
90089.88 | 110.3 | 1591.5 | 0.01
90089.93 | 110.1 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
90089.9263 | 0.045232 | 110.885 | 0.39532
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail X
2020-10-12 | 18:41
100101.41 | 133.6 | 1591.5 | 0.01
100101.376 | 132.9 | 1591.5 | 0.01
100101.365 | 133.3 | 1591.5 | 0.01
100101.378 | 132.9 | 1591.5 | 0.01
100101.426 | 133 | 1591.5 | 0.01
100101.41 | 134 | 1591.5 | 0.01
100101.376 | 132.9 | 1591.5 | 0.01
100101.42 | 133 | 1591.5 | 0.01
100101.426 | 133.8 | 1591.5 | 0.01
100101.398 | 133.3 | 1591.5 | 0.01
100101.43 | 132.3 | 1591.5 | 0.01
100101.33 | 132.2 | 1591.5 | 0.01
100101.34 | 133.4 | 1591.5 | 0.01
100101.412 | 132.1 | 1591.5 | 0.01
100101.419 | 134.1 | 1591.5 | 0.01
100101.322 | 132.9 | 1591.5 | 0.01
100101.354 | 132.9 | 1591.5 | 0.01
100101.36 | 132.9 | 1591.5 | 0.01
100101.356 | 132.4 | 1591.5 | 0.01
100101.37 | 133.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100101.3839 | 0.033297 | 133.075 | 0.55486
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 1
2020-10-12 | 18:44
Capacitor description: | GR sn:4556 0.1uF Dail 1
2020-10-12 | 18:45
100016.06 | 136.1 | 1591.5 | 0.01
100015.938 | 136.2 | 1591.5 | 0.01
100015.84 | 135.1 | 1591.5 | 0.01
100016.02 | 134.9 | 1591.5 | 0.01
100015.907 | 135.2 | 1591.5 | 0.01
100015.977 | 135.1 | 1591.5 | 0.01
100016 | 135.3 | 1591.5 | 0.01
100015.91 | 134.8 | 1591.5 | 0.01
100015.8 | 136 | 1591.5 | 0.01
100016.01 | 136.2 | 1591.5 | 0.01
100015.99 | 135.7 | 1591.5 | 0.01
100015.94 | 134.5 | 1591.5 | 0.01
100015.972 | 136 | 1591.5 | 0.01
100015.98 | 136.1 | 1591.5 | 0.01
100015.94 | 136 | 1591.5 | 0.01
100015.93 | 135.3 | 1591.5 | 0.01
100015.968 | 136 | 1591.5 | 0.01
100015.82 | 136.2 | 1591.5 | 0.01
100015.938 | 135 | 1591.5 | 0.01
100016 | 135.2 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100015.947 | 0.065305 | 135.545 | 0.54173
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 2
2020-10-12 | 18:47
200421.6 | 483.7 | 1591.5 | 0.01
200421.49 | 483.7 | 1591.5 | 0.01
200421.76 | 482.8 | 1591.5 | 0.01
200421.8 | 483.7 | 1591.5 | 0.01
200421.71 | 483.5 | 1591.5 | 0.01
200421.86 | 483 | 1591.5 | 0.01
200421.62 | 481.2 | 1591.5 | 0.01
200421.78 | 482.9 | 1591.5 | 0.01
200421.81 | 482.1 | 1591.5 | 0.01
200421.61 | 483.3 | 1591.5 | 0.01
200421.7 | 483.8 | 1591.5 | 0.01
200421.51 | 481.2 | 1591.5 | 0.01
200421.82 | 482.8 | 1591.5 | 0.01
200421.71 | 481.5 | 1591.5 | 0.01
200421.74 | 482.8 | 1591.5 | 0.01
200421.82 | 481.5 | 1591.5 | 0.01
200421.54 | 484.1 | 1591.5 | 0.01
200421.92 | 482.6 | 1591.5 | 0.01
200421.69 | 483.5 | 1591.5 | 0.01
200421.75 | 481.9 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
200421.712 | 0.11548 | 482.78 | 0.90089
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 3
2020-10-12 | 18:49
300392.5 | 941 | 1591.5 | 0.003
300392.1 | 931 | 1591.5 | 0.003
300391.8 | 938 | 1591.5 | 0.003
300392.25 | 932 | 1591.5 | 0.003
300391.8 | 934 | 1591.5 | 0.003
300392.07 | 933 | 1591.5 | 0.003
300392 | 938 | 1591.5 | 0.003
300392.3 | 938 | 1591.5 | 0.003
300392.7 | 939 | 1591.5 | 0.003
300391.5 | 930 | 1591.5 | 0.003
300391.3 | 939 | 1591.5 | 0.003
300392.5 | 937 | 1591.5 | 0.003
300392.6 | 933 | 1591.5 | 0.003
300392.5 | 935 | 1591.5 | 0.003
300392.1 | 944 | 1591.5 | 0.003
300391.6 | 942 | 1591.5 | 0.003
300392.4 | 939 | 1591.5 | 0.003
300391.9 | 930 | 1591.5 | 0.003
300392.9 | 938 | 1591.5 | 0.003
300391.9 | 940 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
300392.136 | 0.41094 | 936.55 | 3.9683
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 4
2020-10-12 | 18:51
400792 | 1669 | 1591.5 | 0.003
400792.28 | 1673 | 1591.5 | 0.003
400792.2 | 1677 | 1591.5 | 0.003
400792.1 | 1668 | 1591.5 | 0.003
400792.42 | 1667 | 1591.5 | 0.003
400790.9 | 1664 | 1591.5 | 0.003
400791.85 | 1668 | 1591.5 | 0.003
400791.1 | 1666 | 1591.5 | 0.003
400791.48 | 1671 | 1591.5 | 0.003
400792.4 | 1671 | 1591.5 | 0.003
400790.8 | 1675 | 1591.5 | 0.003
400790.4 | 1668 | 1591.5 | 0.003
400792.18 | 1665 | 1591.5 | 0.003
400791.2 | 1665 | 1591.5 | 0.003
400792.3 | 1673 | 1591.5 | 0.003
400791.5 | 1666 | 1591.5 | 0.003
400791.3 | 1680 | 1591.5 | 0.003
400791.7 | 1679 | 1591.5 | 0.003
400792.2 | 1668 | 1591.5 | 0.003
400793.05 | 1674 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
400791.768 | 0.64947 | 1670.35 | 4.6505
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 5
2020-10-12 | 18:54
500613.2 | 2636 | 1591.5 | 0.003
500612.7 | 2621 | 1591.5 | 0.003
500611.5 | 2629 | 1591.5 | 0.003
500612.7 | 2630 | 1591.5 | 0.003
500612.2 | 2636 | 1591.5 | 0.003
500612 | 2632 | 1591.5 | 0.003
500612 | 2632 | 1591.5 | 0.003
500613.6 | 2635 | 1591.5 | 0.003
500612.9 | 2635 | 1591.5 | 0.003
500612.6 | 2633 | 1591.5 | 0.003
500612.1 | 2626 | 1591.5 | 0.003
500611.4 | 2621 | 1591.5 | 0.003
500613.1 | 2623 | 1591.5 | 0.003
500611.8 | 2639 | 1591.5 | 0.003
500612.1 | 2635 | 1591.5 | 0.003
500611.4 | 2632 | 1591.5 | 0.003
500613.4 | 2638 | 1591.5 | 0.003
500612.9 | 2625 | 1591.5 | 0.003
500612.5 | 2636 | 1591.5 | 0.003
500611.5 | 2637 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
500612.38 | 0.66378 | 2631.55 | 5.4907
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 6
2020-10-12 | 18:56
600591.9 | 3681 | 1591.5 | 0.003
600591.6 | 3673 | 1591.5 | 0.003
600590.8 | 3670 | 1591.5 | 0.003
600592.8 | 3681 | 1591.5 | 0.003
600592.3 | 3670 | 1591.5 | 0.003
600593.2 | 3673 | 1591.5 | 0.003
600591.8 | 3679 | 1591.5 | 0.003
600591.7 | 3680 | 1591.5 | 0.003
600591.1 | 3691 | 1591.5 | 0.003
600591.9 | 3691 | 1591.5 | 0.003
600592 | 3674 | 1591.5 | 0.003
600592.9 | 3671 | 1591.5 | 0.003
600592 | 3670 | 1591.5 | 0.003
600591.1 | 3656 | 1591.5 | 0.003
600592.6 | 3679 | 1591.5 | 0.003
600592.4 | 3673 | 1591.5 | 0.003
600590.24 | 3666 | 1591.5 | 0.003
600590.8 | 3662 | 1591.5 | 0.003
600591.8 | 3675 | 1591.5 | 0.003
600591.3 | 3670 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
600591.812 | 0.74936 | 3674.25 | 8.2757
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 7
2020-10-12 | 18:58
701008.3 | 4833 | 1591.5 | 0.001
701004.8 | 4869 | 1591.5 | 0.001
701004 | 4900 | 1591.5 | 0.001
701007.3 | 4870 | 1591.5 | 0.001
701005 | 4880 | 1591.5 | 0.001
701005 | 4890 | 1591.5 | 0.001
701006.8 | 4938 | 1591.5 | 0.001
701006 | 4810 | 1591.5 | 0.001
701007.7 | 4870 | 1591.5 | 0.001
701006.4 | 4846 | 1591.5 | 0.001
701005.3 | 4870 | 1591.5 | 0.001
701010 | 4880 | 1591.5 | 0.001
701004.1 | 4950 | 1591.5 | 0.001
701008.9 | 4855 | 1591.5 | 0.001
701010.9 | 4864 | 1591.5 | 0.001
701003.4 | 4920 | 1591.5 | 0.001
701001.4 | 4922 | 1591.5 | 0.001
701003 | 4880 | 1591.5 | 0.001
701007.7 | 4950 | 1591.5 | 0.001
701009 | 4870 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
701006.25 | 2.4309 | 4883.35 | 36.3803
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 8
2020-10-12 | 19:00
800991.9 | 6460 | 1591.5 | 0.001
800983.6 | 6360 | 1591.5 | 0.001
800991.9 | 6360 | 1591.5 | 0.001
800990 | 6400 | 1591.5 | 0.001
800990.9 | 6370 | 1591.5 | 0.001
800992.4 | 6400 | 1591.5 | 0.001
800988.3 | 6320 | 1591.5 | 0.001
800994.4 | 6340 | 1591.5 | 0.001
801000 | 6410 | 1591.5 | 0.001
800991.9 | 6330 | 1591.5 | 0.001
800989.5 | 6390 | 1591.5 | 0.001
800995.3 | 6420 | 1591.5 | 0.001
800985.8 | 6410 | 1591.5 | 0.001
800982.7 | 6390 | 1591.5 | 0.001
800990 | 6350 | 1591.5 | 0.001
800987.6 | 6350 | 1591.5 | 0.001
800997.6 | 6320 | 1591.5 | 0.001
800993.7 | 6370 | 1591.5 | 0.001
800994.2 | 6390 | 1591.5 | 0.001
800992.8 | 6440 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
800991.225 | 4.1986 | 6379 | 37.8021
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 9
2020-10-12 | 19:02
901406 | 7970 | 1591.5 | 0.001
901403.7 | 8020 | 1591.5 | 0.001
901397 | 8060 | 1591.5 | 0.001
901404 | 7950 | 1591.5 | 0.001
901402 | 8020 | 1591.5 | 0.001
901396.7 | 7990 | 1591.5 | 0.001
901405.7 | 7970 | 1591.5 | 0.001
901405 | 7980 | 1591.5 | 0.001
901406 | 8090 | 1591.5 | 0.001
901399 | 8040 | 1591.5 | 0.001
901405.9 | 7990 | 1591.5 | 0.001
901394 | 7970 | 1591.5 | 0.001
901401 | 8080 | 1591.5 | 0.001
901409.5 | 8040 | 1591.5 | 0.001
901403.2 | 8080 | 1591.5 | 0.001
901399.7 | 8040 | 1591.5 | 0.001
901412 | 8000 | 1591.5 | 0.001
901404.1 | 7980 | 1591.5 | 0.001
901404 | 8090 | 1591.5 | 0.001
901402.9 | 8050 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
901403.07 | 4.2023 | 8020.5 | 43.872
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail X
2020-10-12 | 19:05
1001394 | 9950 | 1591.5 | 0.001
1001381.8 | 9870 | 1591.5 | 0.001
1001392 | 10040 | 1591.5 | 0.001
1001397.2 | 9970 | 1591.5 | 0.001
1001389 | 10010 | 1591.5 | 0.001
1001394.9 | 9990 | 1591.5 | 0.001
1001383 | 9990 | 1591.5 | 0.001
1001388.1 | 10020 | 1591.5 | 0.001
1001394 | 9940 | 1591.5 | 0.001
1001392 | 9930 | 1591.5 | 0.001
1001383.4 | 10000 | 1591.5 | 0.001
1001397 | 9940 | 1591.5 | 0.001
1001389 | 10010 | 1591.5 | 0.001
1001395.8 | 9960 | 1591.5 | 0.001
1001396 | 9940 | 1591.5 | 0.001
1001389.3 | 9970 | 1591.5 | 0.001
1001394 | 9970 | 1591.5 | 0.001
1001397.8 | 10020 | 1591.5 | 0.001
1001391 | 9920 | 1591.5 | 0.001
1001397 | 10000 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1001391.815 | 4.8028 | 9972 | 40.5709
end | of | the | test

## AH2700_cal
Start with internal cal.
AH2700 will be calibrated 'on the day' against known reference capacitors.
Use these results to have a correction factor adequate for this calibration.
[7 empty] | subtract zero |  |  |  | Value from capacitance scale, out.csv |  |  | measured | NMIA
| Reference | C | sdC | G | sdG |  | C | sdC | G | sdG | C | mean | stdev | tan delta | G | tan delta
|  | pF | pF | nS | nS |  | pF | pF | nS | nS | pF | -2.483 | 0.13805 |  |  |  | correction
| zero | -1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06 |  | 0 | 1.9233e-07 | 0 | 2.7294e-06 |  | ppm correction |  | mrad | nS | mrad | mrad
| AH111A | 10 | 1.8e-07 | -0.00012269 | 2.05e-06 |  | 10 | 2.256e-07 | -0.00014989 | 2.8156e-06 | 9.9999 | -2.4644 |  | -0.0014989 | 0.00018 | 0.0018 | 0.0032989
| AH111B | 10 | 2.65e-07 | -6.61e-05 | 2.22e-06 |  | 10 | 2.9786e-07 | -9.33e-05 | 2.9416e-06 | 10 | -2.4315 |  | -0.000933 | 0.00024 | 0.0024 | 0.003333
| AH111C | 99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06 |  | 99.9998 | 2.1693e-07 | 0.0016097 | 2.4943e-06 | 99.9996 | -2.3617 |  | 0.0016097 | 0.0019 | 0.0019 | 0.0002903
| AH111D | 99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06 |  | 99.9998 | 2.7412e-07 | 0.0020538 | 2.8595e-06 | 99.9996 | -2.3869 |  | 0.0020538 | 0.0023 | 0.0023 | 0.00024625
| GR1000A | 1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05 |  | 1000.0317 | 6.4014e-06 | 0.17933 | 5.4534e-05 | 1000.0289 | -2.7431 |  | 0.017932
| GR1000B | 1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05 |  | 1000.0164 | 5.9216e-06 | 0.18442 | 5.7932e-05 | 1000.0139 | -2.5101 |  | 0.018442
| UB1000 | 1000.1471 | 5.89e-06 | 0.142 | 9.76e-05 |  | 1000.1471 | 5.8916e-06 | 0.14197 | 9.7619e-05 |  |  |  | 0.014195
corrected
[6 empty] | UB1000 = | 1000.1446

## pyAH2700readings
Summary of capacitance box measurements extracted from pasted csv file in AH2700_data worksheet.
subtract from relevant zero
| Setting | C | sdC | G | sdG | Corrected | C | G | tan delta
Row |  | pF | pF | nS | nS | pF | pF | nS
235 | main_zero | 49.2216 | 0.00043741 | 1.3317 | 0.0053142 | 49.2215 | 49.2215 | 1.3317 | 0.0027056
261 | E1 | 59.6543 | 0.00014858 | 1.3377 | 0.0022088 | 59.6542 | 10.4327 | 0.0059616 | 5.7143e-05
287 | E2 | 70.2234 | 6.02e-05 | 1.3687 | 0.0020591 | 70.2232 | 21.0017 | 0.036932 | 0.00017585
313 | E3 | 80.8288 | 0.00048247 | 1.3481 | 0.0052107 | 80.8286 | 31.6071 | 0.01633 | 5.1665e-05
339 | E4 | 90.9238 | 0.00021909 | 1.35 | 0.00079946 | 90.9235 | 41.702 | 0.018306 | 4.3898e-05
365 | E5 | 100.4774 | 0.00036577 | 1.3441 | 0.0038393 | 100.4771 | 51.2556 | 0.012348 | 2.4091e-05
391 | E6 | 110.6912 | 0.00050256 | 1.3455 | 0.00079577 | 110.6909 | 61.4694 | 0.013767 | 2.2396e-05
417 | E7 | 120.5081 | 0.00029987 | 1.3463 | 0.00018653 | 120.5078 | 71.2863 | 0.014561 | 2.0426e-05
443 | E8 | 129.974 | 5.59e-05 | 1.3525 | 0.00082738 | 129.9736 | 80.7521 | 0.020738 | 2.5682e-05
469 | E9 | 139.7386 | 0.0005509 | 1.3521 | 0.0006845 | 139.7382 | 90.5167 | 0.020384 | 2.252e-05
495 | E10 | 149.5767 | 0.00025146 | 1.3555 | 0.00075452 | 149.5763 | 100.3548 | 0.023795 | 2.3711e-05
521 | D1 | 145.8887 | 0.00046896 | 1.6176 | 0.015335 | 145.8884 | 96.6669 | 0.28585 | 0.0002957
547 | D2 | 244.9349 | 0.00010611 | 1.4339 | 0.0016017 | 244.9343 | 195.7128 | 0.10215 | 5.2193e-05
573 | D3 | 341.4798 | 0.00032198 | 1.5154 | 0.0029831 | 341.479 | 292.2575 | 0.18372 | 6.2861e-05
599 | D4 | 441.2625 | 0.0001467 | 1.5405 | 0.0012842 | 441.2614 | 392.0399 | 0.20879 | 5.3258e-05
625 | D5 | 537.7661 | 6.53e-05 | 1.6209 | 0.0006773 | 537.7648 | 488.5433 | 0.28917 | 5.9191e-05
651 | D6 | 640.7041 | 0.00010485 | 1.6665 | 0.00040726 | 640.7025 | 591.481 | 0.33476 | 5.6597e-05
677 | D7 | 739.8078 | 0.00025568 | 1.687 | 0.0016552 | 739.806 | 690.5845 | 0.35524 | 5.1441e-05
703 | D8 | 836.3521 | 0.00010751 | 1.7676 | 0.00037086 | 836.35 | 787.1285 | 0.43591 | 5.5379e-05
729 | D9 | 936.1413 | 0.00025167 | 1.7886 | 0.00083413 | 936.1389 | 886.9174 | 0.45689 | 5.1515e-05
755 | D10 | 1032.4678 | 0.00013984 | 1.8805 | 0.0014275 | 1032.4652 | 983.2437 | 0.54878 | 5.5814e-05
781 | C1 | 1044.2431 | 0.00020482 | 1.8803 | 0.00085081 | 1044.2405 | 995.019 | 0.54861 | 5.5136e-05
807 | C2 | 2046.623 | 7.55e-05 | 2.2018 | 0.00016795 | 2046.6179 | 1997.3964 | 0.87012 | 4.3563e-05
833 | C3 | 3041.3243 | 0.0001986 | 2.7758 | 0.00070788 | 3041.3167 | 2992.0952 | 1.4441 | 4.8264e-05
859 | C4 | 4043.9175 | 0.00020302 | 3.2162 | 0.00073231 | 4043.9074 | 3994.6859 | 1.8845 | 4.7174e-05
885 | C5 | 5047.2787 | 0.00034511 | 3.7846 | 0.00098189 | 5047.2662 | 4998.0447 | 2.4528 | 4.9076e-05
911 | C6 | 6041.4577 | 0.00036621 | 4.3636 | 0.0010893 | 6041.4427 | 5992.2212 | 3.0319 | 5.0597e-05
937 | C7 | 7043.9733 | 0.00037834 | 4.8052 | 0.0016222 | 7043.9558 | 6994.7343 | 3.4735 | 4.9659e-05
963 | C8 | 8037.9916 | 0.00041509 | 5.5054 | 0.0081515 | 8037.9717 | 7988.7502 | 4.1737 | 5.2245e-05
989 | C9 | 9040.2665 | 0.00043352 | 5.9494 | 0.0042009 | 9040.244 | 8991.0225 | 4.6177 | 5.1359e-05
1015 | C10 | 10034.2269 | 0.00082721 | 6.7192 | 0.0057671 | 10034.202 | 9984.9805 | 5.3875 | 5.3956e-05
1041 | B1 | 10060.9476 | 0.0008394 | 6.7161 | 0.0064535 | 10060.9226 | 10011.7011 | 5.3843 | 5.378e-05
1067 | B2 | 20052.3191 | 0.0013909 | 14.1209 | 0.0099412 | 20052.2693 | 20003.0478 | 12.7891 | 6.3936e-05
1093 | B3 | 30063.9931 | 0.0056836 | 21.596 | 0.051904 | 30063.9184 | 30014.6969 | 20.2643 | 6.7515e-05
1119 | B4 | 40066.9257 | 0.00735 | 32.053 | 0.067164 | 40066.8262 | 40017.6047 | 30.7213 | 7.6769e-05
1145 | B5 | 50072.9504 | 0.007224 | 44.6475 | 0.06774 | 50072.8261 | 50023.6046 | 43.3158 | 8.6591e-05
1171 | B6 | 60084.2973 | 0.009408 | 58.51 | 0.099499 | 60084.1481 | 60034.9266 | 57.1783 | 9.5242e-05
1197 | B7 | 70075.9355 | 0.033996 | 73.07 | 0.27586 | 70075.7615 | 70026.5399 | 71.7383 | 0.00010244
1223 | B8 | 80087.3182 | 0.04288 | 91.56 | 0.36387 | 80087.1193 | 80037.8978 | 90.2283 | 0.00011273
1249 | B9 | 90089.9263 | 0.045232 | 110.885 | 0.39532 | 90089.7026 | 90040.4811 | 109.5533 | 0.00012167
1275 | B10 | 100101.3839 | 0.033297 | 133.075 | 0.55486 | 100101.1354 | 100051.9138 | 131.7433 | 0.00013167
1303 | A1 | 100015.947 | 0.065305 | 135.545 | 0.54173 | 100015.6987 | 99966.4772 | 134.2133 | 0.00013426
1329 | A2 | 200421.712 | 0.11548 | 482.78 | 0.90089 | 200421.2144 | 200371.9928 | 481.4483 | 0.00024028
1355 | A3 | 300392.136 | 0.41094 | 936.55 | 3.9683 | 300391.3901 | 300342.1686 | 935.2183 | 0.00031138
1381 | A4 | 400791.768 | 0.64947 | 1670.35 | 4.6505 | 400790.7729 | 400741.5513 | 1669.0183 | 0.00041648
1407 | A5 | 500612.38 | 0.66378 | 2631.55 | 5.4907 | 500611.137 | 500561.9155 | 2630.2183 | 0.00052545
1433 | A6 | 600591.812 | 0.74936 | 3674.25 | 8.2757 | 600590.3208 | 600541.0992 | 3672.9183 | 0.0006116
1459 | A7 | 701006.25 | 2.4309 | 4883.35 | 36.3803 | 701004.5094 | 700955.2879 | 4882.0183 | 0.00069648
1485 | A8 | 800991.225 | 4.1986 | 6379 | 37.8021 | 800989.2362 | 800940.0147 | 6377.6683 | 0.00079627
1511 | A9 | 901403.07 | 4.2023 | 8020.5 | 43.872 | 901400.8319 | 901351.6103 | 8019.1683 | 0.00088968
1537 | A10 | 1001391.815 | 4.8028 | 9972 | 40.5709 | 1001389.3286 | 1001340.1071 | 9970.6683 | 0.00099573

## report
| Universal Bridge | [7 empty] | AH2700
[17 empty] | En | En
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | C | tand
E1 | 1pF | 10 | 10.4647 | 1.2853 | 2.5704 | 2.3509 | 0.48048 | 2.0217 | 1pF | 10 | 10.4327 | 1.2853 | 2.5704 | 0.057143 | 0.48048 | 2.0217 | 0.017607 | 3.3756
E2 | DIV | 20 | 21.2794 | 1.2853 | 2.5704 | 1.3911 | 0.21955 | 2.0131 | DIV | 20 | 21.0017 | 1.2853 | 2.5704 | 0.17585 | 0.21955 | 2.0131 | 0.15275 | 3.914
E3 |  | 30 | 31.498 | 1.2853 | 2.5704 | 0.75247 | 0.14576 | 2.0137 |  | 30 | 31.6071 | 1.2853 | 2.5704 | 0.051665 | 0.14576 | 2.0137 | 0.060019 | 3.3998
E4 |  | 40 | 41.7956 | 1.2853 | 2.5703 | 0.54793 | 0.11131 | 2.0119 |  | 40 | 41.702 | 1.2853 | 2.5703 | 0.043898 | 0.11131 | 2.0119 | 0.051494 | 3.2018
E5 |  | 50 | 51.6942 | 1.2853 | 2.5703 | 0.46236 | 0.091936 | 2.01 |  | 50 | 51.2556 | 1.2853 | 2.5703 | 0.024091 | 0.091936 | 2.01 | 0.2413 | 3.3709
E6 |  | 60 | 61.5428 | 1.2853 | 2.5702 | 0.39487 | 0.079262 | 2.0091 |  | 60 | 61.4694 | 1.2853 | 2.5702 | 0.022396 | 0.079262 | 2.0091 | 0.040378 | 3.3229
E7 |  | 70 | 70.9773 | 1.2853 | 2.5702 | 0.31561 | 0.070714 | 2.0094 |  | 70 | 71.2863 | 1.2853 | 2.5702 | 0.020426 | 0.070714 | 2.0094 | 0.17 | 2.9517
E8 |  | 80 | 80.6608 | 1.2853 | 2.5701 | 0.27276 | 0.06431 | 2.0111 |  | 80 | 80.7521 | 1.2853 | 2.5701 | 0.025682 | 0.06431 | 2.0111 | 0.050254 | 2.7167
E9 |  | 90 | 90.6844 | 1.2852 | 2.57 | 0.2371 | 0.059353 | 2.0144 |  | 90 | 90.5167 | 1.2852 | 2.57 | 0.02252 | 0.059353 | 2.0144 | 0.092245 | 2.5564
E10 |  | '100' | 100.0589 | 1.2852 | 2.57 | 0.20889 | 0.055792 | 2.0186 |  | '100' | 100.3548 | 1.2852 | 2.57 | 0.023711 | 0.055792 | 2.0186 | 0.16282 | 2.3469
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
D1 | 100pF | 1 | 96.6323 | 0.020514 | 2.0807 | 0.05692 | 0.056946 | 2.0171 | 100pF | 1 | 96.6669 | 0.020514 | 2.0807 | 0.2957 | 0.056946 | 2.0171 | 1.1909 | 2.965
D2 | STEPS | 2 | 195.736 | 0.033875 | 2.0869 | 0.02197 | 0.042311 | 2.0883 | STEPS | 2 | 195.7128 | 0.033875 | 2.0869 | 0.052193 | 0.042311 | 2.0883 | 0.48566 | 0.50509
D3 |  | 3 | 292.2363 | 0.047872 | 2.0888 | 0.025666 | 0.039614 | 2.1426 |  | 3 | 292.2575 | 0.047872 | 2.0888 | 0.062861 | 0.039614 | 2.1426 | 0.31206 | 0.66394
D4 |  | 4 | 392.0512 | 0.060946 | 2.085 | 0.0071423 | 0.038703 | 2.1738 |  | 4 | 392.0399 | 0.060946 | 2.085 | 0.053258 | 0.038703 | 2.1738 | 0.13081 | 0.84254
D5 |  | 5 | 488.5535 | 0.075221 | 2.0858 | 0.0092114 | 0.03832 | 2.1909 |  | 5 | 488.5433 | 0.075221 | 2.0858 | 0.059191 | 0.03832 | 2.1909 | 0.095857 | 0.92225
D6 |  | 6 | 591.5278 | 0.089862 | 2.0851 | -0.0062553 | 0.038108 | 2.2017 |  | 6 | 591.481 | 0.089862 | 2.0851 | 0.056597 | 0.038108 | 2.2017 | 0.36822 | 1.1662
D7 |  | 7 | 690.6225 | 0.10397 | 2.0846 | 0.0099915 | 0.03799 | 2.2083 |  | 7 | 690.5845 | 0.10397 | 2.0846 | 0.051441 | 0.03799 | 2.2083 | 0.25872 | 0.77151
D8 |  | 8 | 787.1568 | 0.11558 | 2.0813 | 0.027061 | 0.037916 | 2.2126 |  | 8 | 787.1285 | 0.11558 | 2.0813 | 0.055379 | 0.037916 | 2.2126 | 0.1733 | 0.52812
D9 |  | 9 | 886.9637 | 0.12955 | 2.0809 | 0.03732 | 0.037864 | 2.2158 |  | 9 | 886.9174 | 0.12955 | 2.0809 | 0.051515 | 0.037864 | 2.2158 | 0.25238 | 0.26507
D10 |  | 10 | 983.2789 | 0.14304 | 2.0806 | 0.044649 | 0.037829 | 2.218 |  | 10 | 983.2437 | 0.14304 | 2.0806 | 0.055814 | 0.037829 | 2.218 | 0.17403 | 0.20869
| Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
C1 | 1000pF | 1 | 0.99511 | 0.00014883 | 2.0582 | 0.040196 | 0.055932 | 2.0185 | 1000pF | 1 | 0.99502 | 0.00014883 | 2.0582 | 0.055136 | 0.055932 | 2.0185 | 0.40903 | 0.18887
C2 | STEPS | 2 | 1.9975 | 0.00028792 | 2.073 | 0.017021 | 0.042112 | 2.0912 | STEPS | 2 | 1.9974 | 0.00028792 | 2.073 | 0.043563 | 0.042112 | 2.0912 | 0.23226 | 0.44565
C3 |  | 3 | 2.9923 | 0.00042654 | 2.0757 | 0.012365 | 0.039517 | 2.1455 |  | 3 | 2.9921 | 0.00042654 | 2.0757 | 0.048264 | 0.039517 | 2.1455 | 0.33287 | 0.64237
C4 |  | 4 | 3.9949 | 0.00056671 | 2.0766 | 0.0057573 | 0.038664 | 2.1754 |  | 4 | 3.9947 | 0.00056671 | 2.0766 | 0.047174 | 0.038664 | 2.1754 | 0.24447 | 0.75745
C5 |  | 5 | 4.9983 | 0.00070718 | 2.077 | -0.0014005 | 0.038292 | 2.1922 |  | 5 | 4.998 | 0.00070718 | 2.077 | 0.049076 | 0.038292 | 2.1922 | 0.28308 | 0.93211
C6 |  | 6 | 5.9926 | 0.00084646 | 2.0772 | -0.0093448 | 0.038099 | 2.2022 |  | 6 | 5.9922 | 0.00084646 | 2.0772 | 0.050597 | 0.038099 | 2.2022 | 0.32714 | 1.1125
C7 |  | 7 | 6.995 | 0.00099568 | 2.0786 | 0.0020014 | 0.037985 | 2.2085 |  | 7 | 6.9947 | 0.00099568 | 2.0786 | 0.049659 | 0.037985 | 2.2085 | 0.2232 | 0.88716
C8 |  | 8 | 7.9891 | 0.0011464 | 2.08 | 0.014394 | 0.037914 | 2.2128 |  | 8 | 7.9888 | 0.0011464 | 2.08 | 0.052245 | 0.037914 | 2.2128 | 0.23653 | 0.70593
C9 |  | 9 | 8.9914 | 0.0012668 | 2.0774 | 0.022911 | 0.037865 | 2.2157 |  | 9 | 8.991 | 0.0012668 | 2.0774 | 0.051359 | 0.037865 | 2.2157 | 0.22158 | 0.53126
C10 |  | 10 | 9.9856 | 0.0014186 | 2.0787 | 0.037053 | 0.037832 | 2.2178 |  | 10 | 9.985 | 0.0014186 | 2.0787 | 0.053956 | 0.037832 | 2.2178 | 0.29606 | 0.31592
| Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
B1 | 0.01µF | 1 | 10.0128 | 0.0015272 | 2.0632 | 0.036953 | 0.055732 | 2.0188 | 0.01µF | 1 | 10.0117 | 0.0015272 | 2.0632 | 0.05378 | 0.055732 | 2.0188 | 0.48892 | 0.2135
B2 | STEPS | 2 | 20.0045 | 0.002914 | 2.0754 | 0.028493 | 0.042136 | 2.0907 | STEPS | 2 | 20.003 | 0.002914 | 2.0754 | 0.063936 | 0.042136 | 2.0907 | 0.35534 | 0.59478
B3 |  | 3 | 30.017 | 0.0043474 | 2.0786 | 0.019989 | 0.039587 | 2.1432 |  | 3 | 30.0147 | 0.0043474 | 2.0786 | 0.067515 | 0.039587 | 2.1432 | 0.36959 | 0.84892
B4 |  | 4 | 40.0205 | 0.005784 | 2.0798 | 0.013493 | 0.0388 | 2.1698 |  | 4 | 40.0176 | 0.005784 | 2.0798 | 0.076769 | 0.0388 | 2.1698 | 0.35706 | 1.1532
B5 |  | 5 | 50.0297 | 0.0070958 | 2.0776 | 0.0093944 | 0.038503 | 2.1823 |  | 5 | 50.0236 | 0.0070958 | 2.0776 | 0.086591 | 0.038503 | 2.1823 | 0.60564 | 1.4177
B6 |  | 6 | 60.0424 | 0.0086638 | 2.0806 | -0.011825 | 0.038401 | 2.1873 |  | 6 | 60.0349 | 0.0086638 | 2.0806 | 0.095242 | 0.038401 | 2.1873 | 0.61315 | 1.9715
B7 |  | 7 | 70.0335 | 0.010012 | 2.0794 | 0.012708 | 0.038397 | 2.188 |  | 7 | 70.0265 | 0.010012 | 2.0794 | 0.10244 | 0.038397 | 2.188 | 0.49116 | 1.6525
B8 |  | 8 | 80.0456 | 0.011339 | 2.0782 | 0.022362 | 0.038456 | 2.1861 |  | 8 | 80.0379 | 0.011339 | 2.0782 | 0.11273 | 0.038456 | 2.1861 | 0.47726 | 1.6617
B9 |  | 9 | 90.0487 | 0.012533 | 2.0757 | 0.035425 | 0.038561 | 2.1826 |  | 9 | 90.0405 | 0.012533 | 2.0757 | 0.12167 | 0.038561 | 2.1826 | 0.46412 | 1.5815
B10 |  | 10 | 100.0607 | 0.014046 | 2.077 | 0.050369 | 0.038703 | 2.178 |  | 10 | 100.0519 | 0.014046 | 2.077 | 0.13167 | 0.038703 | 2.178 | 0.44051 | 1.4854
| Dial Label | Dial Setting | Capacitance Change (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
A1 | 0.1µF | 1 | 0.099974 | 1.4252e-05 | 2.0519 | 0.04001 | 0.056589 | 2.0154 | 0.1µF | 1 | 0.099966 | 1.4252e-05 | 2.0519 | 0.13426 | 0.056589 | 2.0154 | 0.35121 | 1.1777
A2 | STEPS | 2 | 0.20038 | 2.7687e-05 | 2.0673 | 0.044914 | 0.046296 | 2.0596 | STEPS | 2 | 0.20037 | 2.7687e-05 | 2.0673 | 0.24028 | 0.046296 | 2.0596 | 0.28504 | 2.9839
A3 |  | 3 | 0.30036 | 4.2323e-05 | 2.0745 | 0.039619 | 0.049721 | 2.0927 |  | 3 | 0.30034 | 4.2323e-05 | 2.0745 | 0.31138 | 0.049721 | 2.0927 | 0.22624 | 3.8649
A4 |  | 4 | 0.40076 | 5.5875e-05 | 2.0744 | 0.052151 | 0.057201 | 2.1357 |  | 4 | 0.40074 | 5.5875e-05 | 2.0744 | 0.41648 | 0.057201 | 2.1357 | 0.18673 | 4.5038
A5 |  | 5 | 0.50057 | 6.9133e-05 | 2.0737 | 0.06912 | 0.066636 | 2.1832 |  | 5 | 0.50056 | 6.9133e-05 | 2.0737 | 0.52545 | 0.066636 | 2.1832 | 0.13227 | 4.8424
A6 |  | 6 | 0.60057 | 8.3613e-05 | 2.0753 | 0.013987 | 0.077184 | 2.2271 |  | 6 | 0.60054 | 8.3613e-05 | 2.0753 | 0.6116 | 0.077184 | 2.2271 | 0.23093 | 5.4749
A7 |  | 7 | 0.70096 | 9.756e-05 | 2.0754 | 0.074469 | 0.088403 | 2.2643 |  | 7 | 0.70096 | 9.756e-05 | 2.0754 | 0.69648 | 0.088403 | 2.2643 | 0.019669 | 4.9753
A8 |  | 8 | 0.80083 | 0.00011144 | 2.0756 | 0.091904 | 0.099932 | 2.2944 |  | 8 | 0.80094 | 0.00011144 | 2.0756 | 0.79627 | 0.099932 | 2.2944 | 0.6946 | 4.984
A9 |  | 9 | 0.90133 | 0.0001265 | 2.0769 | 0.11239 | 0.11176 | 2.3187 |  | 9 | 0.90135 | 0.0001265 | 2.0769 | 0.88968 | 0.11176 | 2.3187 | 0.10875 | 4.918
A10 |  | 10 | 1.0013 | 0.0001393 | 2.0757 | 0.13253 | 0.12367 | 2.3381 |  | 10 | 1.0013 | 0.0001393 | 2.0757 | 0.99573 | 0.12367 | 2.3381 | 0.17968 | 4.9357
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
main_zero | All dials | 0 | 49.1438 | 1.5423 | 2.5705 | 2.1978 | 0.11021 | 2.0182 | All dials | 0 | 49.2215 | 1.5423 | 2.5705 | 2.7056 | 0.11021 | 2.0182 | 0.035632 | 3.2582

## Previous
dial code | Dial | Dial | Capacitance | Expanded | k | Loss angle | Expanded | k
| Label | Setting | change | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian) | [5 empty] | AH2700
E16Y | 1pF | 10 | 10.3256 | 1.2348 | 2.4695 | 0.092024 | 0.34495 | 2.082 | 10 | 10.4647 | 1.2853 | 2.5704 | 0.057143 | 0.48048 | 2.0217 | 0.078025 | 0.05897
E26Y | DIV | 20 | 21.0663 | 1.2348 | 2.4695 | 0.13261 | 0.17817 | 2.0702 | 20 | 21.2794 | 1.2853 | 2.5704 | 0.17585 | 0.21955 | 2.0131 | 0.11958 | 0.15296
E36Y |  | 30 | 31.7019 | 1.2348 | 2.4694 | 0.052217 | 0.12525 | 2.0637 | 30 | 31.498 | 1.2853 | 2.5704 | 0.051665 | 0.14576 | 2.0137 | 0.1144 | 0.0028687
E46Y |  | 40 | 41.8955 | 1.2348 | 2.4694 | 0.049209 | 0.10028 | 2.0612 | 40 | 41.7956 | 1.2853 | 2.5703 | 0.043898 | 0.11131 | 2.0119 | 0.056017 | 0.035447
E56Y |  | 50 | 51.636 | 1.2348 | 2.4693 | 0.038303 | 0.085959 | 2.0614 | 50 | 51.6942 | 1.2853 | 2.5703 | 0.024091 | 0.091936 | 2.01 | 0.032669 | 0.11292
E66Y |  | 60 | 61.1184 | 1.2348 | 2.4693 | 0.032946 | 0.076623 | 2.0633 | 60 | 61.5428 | 1.2853 | 2.5702 | 0.022396 | 0.079262 | 2.0091 | 0.23808 | 0.095699
E76Y |  | 70 | 70.6509 | 1.2348 | 2.4692 | 0.048452 | 0.069923 | 2.0663 | 70 | 70.9773 | 1.2853 | 2.5702 | 0.020426 | 0.070714 | 2.0094 | 0.18311 | 0.28182
E86Y |  | 80 | 80.2034 | 1.2348 | 2.4691 | 0.04799 | 0.064926 | 2.0699 | 80 | 80.6608 | 1.2853 | 2.5701 | 0.025682 | 0.06431 | 2.0111 | 0.25664 | 0.24412
E96Y |  | 90 | 90.3229 | 1.2348 | 2.4691 | 0.041556 | 0.060881 | 2.0742 | 90 | 90.6844 | 1.2852 | 2.57 | 0.02252 | 0.059353 | 2.0144 | 0.20279 | 0.2239
E106Y |  | '100' | 100.1625 | 1.2348 | 2.469 | 0.025056 | 0.057803 | 2.0785 | '100' | 100.0589 | 1.2852 | 2.57 | 0.023711 | 0.055792 | 2.0186 | 0.058124 | 0.016735
D16Y | 100pF | 1 | 96.96 | 0.02408 | 2.2338 | 0.32282 | 0.058729 | 2.0771 | 1 | 96.6323 | 0.020514 | 2.0807 | 0.2957 | 0.056946 | 2.0171 | 10.3573 | 0.33147
D26Y | STEPS | 2 | 195.9392 | 0.040183 | 2.2424 | 0.047648 | 0.045048 | 2.1123 | 2 | 195.736 | 0.033875 | 2.0869 | 0.052193 | 0.042311 | 2.0883 | 3.8665 | 0.073543
D36Y |  | 3 | 292.7762 | 0.055966 | 2.2449 | 0.13935 | 0.04117 | 2.1293 | 3 | 292.2363 | 0.047872 | 2.0888 | 0.062861 | 0.039614 | 2.1426 | 7.3302 | 1.3387
D46Y |  | 4 | 392.4936 | 0.072229 | 2.246 | 0.04977 | 0.039311 | 2.1381 | 4 | 392.0512 | 0.060946 | 2.085 | 0.053258 | 0.038703 | 2.1738 | 4.6813 | 0.063237
D56Y |  | 5 | 489.3095 | 0.088022 | 2.2465 | 0.10453 | 0.038276 | 2.1427 | 5 | 488.5535 | 0.075221 | 2.0858 | 0.059191 | 0.03832 | 2.1909 | 6.5299 | 0.83716
D66Y |  | 6 | 592.1304 | 0.1048 | 2.2468 | 0.094743 | 0.037568 | 2.1457 | 6 | 591.5278 | 0.089862 | 2.0851 | 0.056597 | 0.038108 | 2.2017 | 4.3651 | 0.71286
D76Y |  | 7 | 691.1667 | 0.12096 | 2.247 | 0.050106 | 0.037095 | 2.1475 | 7 | 690.6225 | 0.10397 | 2.0846 | 0.051441 | 0.03799 | 2.2083 | 3.4118 | 0.025148
D86Y |  | 8 | 787.9986 | 0.13676 | 2.2471 | 0.083688 | 0.036753 | 2.1487 | 8 | 787.1568 | 0.11558 | 2.0813 | 0.055379 | 0.037916 | 2.2126 | 4.7014 | 0.5361
D96Y |  | 9 | 887.734 | 0.15303 | 2.2472 | 0.050688 | 0.036482 | 2.1496 | 9 | 886.9637 | 0.12955 | 2.0809 | 0.051515 | 0.037864 | 2.2158 | 3.8422 | 0.015728
D106Y |  | X | 984.3529 | 0.1688 | 2.2473 | 0.077226 | 0.036273 | 2.1502 | 10 | 983.2789 | 0.14304 | 2.0806 | 0.055814 | 0.037829 | 2.218 | 4.8543 | 0.40855
| Dial | Dial | Capacitance | Expanded | k | Loss angle | Expanded | k
| Label | Setting | change | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian)
C15Y | 1000pF | 1 | 0.99596 | 0.00017247 | 2.222 | 0.024097 | 0.045354 | 2.0619 | 1 | 0.99511 | 0.00014883 | 2.0582 | 0.040196 | 0.055932 | 2.0185 | 3.7528 | 0.22358
C25Y | STEPS | 2 | 1.9991 | 0.00033526 | 2.2403 | 0.010004 | 0.03751 | 2.1004 | 2 | 1.9975 | 0.00028792 | 2.073 | 0.017021 | 0.042112 | 2.0912 | 3.7331 | 0.12443
C35Y |  | 3 | 2.9947 | 0.00049746 | 2.2442 | 0.0083478 | 0.035976 | 2.124 | 3 | 2.9923 | 0.00042654 | 2.0757 | 0.012365 | 0.039517 | 2.1455 | 3.6849 | 0.075172
C45Y |  | 4 | 3.9982 | 0.00066111 | 2.2456 | 0.00075031 | 0.035407 | 2.1354 | 4 | 3.9949 | 0.00056671 | 2.0766 | 0.0057573 | 0.038664 | 2.1754 | 3.8353 | 0.095505
C55Y |  | 5 | 5.0024 | 0.00082491 | 2.2463 | -0.0047976 | 0.035129 | 2.1413 | 5 | 4.9983 | 0.00070718 | 2.077 | -0.0014005 | 0.038292 | 2.1922 | 3.7223 | 0.065374
C65Y |  | 6 | 5.9974 | 0.00098726 | 2.2467 | -0.0061691 | 0.034971 | 2.1446 | 6 | 5.9926 | 0.00084646 | 2.0772 | -0.0093448 | 0.038099 | 2.2022 | 3.6982 | 0.061405
C75Y |  | 7 | 7.0007 | 0.001151 | 2.2469 | -0.00028568 | 0.034869 | 2.1468 | 7 | 6.995 | 0.00099568 | 2.0786 | 0.0020014 | 0.037985 | 2.2085 | 3.735 | 0.044355
C85Y |  | 8 | 7.9956 | 0.0013133 | 2.2471 | 0.009755 | 0.0348 | 2.1481 | 8 | 7.9891 | 0.0011464 | 2.08 | 0.014394 | 0.037914 | 2.2128 | 3.7281 | 0.090149
C95Y |  | 9 | 8.9988 | 0.001477 | 2.2472 | 0.022002 | 0.034751 | 2.149 | 9 | 8.9914 | 0.0012668 | 2.0774 | 0.022911 | 0.037865 | 2.2157 | 3.8101 | 0.017674
C105Y |  | X | 9.9938 | 0.0016394 | 2.2472 | 0.036322 | 0.034714 | 2.1497 | 10 | 9.9856 | 0.0014186 | 2.0787 | 0.037053 | 0.037832 | 2.2178 | 3.7818 | 0.014248
B14Y | 0.01µF | 1 | 10.0212 | 0.0016624 | 2.2202 | 0.040912 | 0.044109 | 2.0612 | 1 | 10.0128 | 0.0015272 | 2.0632 | 0.036953 | 0.055732 | 2.0188 | 3.7427 | 0.055708
B24Y | STEPS | 2 | 20.0207 | 0.0032845 | 2.24 | 0.029469 | 0.036828 | 2.098 | 2 | 20.0045 | 0.002914 | 2.0754 | 0.028493 | 0.042136 | 2.0907 | 3.689 | 0.017428
B34Y |  | 3 | 30.0416 | 0.004917 | 2.2441 | 0.020637 | 0.035565 | 2.1208 | 3 | 30.017 | 0.0043474 | 2.0786 | 0.019989 | 0.039587 | 2.1432 | 3.7552 | 0.012194
B44Y |  | 4 | 40.053 | 0.0065495 | 2.2456 | 0.015229 | 0.035196 | 2.1298 | 4 | 40.0205 | 0.005784 | 2.0798 | 0.013493 | 0.0388 | 2.1698 | 3.7189 | 0.033147
B54Y |  | 5 | 50.0693 | 0.0081834 | 2.2463 | 0.010785 | 0.03509 | 2.1326 | 5 | 50.0297 | 0.0070958 | 2.0776 | 0.0093944 | 0.038503 | 2.1823 | 3.6601 | 0.02669
B64Y |  | 6 | 60.0901 | 0.0098184 | 2.2467 | 0.0048259 | 0.035101 | 2.1324 | 6 | 60.0424 | 0.0086638 | 2.0806 | -0.011825 | 0.038401 | 2.1873 | 3.6422 | 0.32005
B74Y |  | 7 | 70.0893 | 0.01145 | 2.2469 | 0.01498 | 0.035179 | 2.1304 | 7 | 70.0335 | 0.010012 | 2.0794 | 0.012708 | 0.038397 | 2.188 | 3.6714 | 0.043635
B84Y |  | 8 | 80.1099 | 0.013085 | 2.2471 | 0.024341 | 0.035307 | 2.1273 | 8 | 80.0456 | 0.011339 | 2.0782 | 0.022362 | 0.038456 | 2.1861 | 3.7188 | 0.037903
B94Y |  | 9 | 90.1209 | 0.014719 | 2.2472 | 0.037948 | 0.035475 | 2.1237 | 9 | 90.0487 | 0.012533 | 2.0757 | 0.035425 | 0.038561 | 2.1826 | 3.7368 | 0.048151
B104Y |  | X | 100.1418 | 0.016354 | 2.2472 | 0.053323 | 0.035679 | 2.1198 | 10 | 100.0607 | 0.014046 | 2.077 | 0.050369 | 0.038703 | 2.178 | 3.7614 | 0.056114
A13Y | 0.1µF | 1 | 0.10003 | 1.6523e-05 | 2.2199 | 0.042985 | 0.045009 | 2.052 | 1 | 0.099974 | 1.4252e-05 | 2.0519 | 0.04001 | 0.056589 | 2.0154 | 2.7262 | 0.041135
A23Y | STEPS | 2 | 0.20049 | 3.2817e-05 | 2.24 | 0.051374 | 0.04162 | 2.0647 | 2 | 0.20038 | 2.7687e-05 | 2.0673 | 0.044914 | 0.046296 | 2.0596 | 2.3982 | 0.10377
A33Y |  | 3 | 0.30052 | 4.9114e-05 | 2.2441 | 0.05091 | 0.04711 | 2.0985 | 3 | 0.30036 | 4.2323e-05 | 2.0745 | 0.039619 | 0.049721 | 2.0927 | 2.5212 | 0.16484
A43Y |  | 4 | 0.40097 | 6.5493e-05 | 2.2456 | 0.060602 | 0.055685 | 2.1506 | 4 | 0.40076 | 5.5875e-05 | 2.0744 | 0.052151 | 0.057201 | 2.1357 | 2.4386 | 0.10586
A53Y |  | 5 | 0.50087 | 8.179e-05 | 2.2463 | 0.079659 | 0.065807 | 2.2044 | 5 | 0.50057 | 6.9133e-05 | 2.0737 | 0.06912 | 0.066636 | 2.1832 | 2.7679 | 0.11253
A63Y |  | 6 | 0.60091 | 9.8112e-05 | 2.2467 | 0.084869 | 0.076824 | 2.2514 | 6 | 0.60057 | 8.3613e-05 | 2.0753 | 0.013987 | 0.077184 | 2.2271 | 2.6214 | 0.6509
A73Y |  | 7 | 0.70136 | 0.0001145 | 2.2469 | 0.10123 | 0.088393 | 2.29 | 7 | 0.70096 | 9.756e-05 | 2.0754 | 0.074469 | 0.088403 | 2.2643 | 2.6617 | 0.21406
A83Y |  | 8 | 0.80139 | 0.00013083 | 2.2471 | 0.12079 | 0.10021 | 2.3206 | 8 | 0.80083 | 0.00011144 | 2.0756 | 0.091904 | 0.099932 | 2.2944 | 3.258 | 0.20409
A93Y |  | 9 | 0.90184 | 0.00014722 | 2.2472 | 0.1437 | 0.11226 | 2.3449 | 9 | 0.90133 | 0.0001265 | 2.0769 | 0.11239 | 0.11176 | 2.3187 | 2.6139 | 0.19768
A103Y |  | X | 1.0019 | 0.00016354 | 2.2472 | 0.16519 | 0.12437 | 2.3641 | 10 | 1.0013 | 0.0001393 | 2.0757 | 0.13253 | 0.12367 | 2.3381 | 2.6526 | 0.18621
dial code | Dial | Dial | Capacitance | Expanded | k | Loss angle | Expanded | k
|  | Setting |  | uncertainty |  |  | uncertainty
|  |  | (pF) | (pF) |  | (milliradian) | (milliradian)
E06Y | All dials | 0 | 48.9946 | 1.4818 | 2.4696 | 2.9562 | 0.069554 | 2.0988 | 0 | 49.1438 | 1.5423 | 2.5705 | 2.7056 | 0.11021 | 2.0182 | 0.069773 | 1.9232