## Commentary
As S21815.xls
Start with S21580a
Update history page for easier plotting.
Copy 'Universal Bridge' from most recent S21814 for RLC standards.
Carefully redefine critical block names.
Alter label structure on Results page, could try to use vlookup to find range zeros on UniversalBridge sheet?
Replace uncertainty sheet with the uncertainty sheet from S21814
Alter treatment of range 7Y and 6Y by applying corrections as at bottom of Results sheet
Modify Report to pick up corrected results
As S21942.xlsx
Start with S21815.xls
Enter Ubridge calibration data of 7/7/14
Add graph history of UB cal constants, just for interest.
Delays in calibrating the UB (and repairing the permutable) are the reason for having the standards so long, and hence measuring them twice.
Range 7 results remain a problem for loss angle measurements, see 'Range7' worksheet.
Therefore for reporting loss angle for the 1, 10 and 100 pF capacitors, use AH2700 results with UB uncertainties (for CMC compliance).
As there is doubt about the 1 June loss angle measurements, only the 12 August results are reported

## pyUBreadings
Template for capacitors
Capacitors |  |  |  | Client | [5 empty] | Enter, say, 5 ppm as 5e-6
Date |  |  |  | Job number | [5 empty] | 1 sigma uncertainties from specifications
|  |  |  | S Number | [5 empty] | specs say -140 ppm for capacitance, assume 100 ppm for G
|  |  | Data from lab book |  |  |  | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff C | u | df | Temp coeff G | u | df | Other | u | df
| Hz |  |  |  | Hz | °C
coax zero | 1591.5 | 3Y | 0 | -4 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1 microF | 1591.5 | 3Y | 999801 | 1696 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10
coax zero | 1591.5 | 4Y | 0 | -3 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
0.1 microF | 1591.5 | 4Y | 999928 | 394 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10
coax zero | 1591.5 | 5Y | -1 | 3 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
0.01 microF | 1591.5 | 5Y | 999986 | 496 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10
coax zero | 1591.5 | 6Y | 0 | 23 | 1591.5 | 19.65 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1000 pF | 1591.5 | 6Y | 999660 | -205 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10
GR1000B | 1591.5 | 6Y | 999863 | -136 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10
GR1000A | 1591.5 | 6Y | 999878 | -170 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10
coax zero | 1591.5 | 7Y | 6 | 14 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
100 pF | 1591.5 | 7Y | 999865 | 2540 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10
10 pF | 1591.5 | 7Y | 99969 | 260 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10
1 pF | 1591.5 | 7Y | 10004 | 40 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 1.2e-05 | 10 | 0 | 0 | 10
AH11A | 1591.5 | 7Y | 99989 | 270 | 1591.5 | 19.58 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10
AH11C | 1591.5 | 7Y | 999840 | 2540 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | td | utd | ktd | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for G | G | uG | k | multiplier for C | C | uC | k | nom Freq | td | utd | ktd
coax zero | 1591.5 | 3Y | 0 | -4 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2567e-16 | 0 | 0 | #DIV/0! | coax zero1591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5 | 0 | 0 | 0
1 microF | 1591.5 | 3Y | 999801 | 1696 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10 | 9.999e-07 | 4.9654e-11 | 2.0261 | 1.6999e-06 | 3.7819e-07 | 2.2184 | 0.17002 | 0.037824 | 2.2184 | 2.1998e-11 | 2.6997e-07 | 2.2572 | 1.4009 | 1 microF1591.5 | 1000 | 0.0016999 | 0.00037819 | 2.2184 | 1000000 | 0.9999 | 4.9654e-05 | 2.0261 | 1591.5 | 0.17002 | 0.037824 | 2.2184
coax zero | 1591.5 | 4Y | 0 | -3 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2567e-16 | 0 | 0 | #DIV/0! | coax zero1591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5 | 0 | 0 | 0
0.1 microF | 1591.5 | 4Y | 999928 | 394 | 1591.5 | 19.66 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10 | 1e-07 | 4.8964e-12 | 2.025 | 3.9698e-08 | 3.7824e-08 | 2.2184 | 0.039698 | 0.037824 | 2.2184 | 2.2002e-12 | 2.7003e-08 | 2.2254 | 1.4007 | 0.1 microF1591.5 | 1000 | 3.9698e-05 | 3.7824e-05 | 2.2184 | 1000000 | 0.1 | 4.8964e-06 | 2.025 | 1591.5 | 0.039698 | 0.037824 | 2.2184
coax zero | 1591.5 | 5Y | -1 | 3 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2545e-16 | 2.6972e-12 | 0 | 0 | coax zero1591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5 | 0 | 0 | 0
0.01 microF | 1591.5 | 5Y | 999986 | 496 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 2.75e-05 | 10 | 0 | 0 | 10 | 1.0001e-08 | 4.9425e-13 | 2.0257 | 4.9298e-09 | 3.7826e-09 | 2.2184 | 0.049295 | 0.037824 | 2.2184 | 2.2015e-13 | 2.7029e-09 | 2.2451 | 1.3995 | 0.01 microF1591.5 | 1000 | 4.9298e-06 | 3.7826e-06 | 2.2184 | 1000000 | 0.010001 | 4.9425e-07 | 2.0257 | 1591.5 | 0.049295 | 0.037824 | 2.2184
coax zero | 1591.5 | 6Y | 0 | 23 | 1591.5 | 19.65 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2567e-16 | 0 | 0 | #DIV/0! | coax zero1591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5 | 0 | 0 | 0
1000 pF | 1591.5 | 6Y | 999660 | -205 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10 | 9.9983e-10 | 3.6019e-14 | 2.0709 | -2.2801e-10 | 3.7816e-10 | 2.2184 | -0.022805 | 0.037824 | 2.2184 | 2.2122e-14 | 2.7264e-10 | 1.6282 | 1.387 | 1000 pF1591.5 | 1000 | -2.2801e-07 | 3.7816e-07 | 2.2184 | 1000000000000 | 999.8258 | 0.036019 | 2.0709 | 1591.5 | -0.022805 | 0.037824 | 2.2184
GR1000B | 1591.5 | 6Y | 999863 | -136 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10 | 1e-09 | 3.4258e-14 | 2.1042 | -1.59e-10 | 3.7823e-10 | 2.2184 | -0.0159 | 0.037824 | 2.2184 | 2.2126e-14 | 2.727e-10 | 1.5483 | 1.387 | GR1000B1591.5 | 1000 | -1.59e-07 | 3.7823e-07 | 2.2184 | 1000000000000 | 1000.0289 | 0.034258 | 2.1042 | 1591.5 | -0.0159 | 0.037824 | 2.2184
GR1000A | 1591.5 | 6Y | 999878 | -170 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10 | 1e-09 | 3.4258e-14 | 2.1042 | -1.93e-10 | 3.7824e-10 | 2.2184 | -0.0193 | 0.037824 | 2.2184 | 2.2127e-14 | 2.727e-10 | 1.5483 | 1.387 | GR1000A1591.5 | 1000 | -1.93e-07 | 3.7824e-07 | 2.2184 | 1000000000000 | 1000.0439 | 0.034258 | 2.1042 | 1591.5 | -0.0193 | 0.037824 | 2.2184
coax zero | 1591.5 | 7Y | 6 | 14 | 1591.5 | 19.63 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2568e-16 | 2.7001e-12 | 0 | 0 | coax zero1591.5 | 1000 | 0 | 0 | 0 | 1000000000000 | 0 | 0 | 0 | 1591.5 | 0 | 0 | 0
100 pF | 1591.5 | 7Y | 999865 | 2540 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10 | 9.9995e-11 | 3.6024e-15 | 2.0709 | 2.5259e-10 | 3.7821e-11 | 2.2184 | 0.25261 | 0.037824 | 2.2184 | 2.3256e-15 | 2.9698e-11 | 1.549 | 1.2735 | 100 pF1591.5 | 1000 | 2.5259e-07 | 3.7821e-08 | 2.2184 | 1000000000000 | 99.995 | 0.0036024 | 2.0709 | 1591.5 | 0.25261 | 0.037824 | 2.2184
10 pF | 1591.5 | 7Y | 99969 | 260 | 1591.5 | 19.64 | 0.2 | 10 | 0 | 9e-06 | 10 | 0 | 0 | 10 | 9.9972e-12 | 5.5672e-16 | 1.9998 | 2.4599e-11 | 5.5759e-12 | 2.0187 | 0.24606 | 0.055776 | 2.0187 | 3.4562e-16 | 5.3992e-12 | 1.6108 | 1.0327 | 10 pF1591.5 | 1000 | 2.4599e-08 | 5.5759e-09 | 2.0187 | 1000000000000 | 9.9972 | 0.00055672 | 1.9998 | 1591.5 | 0.24606 | 0.055776 | 2.0187
1 pF | 1591.5 | 7Y | 10004 | 40 | 1591.5 | 19.61 | 0.2 | 10 | 0 | 1.2e-05 | 10 | 0 | 0 | 10 | 9.9989e-13 | 4.4323e-16 | 2.0218 | 2.5999e-12 | 4.4299e-12 | 2.022 | 0.26002 | 0.44305 | 2.022 | 1.4768e-16 | 2.97e-12 | 3.0013 | 1.4915 | 1 pF1591.5 | 1000 | 2.5999e-09 | 4.4299e-09 | 2.022 | 1000000000000 | 0.99989 | 0.00044323 | 2.0218 | 1591.5 | 0.26002 | 0.44305 | 2.022
AH11A | 1591.5 | 7Y | 99989 | 270 | 1591.5 | 19.58 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10 | 9.9992e-12 | 5.4418e-16 | 2.0038 | 2.5599e-11 | 5.5763e-12 | 2.0187 | 0.25602 | 0.055769 | 2.0187 | 3.4566e-16 | 5.3998e-12 | 1.5743 | 1.0327 | AH11A1591.5 | 1000 | 2.5599e-08 | 5.5763e-09 | 2.0187 | 1000000000000 | 9.9992 | 0.00054418 | 2.0038 | 1591.5 | 0.25602 | 0.055769 | 2.0187
AH11C | 1591.5 | 7Y | 999840 | 2540 | 1591.5 | 19.6 | 0.2 | 10 | 0 | 1e-06 | 10 | 0 | 0 | 10 | 9.9993e-11 | 3.4256e-15 | 2.1042 | 2.5259e-10 | 3.782e-11 | 2.2184 | 0.25261 | 0.037824 | 2.2184 | 2.3255e-15 | 2.9697e-11 | 1.473 | 1.2735 | AH11C1591.5 | 1000 | 2.5259e-07 | 3.782e-08 | 2.2184 | 1000000000000 | 99.9925 | 0.0034256 | 2.1042 | 1591.5 | 0.25261 | 0.037824 | 2.2184

## pyReport
[9 empty] | Label | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
| Table 1 | [10 empty] | uncertainty | factor |  | uncertainty | factor
|  | Nominal Value (pF) | Capacitance (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor |  | (pF) | (pF) | (pF) |  | (milliradian) | (milliradian)
1000 pF |  | 1000 | 999.8258 | 0.036019 | 2.0709 | -0.022805 | 0.037824 | 2.2184 | rpt 1000 pF6Y | 1000 | 999.8274 | 0.050849 | 2.1617 | 0.00028984 | 0.034531 | 2.15
100 pF |  | 100 | 99.995 | 0.0036024 | 2.0709 | 0.25261 | 0.037824 | 2.2184 | rpt 100 pF7Y | 100 | 100.0022 | 0.0050861 | 2.1617 | 0.00040501 | 0.034534 | 2.15
10 pF |  | 10 | 9.9972 | 0.00055672 | 1.9998 | 0.24606 | 0.055776 | 2.0187 | rpt 10 pF7Y | 10 | 9.9981 | 0.00059909 | 2.1066 | -0.0015382 | 0.044023 | 2.0613
1 pF |  | 1 | 0.99989 | 0.00044323 | 2.0218 | 0.26002 | 0.44305 | 2.022 | rpt 1 pF7Y | 1 | 0.99989 | 0.00041579 | 2.5532 | -0.009751 | 0.31071 | 2.1827
[10 empty] | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
| Table 2 | [10 empty] | uncertainty | factor |  | uncertainty | factor
|  | Nominal Value (µF) | Capacitance (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor |  | (mF) | (mF) | (mF) |  | (milliradian) | (milliradian)
1 microF |  | 1 | 0.9999 | 4.9654e-05 | 2.0261 | 0.17002 | 0.037824 | 2.2184 | rpt 1 microF3Y | 1 | 0.9999 | 4.2891e-05 | 2.2948 | 0.17091 | 0.034534 | 2.15
0.1 microF |  | 0.1 | 0.1 | 4.8964e-06 | 2.025 | 0.039698 | 0.037824 | 2.2184 | rpt 0.1 microF4Y | 0.1 | 0.1 | 4.2896e-06 | 2.2948 | 0.041297 | 0.034534 | 2.15
0.01 microF |  | 0.01 | 0.010001 | 4.9425e-07 | 2.0257 | 0.049295 | 0.037824 | 2.2184 | rpt 0.01 microF5Y | 0.01 | 0.010001 | 4.2899e-07 | 2.2948 | 0.050094 | 0.034533 | 2.15

## Results
Template for decade and individual capacitors measured on Universal Bridge | [8 empty] | Checked by:
Capacitor(s) | HP 1638XX series |  |  | Client | Babcock Fitzroy
Date | 2014-06-01 |  |  | Job number | 96040215
repeat | 2014-08-12 |  |  | S Number | S21942 | [7 empty] | AH2700
[10 empty] | loss angle | Corrected C | Corrected loss | C | G | Loss angle
Item | Label | Range | a | b | Capacitance |  | Conductance |  | Frequency | milliradians
Hz
coax open | coax open3Y | 3Y | 0 | -4 | 0 | microF | 0 | milliS | 1591.5 | #DIV/0!
1 microF | 1 microF3Y | 3Y | 999801 | 1696 | 0.99987 | microF | 0.0016999 | milliS | 1591.5 | 0.17002
coax open | coax open4Y | 4Y | 0 | -3 | 0 | nF | 0 | microS | 1591.5 | #DIV/0!
0.1 microF | 0.1 microF4Y | 4Y | 999928 | 394 | 99.9999 | nF | 0.039698 | microS | 1591.5 | 0.039699
coax open | coax open5Y | 5Y | -1 | 3 | 0 | nF | 0 | microS | 1591.5 | #DIV/0!
0.01 microF | 0.01 microF5Y | 5Y | 999986 | 496 | 10.0006 | nF | 0.0049297 | microS | 1591.5 | 0.049296
coax open | coax open6Y | 6Y | 0 | 23 | 0 | pF | 0 | nS | 1591.5 | #DIV/0!
1000 pF | 1000 pF6Y | 6Y | 999660 | -205 | 999.7944 | pF | -0.228 | nS | 1591.5 | -0.022805 |  | -0.0018113
GR1000B | GR1000B6Y | 6Y | 999863 | -136 | 999.9975 | pF | -0.159 | nS | 1591.5 | -0.015901 |  | 0.0050935
GR1000A | GR1000A6Y | 6Y | 999878 | -170 | 1000.0125 | pF | -0.193 | nS | 1591.5 | -0.0193 |  | 0.0016937
coax open | coax open7Y | 7Y | 6 | 14 | 0 | pF | 0 | nS | 1591.5 | #DIV/0!
100 pF | 100 pF7Y | 7Y | 999865 | 2540 | 99.9916 | pF | 0.25258 | nS | 1591.5 | 0.25261 | 99.9949 | 0.10866
10 pF | 10 pF7Y | 7Y | 99969 | 260 | 9.9969 | pF | 0.024598 | nS | 1591.5 | 0.24607 | 9.9972 | 0.10212
1 pF | 1 pF7Y | 7Y | 10004 | 40 | 0.99986 | pF | 0.0025998 | nS | 1591.5 | 0.26003 | 0.99989 | 0.11608
AH11A | AH11A7Y | 7Y | 99989 | 270 | 9.9989 | pF | 0.025598 | nS | 1591.5 | 0.25602 | 9.9992 | 0.11207
AH11C | AH11C7Y | 7Y | 999840 | 2540 | 99.9891 | pF | 0.25258 | nS | 1591.5 | 0.25262 | 99.9924 | 0.10867
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A | #N/A | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A | #N/A | #N/A
100 pF | 100 pF6Y | 6Y | 99987 | -10 | 100.0004 | pF | -0.033 | nS | 1591.5 | -0.033001 | 100.0004
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
rpt 1 pF | rpt 1 pF7Y | 7Y | 10004 | 35 | 0.99986 | pF | 0.0020998 | nS | 1591.5 | 0.21002 | 0.99989 | 0.066072 | 0.99993 | -9.75e-05 | -0.009751
rpt 10 pF | rpt 10 pF7Y | 7Y | 99978 | 202 | 9.9978 | pF | 0.018799 | nS | 1591.5 | 0.18803 | 9.9981 | 0.044085 | 9.9987 | -0.0001538 | -0.0015382
rpt 100 pF | rpt 100 pF7Y | 7Y | 999938 | 1872 | 99.9989 | pF | 0.18579 | nS | 1591.5 | 0.18579 | 100.0022 | 0.041845 | 100.0004 | 0.000405 | 0.00040501
rpt open | rpt open7Y | 7Y | 6 | 14 | 0 | pF | 0 | nS | 1591.5 | #DIV/0!
rpt open | rpt open6Y | 6Y | 0 | 13 | 0 | pF | -0.01 | nS | 1591.5 | #DIV/0!
rpt 100 pF | rpt 100 pF6Y | 6Y | 99987 | -10 | 100.0004 | pF | -0.033 | nS | 1591.5 | -0.033001 | 100.0004 | -0.012007
rpt 1000 pF | rpt 1000 pF6Y | 6Y | 999693 | -184 | 999.8274 | pF | -0.207 | nS | 1591.5 | -0.020704 | 999.8274 | 0.00028984 | 999.8347 | 0.0111 | 0.0011102
rpt open | rpt open5Y | 5Y | -1 | 4 | 0 | nF | 9.9994e-06 | microS | 1591.5 | #DIV/0!
rpt 0.01 microF | rpt 0.01 microF5Y | 5Y | 1000018 | 504 | 10.0009 | nF | 0.0050097 | microS | 1591.5 | 0.050094
rpt 0.1 microF | rpt 0.1 microF4Y | 4Y | 999962 | 410 | 100.0033 | nF | 0.041298 | microS | 1591.5 | 0.041297
rpt open | rpt open4Y | 4Y | 0 | -2 | 0 | nF | 9.9994e-05 | microS | 1591.5 | #DIV/0!
rpt open | rpt open3Y | 3Y | 0 | -1 | 0 | microF | 2.9998e-06 | milliS | 1591.5 | #DIV/0!
rpt 1 microF | rpt 1 microF3Y | 3Y | 999834 | 1705 | 0.9999 | microF | 0.0017089 | milliS | 1591.5 | 0.17091
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A |  | ppm
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A | 7Y cap corrn | 32.1212 | from cell R21 of the "L&C" sheet of I:\MSL\Private\Electricity\Quality\TECHPROC\E023-Impedance\UBcalibrationJuly2014.xlsx
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A |  | millirad
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A | 7Y loss corrn | -0.14395 | from cell Q21 of the "L&C" sheet of I:\MSL\Private\Electricity\Quality\TECHPROC\E023-Impedance\UBcalibrationJuly2014.xlsx
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A | 6Y loss corrn | 0.020994 | from cell Q20 of the "L&C" sheet of I:\MSL\Private\Electricity\Quality\TECHPROC\E023-Impedance\UBcalibrationJuly2014.xlsx
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A | 1591.5 | #N/A
Value of MSL Capacitors |  |  | C (ppm) | milliradians | UB Cap correction (ppm) |  |  |  |  | UB correction (millirad)
GR1000A |  |  | 29.6 | 0 | 17.1277 |  |  |  |  | 0.0193
GR1000B |  |  | 15.9 | 0 | 18.4298 |  |  |  |  | 0.015901
AH11A |  |  | -4.9 | 0 | 107.6691 |  |  |  |  | -0.25602
AH11C |  |  | -4.4 | 0 | 104.1689 |  |  |  |  | -0.25262
Note on K55: Andrew Corney: apply phase angle correction
Note on K56: Andrew Corney: apply phase angle correction
Note on K57: Andrew Corney: apply phase angle correction
Note on D75: k.jones: from Post_BIPM_ES.xls, 1 t functions.
Note on E77: k.jones: 3 microradian is max spec for AH11

## Universal Bridge
New UB calibrations need to be entered in Data block and accessed through the Date |  |  | This table is the usual form for calculating spreadsheet corrections | [6 empty] | Date | C1 | R4A | R4B | R4C | G1 | G2 | c | d
|  |  | Component | Value | Date |  |  |  | 1.6 kHz | 1985-02-14 | 32 | 30.1 | -12 | 53 | 13 | 8.7 | <10 | <10 | except for range 4, up to 44? |  |  |  | -91.32 | 0 | 0 | 0 | 0 | 0
|  |  |  | (ppm) | 2014-07-07 |  |  |  | DC for resistors | 1988-11-30 | 123.32 | 35 | -12 | #N/A | 9 | 4 | 0 | 0 |  |  |  |  | 0 | 4.9 | 0 | #N/A | -4 | -4.7 | 0
Date | Record | Comments | C_1 | 143.1 | UBIII |  |  |  | DC for resistors | 1991-03-12 | 133 | 37 | -14 | 61 | 10 | 6 | 0 | 0 |  |  |  |  | 9.68 | 6.9 | -2 | 8 | -3 | -2.7 | 0
1981-05-22 | PEL Meas V | Check Y&Z ranges 6 & 4 @ 400 Hz | R_4A | 54.1261 | UBIII |  |  |  | DC for resistors | 1995-08-22 | 127.406 | 40.34 | -17.21 | 52 | 11.21 | 6.21 | 0 | 0 |  |  |  |  | 4.086 | 10.24 | -5.21 | -1 | -1.79 | -2.49 | 0
1983-12-15 | Universal Bridge II | Full calibration of standards and of all dial ratios and performance checks. | R_4B | -9.8547 | UBIII |  |  |  | DC for resistors | 1998-11-03 | 135.27 | 43.3741 | -16.3523 | 67.0039 | 11.1844 | 6.6891 | 0 | 0 |  |  |  |  | 11.95 | 13.2741 | -4.3523 | 14.0039 | -1.8156 | -2.0109 | 0
1984-02-10 |  | completion of above | R_4C | 67.1933 | UBIII |  |  |  | DC for resistors | 2001-09-08 | 124.1565 | 46.006 | -17.3198 | 64.091 | 12.0228 | 7.4763 | 0 | 0 |  |  |  |  | 0.83653 | 15.906 | -5.3198 | 11.091 | -0.97725 | -1.2237 | 0
1988-11-30 | UB II | component values measured | G_1 | 12.9485 | UBIII |  |  |  | DC for resistors | 2004-01-22 | 131.7148 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 0 | 0 |  |  |  |  | 8.3948 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 0
1991-03-12 | UB II | component values measured | G_2 | 7.4961 | UBIII |  |  |  | DC for resistors | 2004-02-18 | 123.9 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 5.2 | 0 |  |  |  |  | 0.58 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 5.2
1991-05-16 | UB II | problem with '6' dial, repaired! | [6 empty] | DC for resistors | 2005-02-18 | 129.4493 | 46.54 | -16.96 | 61.5 | 13.84 | 9 | 5.7 | 0 |  |  |  |  | 6.1293 | 16.44 | -4.96 | 8.5 | 0.84 | 0.3 | 5.7
1991-06-18 | AI III | problem with 0.01 mF on range 5Y, lack of DC path for current pump, AVOID | Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | assume negligible c,d |  |  | DC for resistors | 2009-08-13 | 130.2987 | 50.8747 | -14.1043 | 64.5247 | 12.5073 | 7.4753 | 2 | 0 |  |  |  |  | 6.9787 | 20.7747 | -2.1043 | 11.5247 | -0.4927 | -1.2247 | 2
1995-08-22 | UB II | bridge calibration | where Z3=R+jwC | correction_c | 10.97 |  |  |  | DC for resistors | 2014-07-07 | 143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 19.78 | 24.0261 | 2.1453 | 14.1933 | -0.051497 | -1.2039 | 10.97
1998-11-03 | UB II | bridge calibration, using HP3458 as transfer instrument | Y4=(G3/G2)(aG1+ßjwC1)(1+c+jd) |  |  | (G3=1/R4*)
2001-09-08 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem fixed | where Y4=G+jwC
2004-01-22 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem REALLY fixed
2004-02-18 | UBII | remeasure capacitor, same dc resistance values as for 22/01/04 | R4A is used for ranges 5 to 1 inclusive
2005-02-18 | UBII | bridge calibration, following MSLT.E.023.004. Uncertainties have increased. | R4B is used on range 6 only
2009-08-13 | UBIII | bridge calibration, following MSLT.E.023.004/5. Uncertainties have increased slightly from 2005. | R4C is used on range 7 only | [16 empty] | AH2700A | ppm | mrad
2014-07-07 | UBIII | bridge calibration, following MSLT.E.023.006. Uncertainties unchanged from 2009. | [9 empty] | Zeros | Range | Frequency | R&F | a | b |  |  | R_4C | 15.0002 | 0.096
[13 empty] | 3Y | 1591.5 | 3Y1591.5 | 0 | -4
[13 empty] | 4Y | 1591.5 | 4Y1591.5 | 0 | -3
[13 empty] | 5Y | 1591.5 | 5Y1591.5 | -1 | 3
[13 empty] | 6Y | 1591.5 | 6Y1591.5 | 0 | 23
|  |  |  | Range | a correction | b correction | [6 empty] | 7Y | 1591.5 | 7Y1591.5 | 6 | 14
[5 empty] | (ppm) | (ppm)
|  |  |  | 1Z | 200.7 | 70.5485
|  |  |  | 2Z | 200.7 | 70.5485
|  |  |  | 3Z | 200.7 | 70.5485
|  |  |  | 4Z | 200.7 | 70.5485
|  |  |  | 5Z | 200.7 | 70.5485
|  |  |  | 6Z | 136.7192 | 6.5677
|  |  |  | 7Z | 213.7672 | 83.6157
|  |  |  | 1Y | 70.5079 | -59.6436
|  |  |  | 2Y | 70.5079 | -59.6436
|  |  |  | 3Y | 70.5079 | -59.6436
|  |  |  | 4Y | 70.5079 | -59.6436
|  |  |  | 5Y | 70.5079 | -59.6436
|  |  |  | 6Y | 134.4887 | 4.3372
|  |  |  | 7Y | 57.4407 | -72.7108
[19 empty] | Symbol font versions |  | Plain version
[5 empty] | Multiplier (C&L) | Unit Name | Unit Multiplier | Multiplier (G&R) | Unit Name | Unit Multiplier | CMC n for L | CMC m for R | [6 empty] | Unit Name | Unit Name | Unit Name | Unit Name
|  |  |  | 1Z | 0.0001 | microH | 1e-06 | 0.0001 | milliohm | 0.001 | 0.2 | 0.002 | [6 empty] | mH | mW | microH | milliohm
|  |  |  | 2Z | 0.001 | microH | 1e-06 | 1e-06 | ohm | 1 | 2 | 0.02 | [6 empty] | mH | W | microH | ohm
|  |  |  | 3Z | 1e-05 | milliH | 0.001 | 1e-05 | ohm | 1 | 20 | 0.2 | [6 empty] | mH | W | milliH | ohm
|  |  |  | 4Z | 0.0001 | milliH | 0.001 | 0.0001 | ohm | 1 | 200 | 2 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 5Z | 0.001 | milliH | 0.001 | 0.001 | ohm | 1 | 2000 | 20 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 6Z | 1e-05 | H | 1 | 1e-05 | kilohm | 1000 | 20000 | 200 | [6 empty] | H | kW | H | kilohm
|  |  |  | 7Z | 0.0001 | H | 1 | 0.0001 | kilohm | 1000 | 200000 | 2000 | [6 empty] | H | kW | H | kilohm
|  |  |  | 1Y | 0.0001 | microF | 1e-06 | 0.0001 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 2Y | 1e-05 | microF | 1e-06 | 1e-05 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 3Y | 1e-06 | microF | 1e-06 | 1e-06 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 4Y | 0.0001 | nF | 1e-09 | 0.0001 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 5Y | 1e-05 | nF | 1e-09 | 1e-05 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 6Y | 0.001 | pF | 1e-12 | 0.001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
|  |  |  | 7Y | 0.0001 | pF | 1e-12 | 0.0001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
Note on M9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on N9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on O9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on P9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on Q9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on R9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on V18: k.jones: AH2700A check of 1 M at 1.5915 k. AI VII, p.42
Note on M41: k.jones: The n/f or m/f factor for CMC calculation.
Note on L47: k.jones: was just 1.0 in source template

## Uncertainty
UNCERTAINTY |  |  |  |  | Information for calculating expanded uncertainties |  |  |  | Parameters dependent on component | [6 empty] | This is a data table with B9 as the column input cell. | [13 empty] | CMC |  |  |  | IANZ
GEUM uncertainty calculation template |  |  |  |  | k-factor from degrees of freedom | [10 empty] | This must be re-calculated with new data sets. | [13 empty] | C uncert = 0.2/f+22C |  | delta uncert = 0.000027+0.00027/C
The component label is the key variable for the table |  |  |  |  | a | 1.96 |  | Model for uncertainty calculation. | Label | R alpha | u_R alpha | L/C alpha | u_L/C alpha | Aux L | Aux R | Label | L/C uncertainty | k | R/G uncertainty | k | G to delta | Range | L/C | mult | n | R/G | mult | m | f | pF | CMC/U |  | CMC/U | C |  | delta
[5 empty] | b | 2.372 |  | Base on formula for Universal bridge....ie assign an uncertainty to each term. | [8 empty] | 4.2891e-05 | 2.2948 | 0.00034529 | 2.15 | mrad
The 95% level of confidence expanded uncertainties are calculated using the Welch-Satterthwaite |  |  |  |  | c | 1.2737 |  |  | 1 microF3Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 1 microF3Y | 4.2889e-05 | 2.2948 | 0.00034528 | 2.15 | 0.034534 | 3Y | 0.99987 | 1e-06 | 0 | 0.0016999 | 0.001 | 0 | 1591.5 | 21.9973 | 0.51288 | 2.7e-05 | 0.78186 | 21.9973 | 0.51288 | 0.03 | 0.86872
formula to combine standard uncertainties. The individual degrees of freedom are derived either |  |  |  |  | d | 7.0962 |  |  | 0.1 microF4Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 0.1 microF4Y | 0.0042895 | 2.2948 | 0.034532 | 2.15 | 0.034534 | 4Y | 99.9999 | 1e-09 | 0 | 0.039698 | 1e-06 | 0 | 1591.5 | 2.2001 | 0.51291 | 2.7003e-05 | 0.78193 | 2.2001 | 0.51291 | 0.03 | 0.86872
from the number of measurements used to derive the uncertainty or an estimate | [8 empty] | 0.01 microF5Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | 0.01 microF5Y | 0.00042897 | 2.2948 | 0.0034534 | 2.15 | 0.034533 | 5Y | 10.0006 | 1e-09 | 0 | 0.0049297 | 1e-06 | 0 | 1591.5 | 0.22014 | 0.51317 | 2.7027e-05 | 0.78263 | 0.22014 | 0.51317 | 0.03 | 0.86872
of the percentage uncertainty in the uncertainty. | [8 empty] | 1000 pF6Y | 0 | 0 | 30 | 9 | 0 | 0 | 1000 pF6Y | 0.050847 | 2.1617 | 0.34523 | 2.15 | 0.034531 | 6Y | 999.7944 | 1e-12 | 0 | -0.228 | 1e-09 | 0 | 1591.5 | 0.022121 | 0.43505 | 2.727e-05 | 0.78972 | 0.022121 | 0.43505 | 0.03 | 0.86878
Component label | rpt 1 microF3Y | [7 empty] | 100 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | 100 pF7Y | 0.0050857 | 2.1617 | 0.03453 | 2.15 | 0.034534 | 7Y | 99.9916 | 1e-12 | 0 | 0.25258 | 1e-09 | 0 | 1591.5 | 0.0023255 | 0.45726 | 2.97e-05 | 0.86002 | 0.0023255 | 0.45726 | 0.03 | 0.8687
Universal bridge reactance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | 10 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | 10 pF7Y | 0.00059905 | 2.1066 | 0.0044009 | 2.0613 | 0.044024 | 7Y | 9.9969 | 1e-12 | 0 | 0.024598 | 1e-09 | 0 | 1591.5 | 0.0003456 | 0.57691 | 5.4008e-05 | 1.2268 | 0.0003456 | 0.57691 | 0.03 | 0.68144
| Standard | for | 999834 | L/C dial setting |  |  |  |  | 1 pF7Y | 0 | 0 | 40 | 12 | 0 | 0 | 1 pF7Y | 0.00041579 | 2.5532 | 0.0031066 | 2.1827 | 0.31071 | 7Y | 0.99986 | 1e-12 | 0 | 0.0025998 | 1e-09 | 0 | 1591.5 | 0.00014766 | 0.35514 | 0.00029704 | 0.95599 | 0.00014766 | 0.35514 | 0.03 | 0.096553
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | rpt 1 microF3Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | rpt 1 microF3Y | 4.2891e-05 | 2.2948 | 0.00034529 | 2.15 | 0.034534 | 3Y | 0.9999 | 1e-06 | 0 | 0.0017089 | 0.001 | 0 | 1591.5
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment |  | rpt 0.1 microF4Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | rpt 0.1 microF4Y | 0.0042896 | 2.2948 | 0.034534 | 2.15 | 0.034534 | 4Y | 100.0033 | 1e-09 | 0 | 0.041298 | 1e-06 | 0 | 1591.5
ac/dc error of standards | 2.4996 | 6.2479 | 30 | 5.5556 | 7.0266 | 5ppm considered two sigma limit, essentially relying on dc values of resistors |  |  | rpt 0.01 microF5Y | 0 | 0 | 7.5 | 27.5 | 0 | 0 | rpt 0.01 microF5Y | 0.00042899 | 2.2948 | 0.0034535 | 2.15 | 0.034533 | 5Y | 10.0009 | 1e-09 | 0 | 0.0050097 | 1e-06 | 0 | 1591.5
dial linearity | 1 | 1 | 30 | 5.5556 | 0.18 | 1 ppm of full scale |  |  | rpt 1000 pF6Y | 0 | 0 | 30 | 9 | 0 | 0 | rpt 1000 pF6Y | 0.050849 | 2.1617 | 0.34524 | 2.15 | 0.034531 | 6Y | 999.8274 | 1e-12 | 0 | -0.207 | 1e-09 | 0 | 1591.5
amplifier/ratio | 16.9972 | 288.9041 | 30 | 5.5556 | 15023.8001 | based on scatter of results in 1984, increased in 2004 and 2005 | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. |  | rpt 100 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | rpt 100 pF7Y | 0.0050861 | 2.1617 | 0.034533 | 2.15 | 0.034534 | 7Y | 99.9989 | 1e-12 | 0 | 0.18579 | 1e-09 | 0 | 1591.5
temperature coefficient of resistors | 1.9997 | 3.9987 | 50 | 2 | 7.9947 | depends on matching of Vishay components |  |  | rpt 10 pF7Y | 0 | 0 | 30 | 9 | 0 | 0 | rpt 10 pF7Y | 0.00059909 | 2.1066 | 0.0044011 | 2.0613 | 0.044023 | 7Y | 9.9978 | 1e-12 | 0 | 0.018799 | 1e-09 | 0 | 1591.5
temperature coefficient of capacitor | 0.99983 | 0.99967 | 30 | 5.5556 | 0.17988 | 2 ppm per degree for capacitor |  |  | rpt 1 pF7Y | 0 | 0 | 40 | 12 | 0 | 0 | rpt 1 pF7Y | 0.00041579 | 2.5532 | 0.0031066 | 2.1827 | 0.31071 | 7Y | 0.99986 | 1e-12 | 0 | 0.0020998 | 1e-09 | 0 | 1591.5
component stability | 4.9992 | 24.9917 | 50 | 2 | 312.2926 | some doubt on stability of capacitor in particular, increased in 2004 | Increase from 2 to 5 ppm to account for short-term instability of 1000 pF capacitor | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
L/C error due to phase error in R/G | 0.002728 | 7.442e-06 | - | 13 | 4.2602e-12 | 16 ppm of resistance dial setting, see comment for resistance, reduced to 5 ppm after 2004 cal | Link to conductance setting, previously just 2 counts | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
resolution | 1.2567 | 1.5792 | 50 | 2 | 1.247 | this might be range and frequency dependent | Previous was 2, use latest value from MSLT.E.023.005, 200/f | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 1 | 5000 | 0 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Sum of variance | 327.7213 |  | Sum of u^4/v | 15352.7208 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u combined (counts) | 18.1031 |  | effective v | 6.9956 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  |  | k | 2.3458 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 42.4667 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ For all frequencies. Additional noise typical at 50 Hz needs to be added | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Phase angle assumed small, ie no dial interaction. | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3/ Consistent with 1984 calibration, the IEEE paper and recent complement checks | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
4/ Updated following 2004/5 UB calibration, see AUDIREGC | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[22 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[22 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Universal bridge resistance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | for | 1705 | conductance dial setting |  | all references to the conductance dial are now to cell D34 (previously separate lookups back to results page) | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [17 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
G/R error due to phase error in L/C | 159.9734 | 25591.5015 | - | 13 | 50378842.254 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | Previous was 16, use latest value of 5 from MSLT.E.023.005 | Now return to 16 to stay consistent with CMCS | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
dial linearity | 10 | 100 | 30 | 5.5556 | 1800 |  | Previous value was 2, now decreased to 1 ppm following MSLT.E023.005 | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ac/dc error of standards | 0.042625 | 0.0018169 | 30 | 5.5556 | 5.942e-07 | 5ppm considered two sigma limit, essentially relying on dc values of resistors | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
resolution | 10 | 100 | 30 | 5.5556 | 1800 | this might be dependent on frequency and range, included below, increased from 0 to 1 in 2006, now 2000/f |  | Increase to 5 for 2011 measurements. | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
amplifier/ratio | 0.028985 | 0.00084013 | 30 | 5.5556 | 1.2705e-07 | based on scatter of results in 1984, increased following 2004 and 2005 calibration | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient of resistors | 0.0341 | 0.0011628 | 50 | 2 | 6.7606e-07 | depends on matching of Vishay components | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
component stability | 0.0341 | 0.0011628 | 30 | 5.5556 | 2.4338e-07 | stability of resistors only | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 50 | 2 | 0 |  | Previous value was 1, now increased to 2 ppm following MSLT.E023.005 | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 1 | 5000 | 0 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Sum of variance | 25791.5065 |  | Sum of u^4/v | 50382442.254 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u combined (counts) | 160.5973 |  | effective v | 13.203 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  |  | k | 2.15 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 345.2916 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ This scatter is consistent with Thompson's phase standards. | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Updated following 200/54 UB calibration see AUDIREGC | [21 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[22 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[22 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Conductance uncertainty for 16380 |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | [20 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [17 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ubridge uncertainty | 160.5973 | 25791.5065 | - | 13.203 | 50382442.254 | ubridge count uncertainty | [15 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3 month drift in resistance | 0 | 0 | 30 | 5.5556 | 0 | [16 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient | 0 | 0 | 20 | 12.5 | 0 | assume u_t = 0.5 degree
auxiliary balance senstivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget | [23 empty] | C uncert = (0.2/f +22C) pF with C in microF, f in Hz
| Sum of variance | 25791.5065 |  | Sum of u^4/v | 50382442.254
| u combined (counts) | 160.5973 |  | effective v | 13.203
|  |  |  | k | 2.15
Assumptions |  |  |  | expanded u (counts) | 345.2916
1/ |  |  |  |  | 0.00034529 | milliS
Reactance uncertainty for 16380 |  |  |  |  | k=a+b/v+c/v^2+d/v^3
Standard
Source of uncertainty | Uncertainties |  |  | Degrees of
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v
ubridge uncertainty | 18.1031 | 327.7213 | - | 6.9956 | 15352.7208
temperature coefficient | 4.6495 | 21.6178 | 30 | 5.5556 | 84.1194 | assume 0.5 deg at one sigma and look up the temp coefficient
3 month drift of standard | 0 | 0 | 30 | 5.5556 | 0
auxiliary balance sensitivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget
| Sum of variance | 349.3391 |  | Sum of u^4/v | 15436.8402
| u combined (count) | 18.6906 |  | effective v | 7.9056
|  |  |  | k | 2.2948
Assumptions |  |  |  | expanded u (count) | 42.8909
1/ |  |  |  |  | 4.2891e-05 | microF
Note on W8: k.jones: 30% of HP quoted TC for uncertainty
Note on W15: k.jones: 30% of HP quoted TC for uncertainty
Note on B18: t.stewart: Assume 0.5 degree as have below? KJ: agreed and modified
Note on B20: t.stewart: 5 or 16 ppm?? Latest comment and MSLT.E023.005 suggests 5 ppm. KJ: agree 5 in 2011 but in 2014, safer to stay with 16
Note on B21: k.jones: Sensitivity falls with f, 1 ppm at 2 kHz
Note on B37: t.stewart: 5 or 16 ppm?? Latest comment and MSLT.E023.005 suggests 5 ppm. KJ: agree 5 in 2011 but revert to 16 in 2014
Note on L37: k.jones: Revert after new CMCs developed and accepted.
Note on D38: k.jones: changed from 50 to 30 %, has impact on dof.
Note on B40: k.jones: say 1/f noise of 1 ppm at 100 Hz
Note on D40: k.jones: Was set to 1%, 30% compatible with CMC calculation
Note on L40: k.jones: 10 in last dial is more consistent with CMCs.
Note on B41: k.jones: was with x10 but!!! the dial setting takes care of the fact that there are 7 dials!!!!
Note on B57: k.jones: Was done in % for the Muirhead inductors.
Note on B59: k.jones: set to zero until rest of calculation is correct.
Note on B60: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B63: t.stewart: Should it be "u combined (%)" ?? KJ: agree %
Note on B72: k.jones: Was % for Muirheads
Note on B73: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.
Note on B74: k.jones: set to zero for now

## Report
Checked by:
2014-06-01 | [8 empty] | GMH51
| Label | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor |  | cap change | temp change | apparent tco
|  | (pF) | (pF) | (pF) |  | (milliradian) | (milliradian) |  |  | ppm
| 1000 pF6Y | 1000 | 999.7944 | 0.050847 | 2.1617 | -0.0018113 | 0.034531 | 2.15 | 19.71 | -33.0112 | -0.38 | 86.8716
| 100 pF7Y | 100 | 99.9949 | 0.0050857 | 2.1617 | 0.10866 | 0.034534 | 2.15 | 19.81 | -73.0103 | -0.28 | 260.7511
| 10 pF7Y | 10 | 9.9972 | 0.00059905 | 2.1066 | 0.10212 | 0.044024 | 2.0613 | 19.71 | -90.0333 | -0.35 | 257.238
| 1 pF7Y | 1 | 0.99989 | 0.00041579 | 2.5532 | 0.11608 | 0.31071 | 2.1827 | 19.7 | 0 | -0.35 | 0
|  | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor
|  | (mF) | (mF) | (mF) |  | (milliradian) | (milliradian)
| 1 microF3Y | 1 | 0.99987 | 4.2889e-05 | 2.2948 | 0.17002 | 0.034534 | 2.15 | 19.68 | -33.0066 | -0.25 | 132.0263
| 0.1 microF4Y | 0.1 | 0.1 | 4.2895e-06 | 2.2948 | 0.039699 | 0.034534 | 2.15 | 19.83 | -34.0024 | 0 | #DIV/0!
| 0.01 microF5Y | 0.01 | 0.010001 | 4.2897e-07 | 2.2948 | 0.049296 | 0.034533 | 2.15 | 19.87 | -32.0004 | 0.11 | -290.9129
2014-08-12
| Label | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor
|  | (pF) | (pF) | (pF) |  | (milliradian) | (milliradian)
| rpt 1000 pF6Y | 1000 | 999.8274 | 0.050849 | 2.1617 | 0.00028984 | 0.034531 | 2.15 | 20.09
| rpt 100 pF7Y | 100 | 100.0022 | 0.0050861 | 2.1617 | 0.00040501 | 0.034534 | 2.15 | 20.09
| rpt 10 pF7Y | 10 | 9.9981 | 0.00059909 | 2.1066 | -0.0015382 | 0.044023 | 2.0613 | 20.06
| rpt 1 pF7Y | 1 | 0.99989 | 0.00041579 | 2.5532 | -0.009751 | 0.31071 | 2.1827 | 20.05
|  | Nominal value | Capacitance | Expanded | Coverage | Loss angle | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor
|  | (mF) | (mF) | (mF) |  | (milliradian) | (milliradian)
| rpt 1 microF3Y | 1 | 0.9999 | 4.2891e-05 | 2.2948 | 0.17091 | 0.034534 | 2.15 | 19.93
| rpt 0.1 microF4Y | 0.1 | 0.1 | 4.2896e-06 | 2.2948 | 0.041297 | 0.034534 | 2.15 | 19.83
| rpt 0.01 microF5Y | 0.01 | 0.010001 | 4.2899e-07 | 2.2948 | 0.050094 | 0.034533 | 2.15 | 19.76
19.9729
0.1345
Note the columns used in vlookup for capacitance value.
1, 10 and 100 pF are column 12 with additonal range 7 correction. The other caps go to column 5, the usual UB calculation.
For loss angle
1, 10 and 100 are column 16, relying on AH2700.
1000 is column 13 using an additional 6Y correction.
The higher value caps are all column 10, the normal UB corrections.

## Range7
Universal Bridge | Range 7 was a design afterthought on the Universal Bridge and is expected to perform less well than the other ranges.
Consider the 2014 calibration of the UB that has, for the first time, included a 1 M Vishay in the calibration.
1 | At 1.6 kHz the G and R ac-dc differences look OK at -10.5 and 0.3 ppm respectively (using the same 'c' correction as for all other ranges, see 'Resistors' page).
We have no resistance phase standard for range 7.
2 | The 1 M Vishay gives a reading of -35 mH, which is much larger than the intrinsic <0.08 uH, but consistent with a mere 0.035 pF in parallel with the 1 M resistor
3 | GR100 is measured ('L and C' sheet) to have a 144 urad loss angle.
4 | AH2700 measures |  | 5.353 | urad for GR100 (see p.23 of UB III)
5 | Calc cap suggests |  | -0.02 | urad for GR100 (see Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls)
Item 1 gives basic confidence of ratio performance.
Item 2 is inconclusive.
Items 3 to 5 suggest a correction is required for phase on Range 7Y of the UB.
Unfortunately the difference in loss angles measured by the UB in July and August exceed the estimated uncertainty.
It is possible that the phase of the UB's 1 M internal standard was changed during the July calibration (physical knock?)
AH2700 | Therefore it seems prudent to rely on the AH2700 for the loss angle of capacitors below 1 nF.
See AH2700 manual page C-6 (appendix C), 'Accuracy of D vs. Frequency' graph
D accuracy
| 1 pF | 6 urad
| 10 pF | 4 urad
| 100 pF/ 1 nF | 3 urad
These are a factor of 10 or more better than our UB based CMCs.
Therefore use AH2700 results, but UB uncertainties (for CMC compliance).

## HP History
HISTORY |  |  | Checked by: | ACC
Instructions
Select rows 24:33 and carry out "Insert Rows". Copy cells in green block A14:N22 and "Paste" to cells A24:N32. Copy cells in blue block and "Paste Special, Values" to green block A14:N22.
2014-08-12 | S21942 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor |  | change | unc ppm |  |  |  |  | Lookup Label | Previous | Current | Days |  |  |  |  | 3 month change, ppm
HP 16387A 1 µF |  | 3Y | 999801 | 1696 |  |  |  | 0.9999 | 4.2891e-05 | 2.2948 | 0.17091 | 0.034534 | 2.15 |  | 16.3006 | 42.895 |  |  |  |  | A1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.43022
HP 16386A 0.1 µF |  | 4Y | 999928 | 394 |  |  |  | 0.1 | 4.2896e-06 | 2.2948 | 0.041297 | 0.034534 | 2.15 |  | 6.2985 | 42.895 |  |  |  |  | B10 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.16624
HP 16385A 0.01 µF |  | 5Y | 999986 | 496 |  |  |  | 0.010001 | 4.2899e-07 | 2.2948 | 0.050094 | 0.034533 | 2.15 |  | 17.298 | 42.8949 |  |  |  |  | C10 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 0.45655
HP 16384A 1000 pF |  | 6Y | 999660 | -205 |  |  |  | 999.8274 | 0.050849 | 2.1617 | 0.00028984 | 0.034531 | 2.15 |  | -7.2241 | 50.8575 |  |  |  |  | D1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | -0.19067
HP 16383A 100 pF |  | 7Y | 999865 | 2540 |  |  |  | 100.0022 | 0.0050861 | 2.1617 | 0.00040501 | 0.034534 | 2.15 |  | 111.3135 | 50.86 |  |  |  |  | E1 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 2.9379
HP 16382A 10 pF |  | 7Y | 99969 | 260 |  |  |  | 9.9981 | 0.00059909 | 2.1066 | -0.0015382 | 0.044023 | 2.0613 |  | -63.7133 | 59.9205 |  |  |  |  | E2 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | -1.6816
HP 16381A 1 pF |  | 7Y | 10004 | 40 |  |  |  | 0.99989 | 0.00041579 | 2.5532 | -0.009751 | 0.31071 | 2.1827 |  | 736.4257 | 415.8365 |  |  |  |  | E3 | 2005-02-22 | 2014-08-12 | 3410 |  |  |  |  | 19.4365
2011-05-17 | S21815 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999831 | 1710 |  |  |  | 0.9999 | 4.2891e-05 | 2.1617 | 0.17111 | 0.034534 | 2.15
HP 16386A 0.1 µF |  | 4Y | 999959 | 397 |  |  |  | 0.1 | 4.2896e-06 | 2.1617 | 0.039798 | 0.034534 | 2.15
HP 16385A 0.01 µF |  | 5Y | 1000016 | 505 |  |  |  | 0.010001 | 4.2899e-07 | 2.1066 | 0.050195 | 0.034533 | 2.15
HP 16384A 1000 pF |  | 6Y | 999702 | -190 |  |  |  | 999.8369 | 0.050849 | 2.1617 | -0.0055039 | 0.034531 | 2.15
HP 16383A 100 pF |  | 7Y | 999888 | 2100 |  |  |  | 100.0007 | 0.0050859 | 2.1617 | 0.0055476 | 0.034534 | 2.15
HP 16382A 10 pF |  | 7Y | 99974 | 213 |  |  |  | 9.9979 | 0.00059907 | 2.1066 | -0.00791 | 0.044023 | 2.0613
HP 16381A 1 pF |  | 7Y | 10005 | 36 |  |  |  | 1 | 0.00041579 | 2.5532 | 0.0040466 | 0.31068 | 2.1827
2005-02-22 | S21580 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
|  | Range | C | G |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999820 | 1653 |  |  |  | 0.99989 | 0.00011135 | 2.1283 | 0.16474 | 0.034393 | 2.15
HP 16386A 0.1 µF |  | 4Y | 999958 | 335 |  |  |  | 0.1 | 1.1136e-05 | 2.1262 | 0.033462 | 0.034393 | 2.15
HP 16385A 0.01 µF |  | 5Y | 1000004 | 500 |  |  |  | 0.010001 | 1.1141e-06 | 2.1038 | 0.049394 | 0.034393 | 2.15
HP 16384A 1000 pF |  | 6Y | 999703 | -227 |  |  |  | 999.8347 | 0.11137 | 2.1283 | -0.013604 | 0.034391 | 2.15
HP 16383A 100 pF |  | 7Y | 999863 | 1834 |  |  |  | 99.991 | 0.011161 | 2.1262 | -0.009865 | 0.034394 | 2.15
HP 16382A 10 pF |  | 7Y | 99988 | 193 |  |  |  | 9.9987 | 0.0011449 | 2.1038 | -0.0046376 | 0.043842 | 2.0613
HP 16381A 1 pF |  | 7Y | 9997 | 42 |  |  |  | 0.99915 | 0.00035273 | 2.0761 | 0.096162 | 0.30966 | 2.1828
2002-07-22 | S21385 | [5 empty] | Nominal Value | Capacitance | Uncertainty | Coverage | Loss Angle | Uncertainty | Coverage
Component |  | Range | Capacitance | Conductance |  |  | (mF) or (pF) | (mF) or (pF) | (mF) or (pF) | Factor | (milliradians) | (milliradians) | Factor
HP 16387A 1 µF |  | 3Y | 999800 | 1564 |  |  | 1 | 0.99989 | 0.00012347 | 2.4531 | 0.17381 | 0.039381 | 2.0872
HP 16386A 0.1 µF |  | 4Y | 999865 | 494 |  |  | 0.1 | 0.1 | 1.2347e-05 | 2.4545 | 0.060496 | 0.039381 | 2.0872
HP 16385A 0.01 µF |  | 4Y | 100000 | 54 |  |  | 0.01 | 0.010001 | 1.255e-06 | 2.3205 | 0.068991 | 0.03431 | 2.1402
HP 16384A 1000 pF |  | 6Y | 999725 | -96 |  |  | 1000 | 999.8515 | 0.074102 | 2.4353 | 0.0029977 | 0.03431 | 2.1402
HP 16383A 100 pF |  | 7Y | 999884 | 1530 |  |  | 100 | 99.9955 | 0.0074095 | 2.4406 | 0.0040066 | 0.03431 | 2.1402
HP 16382A 10 pF |  | 7Y | 100007 | 150 |  |  | 10 | 9.9993 | 0.00076244 | 2.09 | 0.019008 | 0.03431 | 2.1402
HP 16381A 1 pF |  | 7Y | 10027 | 24 |  |  | 1 | 0.99996 | 0.00049631 | 2.3634 | 0.026003 | 0.1071 | 2.1416
15/8/2000
Component |  | Range | C | G | True C |  | True G |  | tan _
HP 16387A 1 µF |  | 3Y | 999806 | 1734 | 0.99954 | µF | 0.0017336 | mS | 0.00017344
HP 16386A 0.1 µF |  | 4Y | 999945 | 600 | 100.0076 | nF | 0.059997 | µS | 5.9992e-05
HP 16385A 0.01 µF |  | 4Y | 100001 | 64 | 10.0014 | nF | 0.0063997 | µS | 6.3988e-05
HP 16384A 1000 pF |  | 6Y | 999733 | 14 | 999.4791 | pF | 0.013997 | nS | 1.4004e-06
HP 16383A 100 pF |  | 7Y | 999906 | 1080 | 103.4 | pF | 0.11178 | nS | 0.0001081
HP 16382A 10 pF |  | 7Y | 100012 | 110 | 10.3422 | pF | 0.011385 | nS | 0.00011008
HP 16381A 1 pF |  | 7Y | 10028 | 15 | 1.037 | pF | 0.0015525 | nS | 0.00014971
Results 27 November 1997
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999768 | 1756
HP 16386A 0.1 µF |  | 4Y | 999907 | 610
HP 16385A 0.01 µF |  | 5Y | 99998 | 56
HP 16384A 1000 pF |  | 6Y | 999756 | 98
HP 16383A 100 pF |  | 7Y | 999911 | 1100
HP 16382A 10 pF |  | 7Y | 100008 | 120
HP 16381A 1 pF |  | 7Y | 10028 | 10
Results 17 October 1995
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999775 | 1750
HP 16386A 0.1 µF |  | 4Y | 999904 | 590
HP 16385A 0.01 µF |  | 5Y | 1000254 | 700
HP 16384A 1000 pF |  | 6Y | 999770 | 69
HP 16383A 100 pF |  | 7Y | 1000019 | 1506
HP 16382A 10 pF |  | 7Y | 100020 | 150
HP 16381A 1 pF |  | 7Y | 10028 | 16
Results 5 August 1993
Component |  | Range | Capcitance | Conductance
HP 16387A 1 µF |  | 3Y | 999746 | 1813
HP 16386A 0.1 µF |  | 4Y | 999882 | 659
HP 16385A 0.01 µF |  | 5Y | 1000131 | 560
HP 16384A 1000 pF |  | 6Y | 999732 | 113
HP 16383A 100 pF |  | 7Y | 999893 | 1520
HP 16382A 10 pF |  | 7Y | 100014 | 150
HP 16381A 1 pF |  | 7Y | 10027 | 16
Date | Job | Report S No.
1993-08-05 |  | 20878
1995-10-17 | 71623 | 21066
1997-11-27 | 71750 | 21250
2000-08-15 | 73040 | 21385
2002-07-22 | 71250 | 21464
2005-02-22 | 7140032 | 21580
| 1993-08-04 |  |  |  | 1995-10-17 |  |  |  | 1997-11-27 |  |  |  | 2000-08-09 |  |  |  | 2002-07-22 |  |  |  | 2005-02-22 |  |  |  | 2011-05-17 |  |  |  | 2014-08-12
| C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta | C | uC | delta | udelta
HP 16387A 1 µF | 0.99984 | 0.00019997 | 0.18 | 0.1 | 0.99986 | 0.00012 | 0.175 | 0.039 | 0.99985 | 0.00012 | 0.176 | 0.039 | 0.99989 | 0.00012 | 0.17 | 0.04 | 0.99988 | 0.00012 | 0.16 | 0.04 | 0.99989 | 0.00011135 | 0.16474 | 0.034393 | 0.9999 | 4.2891e-05 | 0.17111 | 0.034534 | 0.9999 | 4.2891e-05 | 0.17091 | 0.034534
HP 16386A 0.1 µF | 0.099997 | 1.9999e-05 | 0.07 | 0.1 | 0.099998 | 1.2e-05 | 0.059 | 0.039 | 0.099999 | 1.2e-05 | 0.061 | 0.039 | 0.1 | 1.2e-05 | 0.06 | 0.04 | 0.1 | 1.2e-05 | 0.05 | 0.04 | 0.1 | 1.1136e-05 | 0.033462 | 0.034393 | 0.1 | 4.2896e-06 | 0.039798 | 0.034534 | 0.1 | 4.2896e-06 | 0.041297 | 0.034534
HP 16385A 0.01 µF | 0.01 | 2.0001e-06 | 0.08 | 0.1 | 0.010003 | 1.3e-06 | 0.064 | 0.034 | 0.010001 | 1.3e-06 | 0.059 | 0.034 | 0.010001 | 1.3e-06 | 0.07 | 0.04 | 0.010001 | 1.3e-06 | 0.06 | 0.03 | 0.010001 | 1.1141e-06 | 0.049394 | 0.034393 | 0.010001 | 4.2899e-07 | 0.050195 | 0.034533 | 0.010001 | 4.2899e-07 | 0.050094 | 0.034533
HP 16384A 1000 pF | 999.87 | 0.19997 | 0 | 0.1 | 999.9 | 0.07 | 0 | 0.034 | 999.89 | 0.07 | 0 | 0.034 | 999.85 | 0.07 | 0 | 0.03 | 999.85 | 0.07 | 0.01 | 0.03 | 999.8347 | 0.11137 | -0.013604 | 0.034391 | 999.8369 | 0.050849 | -0.0055039 | 0.034531 | 999.8274 | 0.050849 | 0.00028984 | 0.034531
HP 16383A 100 pF | 99.993 | 0.019999 | 0 | 0.1 | 100.005 | 0.008 | 0 | 0.034 | 99.994 | 0.008 | 0 | 0.034 | 99.995 | 0.007 | 0 | 0.03 | 99.992 | 0.007 | 0.07 | 0.03 | 99.991 | 0.011161 | -0.009865 | 0.034394 | 100.0007 | 0.0050859 | 0.0055476 | 0.034534 | 100.0022 | 0.0050861 | 0.00040501 | 0.034534
HP 16382A 10 pF | 9.9989 | 0.0019998 | 0 | 0.1 | 9.9997 | 0.0008 | 0 | 0.034 | 9.9984 | 0.0008 | 0 | 0.034 | 9.9993 | 0.0008 | 0.02 | 0.03 | 9.9986 | 0.0008 | 0.06 | 0.03 | 9.9987 | 0.0011449 | -0.0046376 | 0.043842 | 9.9979 | 0.00059907 | -0.00791 | 0.044023 | 9.9981 | 0.00059909 | -0.0015382 | 0.044023
HP 16381A 1 pF | 0.99957 | 0.00019991 | 0.16 | 0.1 | 0.99996 | 0.00026 | 0.16 | 0.11 | 0.99986 | 0.0005 | 0.1 | 0.11 | 0.99996 | 0.0005 | 0.03 | 0.11 | 0.99986 | 0.0005 | 0.09 | 0.11 | 0.99915 | 0.00035273 | 0.096162 | 0.30966 | 1 | 0.00041579 | 0.0040466 | 0.31068 | 0.99989 | 0.00041579 | -0.009751 | 0.31071
HP 16387A 1 µF
Date | C | uC | delta | udelta
1993-08-05 | 0.99984 | 0.00019997 | 0.18 | 0.1
1995-10-17 | 0.99986 | 0.00012 | 0.175 | 0.039
1997-11-27 | 0.99985 | 0.00012 | 0.176 | 0.039
2000-08-15 | 0.99989 | 0.00012 | 0.17 | 0.04
2002-07-22 | 0.99988 | 0.00012 | 0.16 | 0.04
2005-02-22 | 0.99989 | 0.00011135 | 0.16474 | 0.034393
2011-05-17 | 0.9999 | 4.2891e-05 | 0.17111 | 0.034534
2014-08-12 | 0.9999 | 4.2891e-05 | 0.17091 | 0.034534
Note on X102: k.jones: set to zero in report for all four low value caps.

## AH2700_check
from Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls, A28 to C45
2014-08-11 | [6 empty] | Ref Value |  |  | Correction
1.5915 kHz |  | pF | nS | temperature | ppm | urad | Real | Im |  | Real | Im
| AH11A | 10 | 6e-06 |  |  | 0.060002
| AH11B | 10 | 2.69e-05 |  | -2.86 | 0.26901 | -4.7172 | 0.306 |  | -1.8572 | 0.036991
| AH11C | 99.9998 | 0.001501 |  | -2.37 | 1.5011 | -4.1912 | 0.27809 |  | -1.8212 | -1.223
| ES14 | 0.50005 | 7.26e-05 |  | 94.8 | 14.5191 | 92.7749 | 0.83086 |  | -2.0251 | -13.6882
| GR1000A | 1000.0296 | 0.12759 | 23.994 | 29.6 | 12.759 | 27.8198 | -6.1878 |  | -1.7802 | -18.9468
| GR1000B | 1000.016 | 0.13459 | 23.994 | 16 | 13.4592 | 14.1798 | -6.7618 |  | -1.8202 | -20.221
| GR100 | 100.0039 | 0.0057608 | 23.994 | 38.6 | 5.7608 | 36.5638 | -0.015909 |  | -2.0362 | -5.7767
| GR10 | 10.0002 | 0.0006515 | 23.994 | 18.5 | 6.5151 | 16.3708 | -0.106 |  | -2.1292 | -6.6211
[7 empty] | Corrected, 11 August |  |  |  |  | Report, 1 July |  | Report -AH |  | As fraction of report uncertainty
[7 empty] | Real | Imaginary | pF | mrad |  | pF | mrad | pF | mrad
| HP 1 pF | 0.99993 | -0.0001032 | 19.78 | -65.7 | -10.321 | -67.7251 | -24.0092 | 0.99993 | -0.024009 |  | 0.99989 | 0.11608 | -4.2729e-05 | 0.14009 | -0.10277 | 0.054866
| HP 10 pF | 9.9987 | -0.0001579 |  | -126.3 | -1.5792 | -128.1212 | -2.8022 | 9.9987 | -0.0028022 |  | 9.9972 | 0.10212 | -0.0015235 | 0.10492 | -2.5431 | 0.049805
| HP 100 pF | 100.0007 | 0.000408 |  | 6.6 | 0.40801 | 4.7788 | -0.81495 | 100.0005 | -0.00081495 |  | 99.9949 | 0.10866 | -0.0056228 | 0.10948 | -1.1056 | 0.050643
| HP 1000 pF | 999.8414 | 0.011 | 19.82 | -158.62 | 1.1002 | -160.4002 | -17.8466 | 999.8396 | -0.017847 |  | 999.7944 | -0.0018113 | -0.045157 | 0.016035 | -0.88809 | 0.0074178
Report, 12 August
[12 empty] | pF | mrad
[9 empty] | 0.99993 | -0.024009 |  | 0.99989 | -0.009751 | -4.2729e-05 | 0.014258 | -0.10277 | 0.0055844
[9 empty] | 9.9987 | -0.0028022 |  | 9.9981 | -0.0015382 | -0.0006234 | 0.001264 | -1.0406 | 0.00059999
[9 empty] | 100.0005 | -0.00081495 |  | 100.0022 | 0.00040501 | 0.0016779 | 0.00122 | 0.32992 | 0.00056435
[9 empty] | 999.8396 | -0.017847 |  | 999.8274 | 0.00028984 | -0.012152 | 0.018136 | -0.239 | 0.0083898

## RoomLogger
2014-08-12 | 20.52 | 41.34
2014-08-12 00:05 | 20.53 | 41.22
2014-08-12 00:10 | 20.51 | 41.25
2014-08-12 00:15 | 20.45 | 41.34
2014-08-12 00:20 | 20.5 | 41.1
2014-08-12 00:25 | 20.49 | 41.22
2014-08-12 00:30 | 20.53 | 40.95
2014-08-12 00:35 | 20.51 | 41.22
2014-08-12 00:40 | 20.5 | 41.47
2014-08-12 00:45 | 20.41 | 41.3
2014-08-12 00:50 | 20.44 | 41.44
2014-08-12 00:55 | 20.45 | 41
2014-08-12 01:00 | 20.45 | 41.2
2014-08-12 01:05 | 20.5 | 40.9
2014-08-12 01:10 | 20.52 | 40.98
2014-08-12 01:15 | 20.49 | 40.98
2014-08-12 01:20 | 20.51 | 41.03
2014-08-12 01:25 | 20.52 | 41.15
2014-08-12 01:30 | 20.51 | 41.15
2014-08-12 01:35 | 20.49 | 41.03
2014-08-12 01:40 | 20.53 | 41.12
2014-08-12 01:45 | 20.53 | 41.2
2014-08-12 01:50 | 20.61 | 41.2
2014-08-12 01:55 | 20.61 | 41.12
2014-08-12 02:00 | 20.63 | 41
2014-08-12 02:05 | 20.62 | 41
2014-08-12 02:10 | 20.6 | 41.03
2014-08-12 02:15 | 20.53 | 40.88
2014-08-12 02:20 | 20.56 | 40.93
2014-08-12 02:25 | 20.58 | 40.85
2014-08-12 02:30 | 20.6 | 40.93
2014-08-12 02:35 | 20.57 | 41.03
2014-08-12 02:40 | 20.55 | 40.98
2014-08-12 02:45 | 20.55 | 41.03
2014-08-12 02:50 | 20.58 | 41.17
2014-08-12 02:55 | 20.61 | 40.95
2014-08-12 03:00 | 20.61 | 40.85
2014-08-12 03:05 | 20.61 | 40.78
2014-08-12 03:10 | 20.62 | 40.71
2014-08-12 03:15 | 20.61 | 40.71
2014-08-12 03:20 | 20.61 | 40.66
2014-08-12 03:25 | 20.62 | 40.58
2014-08-12 03:30 | 20.62 | 40.58
2014-08-12 03:35 | 20.62 | 40.51
2014-08-12 03:40 | 20.61 | 40.51
2014-08-12 03:45 | 20.62 | 40.39
2014-08-12 03:50 | 20.61 | 40.41
2014-08-12 03:55 | 20.6 | 40.39
2014-08-12 04:00 | 20.58 | 40.46
2014-08-12 04:05 | 20.65 | 40.46
2014-08-12 04:10 | 20.62 | 40.44
2014-08-12 04:15 | 20.62 | 40.51
2014-08-12 04:20 | 20.58 | 40.58
2014-08-12 04:25 | 20.6 | 40.61
2014-08-12 04:30 | 20.58 | 40.54
2014-08-12 04:35 | 20.57 | 40.46
2014-08-12 04:40 | 20.61 | 40.41
2014-08-12 04:45 | 20.61 | 40.24
2014-08-12 04:50 | 20.58 | 40.34
2014-08-12 04:55 | 20.6 | 40.31
2014-08-12 05:00 | 20.66 | 40.27
2014-08-12 05:05 | 20.66 | 40.17
2014-08-12 05:10 | 20.66 | 40.22
2014-08-12 05:15 | 20.6 | 40.24
2014-08-12 05:20 | 20.66 | 40.14
2014-08-12 05:25 | 20.63 | 40.12
2014-08-12 05:30 | 20.61 | 40.19
2014-08-12 05:35 | 20.67 | 40.22
2014-08-12 05:40 | 20.68 | 40.31
2014-08-12 05:45 | 20.67 | 40.31
2014-08-12 05:50 | 20.66 | 40.36
2014-08-12 05:55 | 20.63 | 40.44
2014-08-12 06:00 | 20.65 | 40.44
2014-08-12 06:05 | 20.65 | 40.41
2014-08-12 06:10 | 20.62 | 40.54
2014-08-12 06:15 | 20.65 | 40.34
2014-08-12 06:20 | 20.69 | 40.41
2014-08-12 06:25 | 20.67 | 40.44
2014-08-12 06:30 | 20.69 | 40.44
2014-08-12 06:35 | 20.68 | 40.41
2014-08-12 06:40 | 20.69 | 40.49
2014-08-12 06:45 | 20.68 | 40.44
2014-08-12 06:50 | 20.68 | 40.49
2014-08-12 06:55 | 20.69 | 40.41
2014-08-12 07:00 | 20.74 | 40.44
2014-08-12 07:05 | 20.72 | 40.36
2014-08-12 07:10 | 20.74 | 40.36
2014-08-12 07:15 | 20.73 | 40.36
2014-08-12 07:20 | 20.68 | 40.58
2014-08-12 07:25 | 20.69 | 40.44
2014-08-12 07:30 | 20.69 | 40.58
2014-08-12 07:35 | 20.69 | 40.54
2014-08-12 07:40 | 20.72 | 40.66
2014-08-12 07:45 | 20.72 | 40.54
2014-08-12 07:50 | 20.67 | 40.76
2014-08-12 07:55 | 20.69 | 40.56
2014-08-12 08:00 | 20.69 | 40.63
2014-08-12 08:05 | 20.74 | 40.54
2014-08-12 08:10 | 20.76 | 40.76
2014-08-12 08:15 | 20.71 | 40.95
2014-08-12 08:20 | 20.69 | 40.9
2014-08-12 08:25 | 20.65 | 41.07
2014-08-12 08:30 | 20.68 | 41.05
2014-08-12 08:35 | 20.66 | 41.07
2014-08-12 08:40 | 20.69 | 41
2014-08-12 08:45 | 20.69 | 40.95
2014-08-12 08:50 | 20.68 | 41.07
2014-08-12 08:55 | 20.68 | 40.95
2014-08-12 09:00 | 20.71 | 40.9
2014-08-12 09:05 | 20.68 | 40.83
2014-08-12 09:10 | 20.69 | 40.9
2014-08-12 09:15 | 20.71 | 40.71
2014-08-12 09:20 | 20.68 | 40.68
2014-08-12 09:25 | 20.67 | 40.88
2014-08-12 09:30 | 20.71 | 40.78
2014-08-12 09:35 | 20.77 | 40.83
2014-08-12 09:40 | 20.84 | 40.76
2014-08-12 09:45 | 20.84 | 40.81
2014-08-12 09:50 | 20.79 | 40.9
2014-08-12 09:55 | 20.82 | 41.05
2014-08-12 10:00 | 20.74 | 40.98
2014-08-12 10:05 | 20.78 | 40.93
2014-08-12 10:10 | 20.8 | 40.73
2014-08-12 10:15 | 20.73 | 40.76
2014-08-12 10:20 | 20.76 | 40.71
2014-08-12 10:25 | 20.79 | 40.56
2014-08-12 10:30 | 20.79 | 40.44
2014-08-12 10:35 | 20.82 | 40.41
2014-08-12 10:40 | 20.85 | 40.29
2014-08-12 10:45 | 20.88 | 40.24
2014-08-12 10:50 | 20.9 | 40.31
2014-08-12 10:55 | 20.89 | 40.41
2014-08-12 11:00 | 20.89 | 40.41 | 20.8527 | 40.2392
2014-08-12 11:05 | 20.84 | 40.46 | 0.031542 | 0.1982
2014-08-12 11:10 | 20.87 | 40.36
2014-08-12 11:15 | 20.85 | 40.39
2014-08-12 11:20 | 20.8 | 40.41
2014-08-12 11:25 | 20.87 | 40.27
2014-08-12 11:30 | 20.88 | 40.22
2014-08-12 11:35 | 20.9 | 40.27
2014-08-12 11:40 | 20.87 | 40.17
2014-08-12 11:45 | 20.94 | 40.07
2014-08-12 11:50 | 20.88 | 40.04
2014-08-12 11:55 | 20.87 | 39.97
2014-08-12 12:00 | 20.88 | 39.95
2014-08-12 12:05 | 20.87 | 39.87
2014-08-12 12:10 | 20.85 | 39.95
2014-08-12 12:15 | 20.84 | 40.07
2014-08-12 12:20 | 20.85 | 40.17
2014-08-12 12:25 | 20.83 | 40.14
2014-08-12 12:30 | 20.84 | 40.22
2014-08-12 12:35 | 20.82 | 40.24
2014-08-12 12:40 | 20.88 | 40.09
2014-08-12 12:45 | 20.88 | 40.09
2014-08-12 12:50 | 20.89 | 39.97
2014-08-12 12:55 | 20.9 | 40.02
2014-08-12 13:00 | 20.85 | 39.92
2014-08-12 13:05 | 20.87 | 40.02
2014-08-12 13:10 | 20.79 | 40.17
2014-08-12 13:15 | 20.8 | 40.27
2014-08-12 13:20 | 20.82 | 40.34
2014-08-12 13:25 | 20.83 | 40.17
2014-08-12 13:30 | 20.92 | 40.14
2014-08-12 13:35 | 20.82 | 40.24
2014-08-12 13:40 | 20.83 | 40.31
2014-08-12 13:45 | 20.87 | 40.07
2014-08-12 13:50 | 20.84 | 40.17
2014-08-12 13:55 | 20.84 | 40.17
2014-08-12 14:00 | 20.85 | 40.22
2014-08-12 14:05 | 20.87 | 40.29
2014-08-12 14:10 | 20.87 | 40.34
2014-08-12 14:15 | 20.82 | 40.46
2014-08-12 14:20 | 20.85 | 40.49
2014-08-12 14:25 | 20.84 | 40.56
2014-08-12 14:30 | 20.83 | 40.44
2014-08-12 14:35 | 20.82 | 40.41
2014-08-12 14:40 | 20.82 | 40.46
2014-08-12 14:45 | 20.87 | 40.41
2014-08-12 14:50 | 20.82 | 40.44
2014-08-12 14:55 | 20.83 | 40.58
2014-08-12 15:00 | 20.82 | 40.81
2014-08-12 15:05 | 20.83 | 40.9
2014-08-12 15:10 | 20.8 | 40.85
2014-08-12 15:15 | 20.83 | 40.78
2014-08-12 15:20 | 20.79 | 40.71
2014-08-12 15:25 | 20.84 | 40.41
2014-08-12 15:30 | 20.83 | 40.41
2014-08-12 15:35 | 20.79 | 40.58
2014-08-12 15:40 | 20.79 | 40.63
2014-08-12 15:45 | 20.85 | 40.73
2014-08-12 15:50 | 20.82 | 40.76
2014-08-12 15:55 | 20.83 | 40.76
2014-08-12 16:00 | 20.82 | 40.63
2014-08-12 16:05 | 20.82 | 40.71
2014-08-12 16:10 | 20.84 | 40.66
2014-08-12 16:15 | 20.83 | 40.76
2014-08-12 16:20 | 20.83 | 40.88
2014-08-12 16:25 | 20.82 | 40.88
2014-08-12 16:30 | 20.78 | 40.81
2014-08-12 16:35 | 20.82 | 40.78
2014-08-12 16:40 | 20.84 | 40.71
2014-08-12 16:45 | 20.84 | 40.63
2014-08-12 16:50 | 20.82 | 40.61
2014-08-12 16:55 | 20.84 | 40.41
2014-08-12 17:00 | 20.79 | 40.71
2014-08-12 17:05 | 20.76 | 40.76
2014-08-12 17:10 | 20.72 | 41.1
2014-08-12 17:15 | 20.79 | 40.9
2014-08-12 17:20 | 20.78 | 40.88
2014-08-12 17:25 | 20.83 | 40.73
2014-08-12 17:30 | 20.79 | 40.76
2014-08-12 17:35 | 20.84 | 40.63
2014-08-12 17:40 | 20.84 | 40.58
2014-08-12 17:45 | 20.84 | 40.46
2014-08-12 17:50 | 20.85 | 40.54
2014-08-12 17:55 | 20.79 | 40.68
2014-08-12 18:00 | 20.83 | 40.88
2014-08-12 18:05 | 20.8 | 40.76
2014-08-12 18:10 | 20.83 | 40.98
2014-08-12 18:15 | 20.84 | 40.95
2014-08-12 18:20 | 20.85 | 40.93
2014-08-12 18:25 | 20.89 | 41
2014-08-12 18:30 | 20.85 | 40.81
2014-08-12 18:35 | 20.88 | 40.83
2014-08-12 18:40 | 20.88 | 40.83
2014-08-12 18:45 | 20.85 | 40.98
2014-08-12 18:50 | 20.85 | 40.81
2014-08-12 18:55 | 20.88 | 40.73
2014-08-12 19:00 | 20.9 | 40.88
2014-08-12 19:05 | 20.92 | 40.98
2014-08-12 19:10 | 20.89 | 41
2014-08-12 19:15 | 20.93 | 41.15
2014-08-12 19:20 | 20.87 | 41.2
2014-08-12 19:25 | 20.9 | 41.1
2014-08-12 19:30 | 20.89 | 41.05
2014-08-12 19:35 | 20.88 | 40.9
2014-08-12 19:40 | 20.89 | 40.95
2014-08-12 19:45 | 20.84 | 41.03
2014-08-12 19:50 | 20.88 | 41.17
2014-08-12 19:55 | 20.93 | 41.25
2014-08-12 20:00 | 20.9 | 41.17
2014-08-12 20:05 | 20.88 | 41.27
2014-08-12 20:10 | 20.83 | 41.3
2014-08-12 20:15 | 20.87 | 41.07
2014-08-12 20:20 | 20.85 | 41.03
2014-08-12 20:25 | 20.85 | 40.98
2014-08-12 20:30 | 20.88 | 41.03
2014-08-12 20:35 | 20.9 | 41.07
2014-08-12 20:40 | 20.88 | 41.15
2014-08-12 20:45 | 20.84 | 41.27
2014-08-12 20:50 | 20.88 | 41.2
2014-08-12 20:55 | 20.89 | 41.17
2014-08-12 21:00 | 20.87 | 41.05
2014-08-12 21:05 | 20.92 | 41.05
2014-08-12 21:10 | 20.89 | 40.85
2014-08-12 21:15 | 20.87 | 41.05
2014-08-12 21:20 | 20.85 | 41.12
2014-08-12 21:25 | 20.84 | 41.12
2014-08-12 21:30 | 20.83 | 41.05
2014-08-12 21:35 | 20.84 | 41.12
2014-08-12 21:40 | 20.83 | 41.07
2014-08-12 21:45 | 20.87 | 40.93
2014-08-12 21:50 | 20.84 | 40.88
2014-08-12 21:55 | 20.87 | 40.83
2014-08-12 22:00 | 20.87 | 40.88
2014-08-12 22:05 | 20.84 | 41.1
2014-08-12 22:10 | 20.84 | 41.03
2014-08-12 22:15 | 20.85 | 41
2014-08-12 22:20 | 20.87 | 40.88
2014-08-12 22:25 | 20.85 | 40.81
2014-08-12 22:30 | 20.88 | 40.66
2014-08-12 22:35 | 20.89 | 40.71
2014-08-12 22:40 | 20.89 | 41
2014-08-12 22:45 | 20.93 | 40.85
2014-08-12 22:50 | 20.85 | 40.95
2014-08-12 22:55 | 20.87 | 40.76
2014-08-12 23:00 | 20.9 | 40.68
2014-08-12 23:05 | 20.92 | 40.68
2014-08-12 23:10 | 20.92 | 40.63
2014-08-12 23:15 | 20.94 | 40.46
2014-08-12 23:20 | 20.94 | 40.34
2014-08-12 23:25 | 20.95 | 40.36
2014-08-12 23:30 | 20.92 | 40.58
2014-08-12 23:35 | 20.88 | 40.63
2014-08-12 23:40 | 20.94 | 40.68
2014-08-12 23:45 | 20.93 | 40.58
2014-08-12 23:50 | 20.93 | 40.56
2014-08-12 23:55 | 20.93 | 40.41