## Overview
In the last report (S21959, 2015) both the AH2700 and the UB were used.
Uncertainty has a formula for series resistance impact on loss angle that should be done in Python.
The final report used UB results for capacitance, but AH2700A results for tan delta up to step 10 on the 100 pF dial.
The two-terminal nature of the box means we stick with about a 1 pF uncertainty which is not challenging for the UB.
There is a 'coax zero', a 'copper jig zero' and a 'box zero' to consider.
The 'box zero' is what should be subtracted off the change values.
The 'copper jig zero' is what is subtracted to give the quoted 'box zero'
The 'box zero' readings will depend on range for the UB while the quoted 'box zero' - 'jig zero' is done on range 6.
It seems sensible to start with the usual UB style template but to also have an AH2700 readings sheet that can be read in.
Perhaps we should start with an AH2700A class structured like the UB class and do a parallel calibration.
Relying on pasting the different result calculations would be OK for now.
Ultimately it can be put together into a generalised UUT class that pulls in both the UB and AH2700 classes.
Expect that AH2700A results will be computer collected so that they come with a standard deviation.
It is good practice to check the coax zero on all ranges and results have this coax zero subtracted
This is only essential if the measured component is coaxially connected (e.g. the cross check with GR1000 set)
When only the box zero is being subtracted we can use dummy values for the coax zero.
The box zero is measured with the copper zero subtracted.

## pyUBreadings
Template for capacitors
Enter, say, 5 ppm as 5e-6
Capacitors |  |  |  | Client | [5 empty] | 1 sigma uncertainties from specifications
Date |  |  |  | Job number | [5 empty] | specs say -140 ppm for capacitance, assume 100 ppm for G
|  |  |  | S Number | [5 empty] | Standard uncertainty of 70 ppm for C and 100 ppm for G.
|  |  | Data from lab book |  |  |  | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff C | u | df | Temp coeff G | u | df | Other | u | df
| Hz |  |  |  | Hz | °C
coax zero | 1591.5 | 3Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
box_zero | 1591.5 | 3Y | 49 | 0 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A1 | 1591.5 | 3Y | 100013 | 40 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A2 | 1591.5 | 3Y | 200413 | 90 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A3 | 1591.5 | 3Y | 300376 | 119 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A4 | 1591.5 | 3Y | 400767 | 209 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A5 | 1591.5 | 3Y | 500576 | 346 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A6 | 1591.5 | 3Y | 600560 | 84 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A7 | 1591.5 | 3Y | 700940 | 522 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A8 | 1591.5 | 3Y | 800803 | 736 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A9 | 1591.5 | 3Y | 901292 | 1013 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
A10 | 1591.5 | 3Y | 1001258 | 1327 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
coax zero | 1591.5 | 4Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
box_zero | 1591.5 | 4Y | 500 | 10 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B1 | 1591.5 | 4Y | 100618 | 47 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B2 | 1591.5 | 4Y | 200526 | 67 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B3 | 1591.5 | 4Y | 300641 | 70 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B4 | 1591.5 | 4Y | 400667 | 64 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B5 | 1591.5 | 4Y | 500749 | 57 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B6 | 1591.5 | 4Y | 600867 | -61 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B7 | 1591.5 | 4Y | 700768 | 99 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B8 | 1591.5 | 4Y | 800879 | 189 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B9 | 1591.5 | 4Y | 900901 | 329 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
B10 | 1591.5 | 4Y | 1001011 | 514 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
box_zero | 1591.5 | 5Y | 5001 | 123 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C1 | 1591.5 | 5Y | 104502 | 163 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C2 | 1591.5 | 5Y | 204731 | 157 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C3 | 1591.5 | 5Y | 304202 | 160 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C4 | 1591.5 | 5Y | 404451 | 146 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C5 | 1591.5 | 5Y | 504786 | 116 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C6 | 1591.5 | 5Y | 604205 | 67 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C7 | 1591.5 | 5Y | 704439 | 137 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C8 | 1591.5 | 5Y | 803838 | 238 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C9 | 1591.5 | 5Y | 904057 | 329 | 1591.5 | 19.78 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
C10 | 1591.5 | 5Y | 1003463 | 493 | 1591.5 | 19.86 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
box_zero | 1591.5 | 6Y | 50024 | 1095 | 1591.5 | 19.67 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D1 | 1591.5 | 6Y | 146641 | 1150 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D2 | 1591.5 | 6Y | 245729 | 1138 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D3 | 1591.5 | 6Y | 342214 | 1170 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D4 | 1591.5 | 6Y | 442013 | 1123 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D5 | 1591.5 | 6Y | 538500 | 1140 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D6 | 1591.5 | 6Y | 641458 | 1058 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D7 | 1591.5 | 6Y | 740537 | 1164 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D8 | 1591.5 | 6Y | 837056 | 1308 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D9 | 1591.5 | 6Y | 936847 | 1426 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
D10 | 1591.5 | 6Y | 1033147 | 1534 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
coax zero | 1591.5 | 6Y | 0 | 17 | 1591.4 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
GR1000A | 1591.5 | 6Y | 999924 | -113 | 1591.4 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
GR1000B | 1591.5 | 6Y | 999907 | -83 | 1591.4 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E1 | 1591.5 | 6Y | 60487 | 1341 | 1591.7 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E2 | 1591.5 | 6Y | 71300 | 1391 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E3 | 1591.5 | 6Y | 81517 | 1332 | 1591.7 | 19.6 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E4 | 1591.5 | 6Y | 91813 | 1324 | 1591.7 | 19.6 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E5 | 1591.5 | 6Y | 101710 | 1334 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E6 | 1591.5 | 6Y | 111557 | 1338 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E7 | 1591.5 | 6Y | 120990 | 1319 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E8 | 1591.5 | 6Y | 130672 | 1315 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E9 | 1591.5 | 6Y | 140694 | 1310 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
E10 | 1591.5 | 6Y | 150067 | 1304 | 1591.7 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
CuZero | 1591.5 | 6Y | 888 | 15 | 1591.6 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10
[5 empty] | average | 19.7139
[5 empty] | stdev | 0.058394

## room_logger
mean | 20.4883 | 46.8025
sd | 0.06234 | 5.3167
2020-10-14 | 20.5 | 53.5
2020-10-14 00:05 | 20.5 | 53.5
2020-10-14 00:10 | 20.6 | 53.4
2020-10-14 00:15 | 20.6 | 53.4
2020-10-14 00:20 | 20.5 | 53.4
2020-10-14 00:25 | 20.6 | 53.3
2020-10-14 00:30 | 20.5 | 53.5
2020-10-14 00:35 | 20.6 | 53.4
2020-10-14 00:40 | 20.5 | 53.4
2020-10-14 00:45 | 20.5 | 53.3
2020-10-14 00:50 | 20.5 | 53.3
2020-10-14 00:55 | 20.6 | 53.2
2020-10-14 01:00 | 20.5 | 53.2
2020-10-14 01:05 | 20.5 | 53.1
2020-10-14 01:10 | 20.6 | 53
2020-10-14 01:15 | 20.5 | 53
2020-10-14 01:20 | 20.5 | 53
2020-10-14 01:25 | 20.5 | 53
2020-10-14 01:30 | 20.5 | 52.9
2020-10-14 01:35 | 20.5 | 53
2020-10-14 01:40 | 20.5 | 52.9
2020-10-14 01:45 | 20.6 | 52.7
2020-10-14 01:50 | 20.6 | 52.7
2020-10-14 01:55 | 20.5 | 52.6
2020-10-14 02:00 | 20.5 | 52.6
2020-10-14 02:05 | 20.5 | 52.6
2020-10-14 02:10 | 20.5 | 52.5
2020-10-14 02:15 | 20.5 | 52.4
2020-10-14 02:20 | 20.5 | 52.3
2020-10-14 02:25 | 20.5 | 52.2
2020-10-14 02:30 | 20.6 | 52.1
2020-10-14 02:35 | 20.5 | 52.1
2020-10-14 02:40 | 20.6 | 52
2020-10-14 02:45 | 20.6 | 52
2020-10-14 02:50 | 20.6 | 51.9
2020-10-14 02:55 | 20.6 | 51.9
2020-10-14 03:00 | 20.6 | 51.9
2020-10-14 03:05 | 20.5 | 51.9
2020-10-14 03:10 | 20.5 | 51.8
2020-10-14 03:15 | 20.6 | 51.7
2020-10-14 03:20 | 20.6 | 51.6
2020-10-14 03:25 | 20.5 | 51.6
2020-10-14 03:30 | 20.5 | 51.6
2020-10-14 03:35 | 20.6 | 51.5
2020-10-14 03:40 | 20.5 | 51.5
2020-10-14 03:45 | 20.5 | 51.5
2020-10-14 03:50 | 20.6 | 51.5
2020-10-14 03:55 | 20.6 | 51.4
2020-10-14 04:00 | 20.5 | 51.4
2020-10-14 04:05 | 20.6 | 51.3
2020-10-14 04:10 | 20.5 | 51.3
2020-10-14 04:15 | 20.5 | 51.3
2020-10-14 04:20 | 20.5 | 51.3
2020-10-14 04:25 | 20.5 | 51.1
2020-10-14 04:30 | 20.5 | 51.1
2020-10-14 04:35 | 20.5 | 51.1
2020-10-14 04:40 | 20.5 | 50.9
2020-10-14 04:45 | 20.6 | 50.8
2020-10-14 04:50 | 20.6 | 50.7
2020-10-14 04:55 | 20.6 | 50.6
2020-10-14 05:00 | 20.6 | 50.6
2020-10-14 05:05 | 20.5 | 50.6
2020-10-14 05:10 | 20.5 | 50.5
2020-10-14 05:15 | 20.6 | 50.2
2020-10-14 05:20 | 20.6 | 50.1
2020-10-14 05:25 | 20.6 | 49.9
2020-10-14 05:30 | 20.5 | 49.9
2020-10-14 05:35 | 20.5 | 49.7
2020-10-14 05:40 | 20.6 | 49.5
2020-10-14 05:45 | 20.6 | 49.5
2020-10-14 05:50 | 20.6 | 49.4
2020-10-14 05:55 | 20.5 | 49.4
2020-10-14 06:00 | 20.5 | 49.4
2020-10-14 06:05 | 20.5 | 49.4
2020-10-14 06:10 | 20.5 | 49.3
2020-10-14 06:15 | 20.6 | 49.3
2020-10-14 06:20 | 20.6 | 49.3
2020-10-14 06:25 | 20.6 | 49.3
2020-10-14 06:30 | 20.6 | 49.2
2020-10-14 06:35 | 20.6 | 49
2020-10-14 06:40 | 20.5 | 49
2020-10-14 06:45 | 20.6 | 48.8
2020-10-14 06:50 | 20.6 | 48.8
2020-10-14 06:55 | 20.6 | 48.8
2020-10-14 07:00 | 20.6 | 48.8
2020-10-14 07:05 | 20.6 | 48.8
2020-10-14 07:10 | 20.6 | 48.8
2020-10-14 07:15 | 20.6 | 48.7
2020-10-14 07:20 | 20.6 | 48.7
2020-10-14 07:25 | 20.5 | 48.7
2020-10-14 07:30 | 20.6 | 48.6
2020-10-14 07:35 | 20.5 | 48.6
2020-10-14 07:40 | 20.5 | 48.6
2020-10-14 07:45 | 20.6 | 48.6
2020-10-14 07:50 | 20.6 | 48.7
2020-10-14 07:55 | 20.5 | 48.8
2020-10-14 08:00 | 20.5 | 48.8
2020-10-14 08:05 | 20.5 | 48.8
2020-10-14 08:10 | 20.5 | 48.6
2020-10-14 08:15 | 20.6 | 48.5
2020-10-14 08:20 | 20.6 | 48.4
2020-10-14 08:25 | 20.5 | 48.4
2020-10-14 08:30 | 20.6 | 48.4
2020-10-14 08:35 | 20.5 | 48.6
2020-10-14 08:40 | 20.5 | 48.7
2020-10-14 08:45 | 20.5 | 48.7
2020-10-14 08:50 | 20.5 | 48.7
2020-10-14 08:55 | 20.5 | 48.7
2020-10-14 09:00 | 20.5 | 48.7
2020-10-14 09:05 | 20.5 | 48.6
2020-10-14 09:10 | 20.5 | 48.7
2020-10-14 09:15 | 20.5 | 48.7
2020-10-14 09:20 | 20.5 | 48.7
2020-10-14 09:25 | 20.5 | 48.7
2020-10-14 09:30 | 20.5 | 48.7
2020-10-14 09:35 | 20.5 | 48.8
2020-10-14 09:40 | 20.5 | 48.8
2020-10-14 09:45 | 20.5 | 48.9
2020-10-14 09:50 | 20.5 | 49
2020-10-14 09:55 | 20.5 | 48.9
2020-10-14 10:00 | 20.5 | 48.8
2020-10-14 10:05 | 20.5 | 48.8
2020-10-14 10:10 | 20.5 | 48.8
2020-10-14 10:15 | 20.5 | 48.7
2020-10-14 10:20 | 20.5 | 48.6
2020-10-14 10:25 | 20.5 | 48.6
2020-10-14 10:30 | 20.5 | 48.6
2020-10-14 10:35 | 20.5 | 48.5
2020-10-14 10:40 | 20.5 | 48.4
2020-10-14 10:45 | 20.4 | 48.4
2020-10-14 10:50 | 20.4 | 48.5
2020-10-14 10:55 | 20.4 | 48.5
2020-10-14 11:00 | 20.4 | 48.6
2020-10-14 11:05 | 20.4 | 48.7
2020-10-14 11:10 | 20.4 | 48.8
2020-10-14 11:15 | 20.5 | 48.8
2020-10-14 11:20 | 20.4 | 48.9
2020-10-14 11:25 | 20.5 | 49
2020-10-14 11:30 | 20.5 | 49.3
2020-10-14 11:35 | 20.5 | 49.3
2020-10-14 11:40 | 20.4 | 49.5
2020-10-14 11:45 | 20.4 | 49.8
2020-10-14 11:50 | 20.5 | 50.1
2020-10-14 11:55 | 20.4 | 50.3
2020-10-14 12:00 | 20.4 | 50.5
2020-10-14 12:05 | 20.4 | 50.7
2020-10-14 12:10 | 20.4 | 50.8
2020-10-14 12:15 | 20.4 | 50.9
2020-10-14 12:20 | 20.5 | 50.9
2020-10-14 12:25 | 20.4 | 51
2020-10-14 12:30 | 20.5 | 51
2020-10-14 12:35 | 20.5 | 51.1
2020-10-14 12:40 | 20.5 | 51.2
2020-10-14 12:45 | 20.5 | 51.7
2020-10-14 12:50 | 20.5 | 51.8
2020-10-14 12:55 | 20.5 | 52
2020-10-14 13:00 | 20.4 | 52.2
2020-10-14 13:05 | 20.4 | 52.3
2020-10-14 13:10 | 20.4 | 52.3
2020-10-14 13:15 | 20.4 | 52.4
2020-10-14 13:20 | 20.4 | 52.5
2020-10-14 13:25 | 20.5 | 52.5
2020-10-14 13:30 | 20.4 | 52.7
2020-10-14 13:35 | 20.4 | 52.8
2020-10-14 13:40 | 20.4 | 52.8
2020-10-14 13:45 | 20.4 | 52.7
2020-10-14 13:50 | 20.5 | 52.7
2020-10-14 13:55 | 20.4 | 52.8
2020-10-14 14:00 | 20.4 | 52.9
2020-10-14 14:05 | 20.4 | 52.8
2020-10-14 14:10 | 20.4 | 52.9
2020-10-14 14:15 | 20.4 | 52.8
2020-10-14 14:20 | 20.5 | 52.7
2020-10-14 14:25 | 20.4 | 52.6
2020-10-14 14:30 | 20.4 | 52.5
2020-10-14 14:35 | 20.5 | 52.4
2020-10-14 14:40 | 20.5 | 52.3
2020-10-14 14:45 | 20.5 | 52.2
2020-10-14 14:50 | 20.4 | 52.2
2020-10-14 14:55 | 20.5 | 52.1
2020-10-14 15:00 | 20.5 | 52
2020-10-14 15:05 | 20.5 | 52
2020-10-14 15:10 | 20.5 | 51.9
2020-10-14 15:15 | 20.4 | 52
2020-10-14 15:20 | 20.5 | 51.9
2020-10-14 15:25 | 20.4 | 51.9
2020-10-14 15:30 | 20.5 | 51.9
2020-10-14 15:35 | 20.5 | 51.8
2020-10-14 15:40 | 20.5 | 51.7
2020-10-14 15:45 | 20.4 | 51.6
2020-10-14 15:50 | 20.4 | 51.5
2020-10-14 15:55 | 20.5 | 51.4
2020-10-14 16:00 | 20.5 | 51.4
2020-10-14 16:05 | 20.4 | 51.4
2020-10-14 16:10 | 20.5 | 51.3
2020-10-14 16:15 | 20.5 | 51.2
2020-10-14 16:20 | 20.4 | 51.3
2020-10-14 16:25 | 20.5 | 51.3
2020-10-14 16:30 | 20.5 | 51.2
2020-10-14 16:35 | 20.5 | 51.2
2020-10-14 16:40 | 20.5 | 51
2020-10-14 16:45 | 20.5 | 50.9
2020-10-14 16:50 | 20.5 | 51
2020-10-14 16:55 | 20.5 | 51
2020-10-14 17:00 | 20.5 | 51.1
2020-10-14 17:05 | 20.5 | 51
2020-10-14 17:10 | 20.5 | 51
2020-10-14 17:15 | 20.5 | 51
2020-10-14 17:20 | 20.5 | 50.9
2020-10-14 17:25 | 20.5 | 50.9
2020-10-14 17:30 | 20.5 | 50.9
2020-10-14 17:35 | 20.5 | 50.8
2020-10-14 17:40 | 20.5 | 50.7
2020-10-14 17:45 | 20.5 | 50.7
2020-10-14 17:50 | 20.5 | 50.7
2020-10-14 17:55 | 20.5 | 50.6
2020-10-14 18:00 | 20.5 | 50.5
2020-10-14 18:05 | 20.5 | 50.5
2020-10-14 18:10 | 20.5 | 50.5
2020-10-14 18:15 | 20.5 | 50.5
2020-10-14 18:20 | 20.5 | 50.4
2020-10-14 18:25 | 20.5 | 50.4
2020-10-14 18:30 | 20.5 | 50.4
2020-10-14 18:35 | 20.5 | 50.4
2020-10-14 18:40 | 20.5 | 50.4
2020-10-14 18:45 | 20.5 | 50.4
2020-10-14 18:50 | 20.5 | 50.3
2020-10-14 18:55 | 20.5 | 50.2
2020-10-14 19:00 | 20.5 | 50
2020-10-14 19:05 | 20.5 | 50
2020-10-14 19:10 | 20.5 | 49.9
2020-10-14 19:15 | 20.5 | 50
2020-10-14 19:20 | 20.5 | 49.9
2020-10-14 19:25 | 20.5 | 49.9
2020-10-14 19:30 | 20.5 | 49.9
2020-10-14 19:35 | 20.5 | 49.9
2020-10-14 19:40 | 20.5 | 49.9
2020-10-14 19:45 | 20.5 | 49.8
2020-10-14 19:50 | 20.6 | 49.7
2020-10-14 19:55 | 20.5 | 49.7
2020-10-14 20:00 | 20.5 | 49.8
2020-10-14 20:05 | 20.5 | 49.7
2020-10-14 20:10 | 20.5 | 49.6
2020-10-14 20:15 | 20.5 | 49.5
2020-10-14 20:20 | 20.5 | 49.5
2020-10-14 20:25 | 20.5 | 49.5
2020-10-14 20:30 | 20.5 | 49.5
2020-10-14 20:35 | 20.5 | 49.4
2020-10-14 20:40 | 20.5 | 49.3
2020-10-14 20:45 | 20.5 | 49.4
2020-10-14 20:50 | 20.5 | 49.3
2020-10-14 20:55 | 20.5 | 49.3
2020-10-14 21:00 | 20.5 | 49.2
2020-10-14 21:05 | 20.5 | 49.1
2020-10-14 21:10 | 20.5 | 49.1
2020-10-14 21:15 | 20.5 | 49.2
2020-10-14 21:20 | 20.5 | 49
2020-10-14 21:25 | 20.5 | 48.9
2020-10-14 21:30 | 20.5 | 48.9
2020-10-14 21:35 | 20.5 | 48.9
2020-10-14 21:40 | 20.5 | 48.9
2020-10-14 21:45 | 20.5 | 48.8
2020-10-14 21:50 | 20.5 | 48.7
2020-10-14 21:55 | 20.5 | 48.7
2020-10-14 22:00 | 20.5 | 48.7
2020-10-14 22:05 | 20.5 | 48.6
2020-10-14 22:10 | 20.5 | 48.6
2020-10-14 22:15 | 20.5 | 48.6
2020-10-14 22:20 | 20.5 | 48.5
2020-10-14 22:25 | 20.5 | 48.5
2020-10-14 22:30 | 20.5 | 48.6
2020-10-14 22:35 | 20.5 | 48.5
2020-10-14 22:40 | 20.5 | 48.4
2020-10-14 22:45 | 20.5 | 48.4
2020-10-14 22:50 | 20.5 | 48.3
2020-10-14 22:55 | 20.5 | 48.3
2020-10-14 23:00 | 20.5 | 48.2
2020-10-14 23:05 | 20.5 | 48.3
2020-10-14 23:10 | 20.5 | 48.2
2020-10-14 23:15 | 20.5 | 48.2
2020-10-14 23:20 | 20.5 | 48.1
2020-10-14 23:25 | 20.6 | 48.1
2020-10-14 23:30 | 20.6 | 48.1
2020-10-14 23:35 | 20.5 | 48.1
2020-10-14 23:40 | 20.5 | 48
2020-10-14 23:45 | 20.5 | 48.1
2020-10-14 23:50 | 20.5 | 48.1
2020-10-14 23:55 | 20.5 | 48
2020-10-15 | 20.5 | 48
2020-10-15 00:05 | 20.5 | 47.9
2020-10-15 00:10 | 20.5 | 47.8
2020-10-15 00:15 | 20.5 | 47.9
2020-10-15 00:20 | 20.5 | 47.7
2020-10-15 00:25 | 20.5 | 47.7
2020-10-15 00:30 | 20.5 | 47.6
2020-10-15 00:35 | 20.5 | 47.6
2020-10-15 00:40 | 20.5 | 47.5
2020-10-15 00:45 | 20.5 | 47.5
2020-10-15 00:50 | 20.5 | 47.5
2020-10-15 00:55 | 20.5 | 47.4
2020-10-15 01:00 | 20.5 | 47.3
2020-10-15 01:05 | 20.5 | 47.3
2020-10-15 01:10 | 20.5 | 47.3
2020-10-15 01:15 | 20.5 | 47.3
2020-10-15 01:20 | 20.5 | 47.2
2020-10-15 01:25 | 20.5 | 47.3
2020-10-15 01:30 | 20.5 | 47.2
2020-10-15 01:35 | 20.5 | 47.1
2020-10-15 01:40 | 20.5 | 47.2
2020-10-15 01:45 | 20.5 | 47.2
2020-10-15 01:50 | 20.5 | 47.1
2020-10-15 01:55 | 20.5 | 47.1
2020-10-15 02:00 | 20.5 | 47.2
2020-10-15 02:05 | 20.5 | 47
2020-10-15 02:10 | 20.5 | 47
2020-10-15 02:15 | 20.5 | 47
2020-10-15 02:20 | 20.5 | 47
2020-10-15 02:25 | 20.5 | 47.1
2020-10-15 02:30 | 20.5 | 47
2020-10-15 02:35 | 20.5 | 46.9
2020-10-15 02:40 | 20.5 | 47
2020-10-15 02:45 | 20.5 | 47
2020-10-15 02:50 | 20.5 | 46.9
2020-10-15 02:55 | 20.5 | 46.9
2020-10-15 03:00 | 20.5 | 46.8
2020-10-15 03:05 | 20.5 | 46.8
2020-10-15 03:10 | 20.5 | 46.8
2020-10-15 03:15 | 20.5 | 46.8
2020-10-15 03:20 | 20.5 | 46.8
2020-10-15 03:25 | 20.5 | 46.7
2020-10-15 03:30 | 20.5 | 46.6
2020-10-15 03:35 | 20.5 | 46.7
2020-10-15 03:40 | 20.6 | 46.6
2020-10-15 03:45 | 20.5 | 46.6
2020-10-15 03:50 | 20.5 | 46.4
2020-10-15 03:55 | 20.5 | 46.4
2020-10-15 04:00 | 20.5 | 46.5
2020-10-15 04:05 | 20.5 | 46.4
2020-10-15 04:10 | 20.5 | 46.4
2020-10-15 04:15 | 20.5 | 46.4
2020-10-15 04:20 | 20.5 | 46.4
2020-10-15 04:25 | 20.5 | 46.4
2020-10-15 04:30 | 20.5 | 46.3
2020-10-15 04:35 | 20.5 | 46.3
2020-10-15 04:40 | 20.5 | 46.3
2020-10-15 04:45 | 20.5 | 46.2
2020-10-15 04:50 | 20.6 | 46.2
2020-10-15 04:55 | 20.5 | 46.1
2020-10-15 05:00 | 20.5 | 46.1
2020-10-15 05:05 | 20.5 | 46
2020-10-15 05:10 | 20.5 | 46
2020-10-15 05:15 | 20.6 | 45.9
2020-10-15 05:20 | 20.5 | 45.9
2020-10-15 05:25 | 20.5 | 45.9
2020-10-15 05:30 | 20.5 | 46
2020-10-15 05:35 | 20.5 | 45.9
2020-10-15 05:40 | 20.6 | 45.8
2020-10-15 05:45 | 20.5 | 45.8
2020-10-15 05:50 | 20.6 | 45.7
2020-10-15 05:55 | 20.5 | 45.6
2020-10-15 06:00 | 20.5 | 45.6
2020-10-15 06:05 | 20.6 | 45.6
2020-10-15 06:10 | 20.5 | 45.6
2020-10-15 06:15 | 20.6 | 45.6
2020-10-15 06:20 | 20.6 | 45.5
2020-10-15 06:25 | 20.5 | 45.5
2020-10-15 06:30 | 20.6 | 45.5
2020-10-15 06:35 | 20.5 | 45.5
2020-10-15 06:40 | 20.6 | 45.4
2020-10-15 06:45 | 20.6 | 45.4
2020-10-15 06:50 | 20.5 | 45.4
2020-10-15 06:55 | 20.6 | 45.4
2020-10-15 07:00 | 20.5 | 45.3
2020-10-15 07:05 | 20.5 | 45.3
2020-10-15 07:10 | 20.5 | 45.4
2020-10-15 07:15 | 20.6 | 45.2
2020-10-15 07:20 | 20.6 | 45.2
2020-10-15 07:25 | 20.5 | 45.2
2020-10-15 07:30 | 20.6 | 45.2
2020-10-15 07:35 | 20.5 | 45.2
2020-10-15 07:40 | 20.5 | 45.1
2020-10-15 07:45 | 20.5 | 45.1
2020-10-15 07:50 | 20.5 | 45
2020-10-15 07:55 | 20.6 | 44.9
2020-10-15 08:00 | 20.5 | 44.9
2020-10-15 08:05 | 20.6 | 44.9
2020-10-15 08:10 | 20.5 | 44.9
2020-10-15 08:15 | 20.5 | 44.9
2020-10-15 08:20 | 20.5 | 44.9
2020-10-15 08:25 | 20.6 | 44.8
2020-10-15 08:30 | 20.6 | 44.8
2020-10-15 08:35 | 20.6 | 44.7
2020-10-15 08:40 | 20.5 | 44.8
2020-10-15 08:45 | 20.5 | 44.8
2020-10-15 08:50 | 20.6 | 44.8
2020-10-15 08:55 | 20.6 | 44.9
2020-10-15 09:00 | 20.6 | 44.8
2020-10-15 09:05 | 20.6 | 44.9
2020-10-15 09:10 | 20.6 | 44.9
2020-10-15 09:15 | 20.6 | 45
2020-10-15 09:20 | 20.6 | 45.3
2020-10-15 09:25 | 20.6 | 45.5
2020-10-15 09:30 | 20.6 | 45.6
2020-10-15 09:35 | 20.6 | 45.6
2020-10-15 09:40 | 20.6 | 45.4
2020-10-15 09:45 | 20.5 | 45.5
2020-10-15 09:50 | 20.6 | 45.3
2020-10-15 09:55 | 20.6 | 45.2
2020-10-15 10:00 | 20.5 | 45.2
2020-10-15 10:05 | 20.5 | 45.1
2020-10-15 10:10 | 20.6 | 45
2020-10-15 10:15 | 20.6 | 45
2020-10-15 10:20 | 20.6 | 44.9
2020-10-15 10:25 | 20.6 | 44.8
2020-10-15 10:30 | 20.5 | 44.8
2020-10-15 10:35 | 20.5 | 44.8
2020-10-15 10:40 | 20.6 | 44.7
2020-10-15 10:45 | 20.6 | 44.7
2020-10-15 10:50 | 20.5 | 44.7
2020-10-15 10:55 | 20.6 | 44.5
2020-10-15 11:00 | 20.6 | 44.6
2020-10-15 11:05 | 20.6 | 44.7
2020-10-15 11:10 | 20.5 | 44.8
2020-10-15 11:15 | 20.5 | 44.8
2020-10-15 11:20 | 20.5 | 44.8
2020-10-15 11:25 | 20.6 | 44.8
2020-10-15 11:30 | 20.6 | 44.8
2020-10-15 11:35 | 20.5 | 44.7
2020-10-15 11:40 | 20.5 | 44.7
2020-10-15 11:45 | 20.6 | 44.6
2020-10-15 11:50 | 20.5 | 44.5
2020-10-15 11:55 | 20.5 | 44.5
2020-10-15 12:00 | 20.5 | 44.5
2020-10-15 12:05 | 20.5 | 44.3
2020-10-15 12:10 | 20.6 | 44.2
2020-10-15 12:15 | 20.6 | 44.1
2020-10-15 12:20 | 20.5 | 44
2020-10-15 12:25 | 20.5 | 44
2020-10-15 12:30 | 20.5 | 44
2020-10-15 12:35 | 20.6 | 43.9
2020-10-15 12:40 | 20.5 | 44
2020-10-15 12:45 | 20.5 | 44.1
2020-10-15 12:50 | 20.5 | 44
2020-10-15 12:55 | 20.5 | 44
2020-10-15 13:00 | 20.5 | 43.9
2020-10-15 13:05 | 20.5 | 43.8
2020-10-15 13:10 | 20.6 | 43.7
2020-10-15 13:15 | 20.5 | 43.7
2020-10-15 13:20 | 20.5 | 43.7
2020-10-15 13:25 | 20.5 | 43.7
2020-10-15 13:30 | 20.5 | 43.8
2020-10-15 13:35 | 20.6 | 43.7
2020-10-15 13:40 | 20.5 | 43.8
2020-10-15 13:45 | 20.6 | 44
2020-10-15 13:50 | 20.5 | 44.2
2020-10-15 13:55 | 20.5 | 44.4
2020-10-15 14:00 | 20.5 | 44.6
2020-10-15 14:05 | 20.5 | 44.5
2020-10-15 14:10 | 20.5 | 44.5
2020-10-15 14:15 | 20.5 | 44.6
2020-10-15 14:20 | 20.5 | 44.7
2020-10-15 14:25 | 20.5 | 44.8
2020-10-15 14:30 | 20.5 | 44.7
2020-10-15 14:35 | 20.5 | 44.7
2020-10-15 14:40 | 20.5 | 44.7
2020-10-15 14:45 | 20.5 | 44.6
2020-10-15 14:50 | 20.5 | 44.4
2020-10-15 14:55 | 20.5 | 44.4
2020-10-15 15:00 | 20.5 | 44.6
2020-10-15 15:05 | 20.5 | 44.6
2020-10-15 15:10 | 20.5 | 44.5
2020-10-15 15:15 | 20.5 | 44.4
2020-10-15 15:20 | 20.5 | 44.4
2020-10-15 15:25 | 20.5 | 44.4
2020-10-15 15:30 | 20.5 | 44.3
2020-10-15 15:35 | 20.5 | 44.2
2020-10-15 15:40 | 20.5 | 44.2
2020-10-15 15:45 | 20.5 | 44.3
2020-10-15 15:50 | 20.5 | 44.2
2020-10-15 15:55 | 20.5 | 44
2020-10-15 16:00 | 20.5 | 44
2020-10-15 16:05 | 20.5 | 44.1
2020-10-15 16:10 | 20.5 | 44.2
2020-10-15 16:15 | 20.5 | 44.2
2020-10-15 16:20 | 20.5 | 44.2
2020-10-15 16:25 | 20.5 | 44
2020-10-15 16:30 | 20.5 | 44
2020-10-15 16:35 | 20.5 | 44
2020-10-15 16:40 | 20.5 | 44
2020-10-15 16:45 | 20.5 | 43.9
2020-10-15 16:50 | 20.5 | 44.1
2020-10-15 16:55 | 20.5 | 44.1
2020-10-15 17:00 | 20.5 | 44.3
2020-10-15 17:05 | 20.5 | 44.4
2020-10-15 17:10 | 20.5 | 44.5
2020-10-15 17:15 | 20.5 | 44.4
2020-10-15 17:20 | 20.5 | 44.4
2020-10-15 17:25 | 20.5 | 44.5
2020-10-15 17:30 | 20.5 | 44.6
2020-10-15 17:35 | 20.4 | 44.5
2020-10-15 17:40 | 20.5 | 44.4
2020-10-15 17:45 | 20.5 | 44.2
2020-10-15 17:50 | 20.5 | 44.2
2020-10-15 17:55 | 20.5 | 44.1
2020-10-15 18:00 | 20.5 | 44.1
2020-10-15 18:05 | 20.5 | 44
2020-10-15 18:10 | 20.5 | 44
2020-10-15 18:15 | 20.5 | 43.8
2020-10-15 18:20 | 20.5 | 43.8
2020-10-15 18:25 | 20.5 | 43.7
2020-10-15 18:30 | 20.5 | 43.6
2020-10-15 18:35 | 20.5 | 43.5
2020-10-15 18:40 | 20.5 | 43.4
2020-10-15 18:45 | 20.5 | 43.3
2020-10-15 18:50 | 20.5 | 43.3
2020-10-15 18:55 | 20.5 | 43.3
2020-10-15 19:00 | 20.5 | 43.2
2020-10-15 19:05 | 20.5 | 43.2
2020-10-15 19:10 | 20.5 | 43.1
2020-10-15 19:15 | 20.5 | 42.9
2020-10-15 19:20 | 20.5 | 42.8
2020-10-15 19:25 | 20.5 | 42.8
2020-10-15 19:30 | 20.5 | 42.8
2020-10-15 19:35 | 20.5 | 42.8
2020-10-15 19:40 | 20.5 | 42.7
2020-10-15 19:45 | 20.5 | 42.7
2020-10-15 19:50 | 20.5 | 42.7
2020-10-15 19:55 | 20.5 | 42.7
2020-10-15 20:00 | 20.5 | 42.6
2020-10-15 20:05 | 20.5 | 42.6
2020-10-15 20:10 | 20.5 | 42.7
2020-10-15 20:15 | 20.5 | 42.6
2020-10-15 20:20 | 20.5 | 42.7
2020-10-15 20:25 | 20.6 | 42.6
2020-10-15 20:30 | 20.5 | 42.7
2020-10-15 20:35 | 20.5 | 42.6
2020-10-15 20:40 | 20.5 | 42.6
2020-10-15 20:45 | 20.5 | 42.6
2020-10-15 20:50 | 20.5 | 42.6
2020-10-15 20:55 | 20.5 | 42.6
2020-10-15 21:00 | 20.5 | 42.6
2020-10-15 21:05 | 20.5 | 42.5
2020-10-15 21:10 | 20.5 | 42.5
2020-10-15 21:15 | 20.5 | 42.5
2020-10-15 21:20 | 20.5 | 42.5
2020-10-15 21:25 | 20.5 | 42.4
2020-10-15 21:30 | 20.5 | 42.4
2020-10-15 21:35 | 20.5 | 42.4
2020-10-15 21:40 | 20.5 | 42.3
2020-10-15 21:45 | 20.5 | 42.3
2020-10-15 21:50 | 20.5 | 42.3
2020-10-15 21:55 | 20.5 | 42.2
2020-10-15 22:00 | 20.5 | 42.2
2020-10-15 22:05 | 20.5 | 42.2
2020-10-15 22:10 | 20.6 | 42.1
2020-10-15 22:15 | 20.6 | 42
2020-10-15 22:20 | 20.5 | 42
2020-10-15 22:25 | 20.5 | 42
2020-10-15 22:30 | 20.5 | 42
2020-10-15 22:35 | 20.5 | 42
2020-10-15 22:40 | 20.5 | 42
2020-10-15 22:45 | 20.5 | 42
2020-10-15 22:50 | 20.5 | 42
2020-10-15 22:55 | 20.6 | 41.9
2020-10-15 23:00 | 20.6 | 42
2020-10-15 23:05 | 20.6 | 41.9
2020-10-15 23:10 | 20.6 | 41.9
2020-10-15 23:15 | 20.5 | 41.9
2020-10-15 23:20 | 20.6 | 41.9
2020-10-15 23:25 | 20.6 | 41.8
2020-10-15 23:30 | 20.6 | 41.7
2020-10-15 23:35 | 20.5 | 41.8
2020-10-15 23:40 | 20.5 | 41.7
2020-10-15 23:45 | 20.5 | 41.7
2020-10-15 23:50 | 20.5 | 41.7
2020-10-15 23:55 | 20.5 | 41.6
2020-10-16 | 20.5 | 41.6
2020-10-16 00:05 | 20.5 | 41.5
2020-10-16 00:10 | 20.6 | 41.4
2020-10-16 00:15 | 20.6 | 41.4
2020-10-16 00:20 | 20.6 | 41.3
2020-10-16 00:25 | 20.6 | 41.3
2020-10-16 00:30 | 20.6 | 41.3
2020-10-16 00:35 | 20.5 | 41.3
2020-10-16 00:40 | 20.5 | 41.2
2020-10-16 00:45 | 20.6 | 41.1
2020-10-16 00:50 | 20.6 | 41.1
2020-10-16 00:55 | 20.6 | 41
2020-10-16 01:00 | 20.6 | 41
2020-10-16 01:05 | 20.6 | 40.9
2020-10-16 01:10 | 20.6 | 40.9
2020-10-16 01:15 | 20.6 | 40.9
2020-10-16 01:20 | 20.6 | 40.9
2020-10-16 01:25 | 20.6 | 40.9
2020-10-16 01:30 | 20.6 | 40.9
2020-10-16 01:35 | 20.6 | 40.9
2020-10-16 01:40 | 20.6 | 40.8
2020-10-16 01:45 | 20.6 | 40.7
2020-10-16 01:50 | 20.6 | 40.6
2020-10-16 01:55 | 20.6 | 40.6
2020-10-16 02:00 | 20.6 | 40.6
2020-10-16 02:05 | 20.6 | 40.5
2020-10-16 02:10 | 20.6 | 40.5
2020-10-16 02:15 | 20.6 | 40.4
2020-10-16 02:20 | 20.6 | 40.4
2020-10-16 02:25 | 20.6 | 40.4
2020-10-16 02:30 | 20.6 | 40.4
2020-10-16 02:35 | 20.6 | 40.2
2020-10-16 02:40 | 20.6 | 40.2
2020-10-16 02:45 | 20.6 | 40.2
2020-10-16 02:50 | 20.6 | 40.1
2020-10-16 02:55 | 20.6 | 40.1
2020-10-16 03:00 | 20.6 | 40.1
2020-10-16 03:05 | 20.6 | 40
2020-10-16 03:10 | 20.6 | 39.9
2020-10-16 03:15 | 20.6 | 39.9
2020-10-16 03:20 | 20.6 | 39.8
2020-10-16 03:25 | 20.6 | 39.7
2020-10-16 03:30 | 20.6 | 39.7
2020-10-16 03:35 | 20.6 | 39.6
2020-10-16 03:40 | 20.6 | 39.5
2020-10-16 03:45 | 20.6 | 39.5
2020-10-16 03:50 | 20.6 | 39.4
2020-10-16 03:55 | 20.6 | 39.5
2020-10-16 04:00 | 20.6 | 39.4
2020-10-16 04:05 | 20.6 | 39.4
2020-10-16 04:10 | 20.6 | 39.3
2020-10-16 04:15 | 20.6 | 39.4
2020-10-16 04:20 | 20.6 | 39.3
2020-10-16 04:25 | 20.6 | 39.3
2020-10-16 04:30 | 20.6 | 39.3
2020-10-16 04:35 | 20.6 | 39.2
2020-10-16 04:40 | 20.6 | 39.2
2020-10-16 04:45 | 20.6 | 39.2
2020-10-16 04:50 | 20.6 | 39.1
2020-10-16 04:55 | 20.6 | 39.1
2020-10-16 05:00 | 20.6 | 39
2020-10-16 05:05 | 20.6 | 39
2020-10-16 05:10 | 20.6 | 38.9
2020-10-16 05:15 | 20.6 | 38.9
2020-10-16 05:20 | 20.6 | 38.8
2020-10-16 05:25 | 20.6 | 38.8
2020-10-16 05:30 | 20.6 | 38.8
2020-10-16 05:35 | 20.6 | 38.7
2020-10-16 05:40 | 20.6 | 38.7
2020-10-16 05:45 | 20.6 | 38.7
2020-10-16 05:50 | 20.6 | 38.7
2020-10-16 05:55 | 20.6 | 38.6
2020-10-16 06:00 | 20.6 | 38.6
2020-10-16 06:05 | 20.6 | 38.5
2020-10-16 06:10 | 20.6 | 38.5
2020-10-16 06:15 | 20.6 | 38.5
2020-10-16 06:20 | 20.6 | 38.4
2020-10-16 06:25 | 20.6 | 38.4
2020-10-16 06:30 | 20.6 | 38.4
2020-10-16 06:35 | 20.6 | 38.4
2020-10-16 06:40 | 20.6 | 38.4
2020-10-16 06:45 | 20.6 | 38.3
2020-10-16 06:50 | 20.6 | 38.3
2020-10-16 06:55 | 20.6 | 38.2
2020-10-16 07:00 | 20.6 | 38.1
2020-10-16 07:05 | 20.6 | 38.2
2020-10-16 07:10 | 20.6 | 38.1
2020-10-16 07:15 | 20.6 | 38
2020-10-16 07:20 | 20.6 | 38
2020-10-16 07:25 | 20.6 | 38
2020-10-16 07:30 | 20.6 | 38
2020-10-16 07:35 | 20.7 | 38
2020-10-16 07:40 | 20.6 | 37.9
2020-10-16 07:45 | 20.6 | 37.9
2020-10-16 07:50 | 20.6 | 37.9
2020-10-16 07:55 | 20.6 | 37.9
2020-10-16 08:00 | 20.6 | 38
2020-10-16 08:05 | 20.6 | 37.9
2020-10-16 08:10 | 20.6 | 38
2020-10-16 08:15 | 20.6 | 38
2020-10-16 08:20 | 20.6 | 38
2020-10-16 08:25 | 20.6 | 38.1
2020-10-16 08:30 | 20.6 | 38.3
2020-10-16 08:35 | 20.6 | 38.4
2020-10-16 08:40 | 20.6 | 38.5
2020-10-16 08:45 | 20.5 | 38.7
2020-10-16 08:50 | 20.5 | 38.7
2020-10-16 08:55 | 20.5 | 38.8
2020-10-16 09:00 | 20.5 | 38.9
2020-10-16 09:05 | 20.5 | 39
2020-10-16 09:10 | 20.5 | 39
2020-10-16 09:15 | 20.5 | 39.1
2020-10-16 09:20 | 20.5 | 39.2
2020-10-16 09:25 | 20.5 | 39.3
2020-10-16 09:30 | 20.5 | 39.3
2020-10-16 09:35 | 20.5 | 39.4
2020-10-16 09:40 | 20.5 | 39.4
2020-10-16 09:45 | 20.5 | 39.6
2020-10-16 09:50 | 20.5 | 39.7
2020-10-16 09:55 | 20.5 | 39.7
2020-10-16 10:00 | 20.5 | 39.7
2020-10-16 10:05 | 20.5 | 39.8
2020-10-16 10:10 | 20.6 | 39.6
2020-10-16 10:15 | 20.5 | 39.6
2020-10-16 10:20 | 20.6 | 39.2
2020-10-16 10:25 | 20.6 | 39.1
2020-10-16 10:30 | 20.5 | 39.1
2020-10-16 10:35 | 20.5 | 39.4
2020-10-16 10:40 | 20.4 | 39.5
2020-10-16 10:45 | 20.4 | 39.6
2020-10-16 10:50 | 20.4 | 39.6
2020-10-16 10:55 | 20.4 | 39.9
2020-10-16 11:00 | 20.4 | 39.7
2020-10-16 11:05 | 20.5 | 39.5
2020-10-16 11:10 | 20.5 | 39.5
2020-10-16 11:15 | 20.5 | 39.4
2020-10-16 11:20 | 20.5 | 39.3
2020-10-16 11:25 | 20.5 | 39.4
2020-10-16 11:30 | 20.4 | 39.5
2020-10-16 11:35 | 20.5 | 39.4
2020-10-16 11:40 | 20.5 | 39.6
2020-10-16 11:45 | 20.5 | 39.6
2020-10-16 11:50 | 20.4 | 39.8
2020-10-16 11:55 | 20.5 | 39.8
2020-10-16 12:00 | 20.5 | 40
2020-10-16 12:05 | 20.5 | 40.1
2020-10-16 12:10 | 20.5 | 40.3
2020-10-16 12:15 | 20.5 | 40.5
2020-10-16 12:20 | 20.5 | 40.6
2020-10-16 12:25 | 20.5 | 40.4
2020-10-16 12:30 | 20.4 | 40.4
2020-10-16 12:35 | 20.5 | 40.4
2020-10-16 12:40 | 20.4 | 40.5
2020-10-16 12:45 | 20.4 | 40.5
2020-10-16 12:50 | 20.4 | 40.5
2020-10-16 12:55 | 20.5 | 40.7
2020-10-16 13:00 | 20.4 | 40.9
2020-10-16 13:05 | 20.4 | 40.9
2020-10-16 13:10 | 20.4 | 41.1
2020-10-16 13:15 | 20.3 | 41.1
2020-10-16 13:20 | 20.4 | 41
2020-10-16 13:25 | 20.4 | 41
2020-10-16 13:30 | 20.4 | 40.8
2020-10-16 13:35 | 20.3 | 40.9
2020-10-16 13:40 | 20.3 | 41
2020-10-16 13:45 | 20.3 | 40.9
2020-10-16 13:50 | 20.3 | 40.8
2020-10-16 13:55 | 20.3 | 40.5
2020-10-16 14:00 | 20.3 | 40.3
2020-10-16 14:05 | 20.3 | 40.1
2020-10-16 14:10 | 20.3 | 40.2
2020-10-16 14:15 | 20.4 | 40.1
2020-10-16 14:20 | 20.4 | 39.6
2020-10-16 14:25 | 20.4 | 39.7
2020-10-16 14:30 | 20.3 | 39.7
2020-10-16 14:35 | 20.3 | 39.9
2020-10-16 14:40 | 20.3 | 40.1
2020-10-16 14:45 | 20.4 | 40
2020-10-16 14:50 | 20.3 | 39.9
2020-10-16 14:55 | 20.4 | 39.7
2020-10-16 15:00 | 20.4 | 39.5
2020-10-16 15:05 | 20.4 | 39.6
2020-10-16 15:10 | 20.4 | 39.6
2020-10-16 15:15 | 20.3 | 39.7
2020-10-16 15:20 | 20.4 | 39.5
2020-10-16 15:25 | 20.5 | 39.4
2020-10-16 15:30 | 20.4 | 39.6
2020-10-16 15:35 | 20.4 | 39.6
2020-10-16 15:40 | 20.4 | 39.8
2020-10-16 15:45 | 20.4 | 39.8
2020-10-16 15:50 | 20.4 | 39.9
2020-10-16 15:55 | 20.4 | 40
2020-10-16 16:00 | 20.4 | 40
2020-10-16 16:05 | 20.4 | 40.2
2020-10-16 16:10 | 20.4 | 40.2
2020-10-16 16:15 | 20.4 | 40.2
2020-10-16 16:20 | 20.4 | 40.2
2020-10-16 16:25 | 20.4 | 40.5
2020-10-16 16:30 | 20.4 | 40.5
2020-10-16 16:35 | 20.4 | 40.5
2020-10-16 16:40 | 20.4 | 40.3
2020-10-16 16:45 | 20.4 | 40.1
2020-10-16 16:50 | 20.4 | 40
2020-10-16 16:55 | 20.4 | 39.8
2020-10-16 17:00 | 20.4 | 39.6
2020-10-16 17:05 | 20.4 | 39.6
2020-10-16 17:10 | 20.4 | 39.3
2020-10-16 17:15 | 20.4 | 39.2
2020-10-16 17:20 | 20.4 | 39.2
2020-10-16 17:25 | 20.4 | 39.2
2020-10-16 17:30 | 20.4 | 39
2020-10-16 17:35 | 20.4 | 38.9
2020-10-16 17:40 | 20.4 | 38.9
2020-10-16 17:45 | 20.4 | 38.9
2020-10-16 17:50 | 20.4 | 38.9
2020-10-16 17:55 | 20.4 | 38.8
2020-10-16 18:00 | 20.4 | 38.8
2020-10-16 18:05 | 20.4 | 38.8
2020-10-16 18:10 | 20.4 | 38.9
2020-10-16 18:15 | 20.4 | 38.9
2020-10-16 18:20 | 20.4 | 38.9
2020-10-16 18:25 | 20.4 | 39
2020-10-16 18:30 | 20.4 | 39.1
2020-10-16 18:35 | 20.4 | 39.2
2020-10-16 18:40 | 20.4 | 39.3
2020-10-16 18:45 | 20.4 | 39.2
2020-10-16 18:50 | 20.4 | 39.3
2020-10-16 18:55 | 20.4 | 39.2
2020-10-16 19:00 | 20.4 | 39.2
2020-10-16 19:05 | 20.5 | 39.2
2020-10-16 19:10 | 20.5 | 39.1
2020-10-16 19:15 | 20.4 | 39.1
2020-10-16 19:20 | 20.5 | 39
2020-10-16 19:25 | 20.5 | 39
2020-10-16 19:30 | 20.4 | 38.9
2020-10-16 19:35 | 20.4 | 38.8
2020-10-16 19:40 | 20.5 | 38.6
2020-10-16 19:45 | 20.5 | 38.5
2020-10-16 19:50 | 20.5 | 38.4
2020-10-16 19:55 | 20.5 | 38.3
2020-10-16 20:00 | 20.5 | 38.2
2020-10-16 20:05 | 20.5 | 38.2
2020-10-16 20:10 | 20.5 | 38.1
2020-10-16 20:15 | 20.5 | 38.1
2020-10-16 20:20 | 20.5 | 38.1
2020-10-16 20:25 | 20.5 | 38.1
2020-10-16 20:30 | 20.5 | 38.1
2020-10-16 20:35 | 20.5 | 38
2020-10-16 20:40 | 20.5 | 38.1
2020-10-16 20:45 | 20.5 | 38.1
2020-10-16 20:50 | 20.5 | 38.2
2020-10-16 20:55 | 20.5 | 38.2
2020-10-16 21:00 | 20.5 | 38.2
2020-10-16 21:05 | 20.5 | 38.2
2020-10-16 21:10 | 20.5 | 38.3
2020-10-16 21:15 | 20.5 | 38.4
2020-10-16 21:20 | 20.5 | 38.5
2020-10-16 21:25 | 20.5 | 38.5
2020-10-16 21:30 | 20.5 | 38.7
2020-10-16 21:35 | 20.5 | 38.7
2020-10-16 21:40 | 20.5 | 38.8
2020-10-16 21:45 | 20.5 | 39
2020-10-16 21:50 | 20.5 | 39
2020-10-16 21:55 | 20.5 | 39.1
2020-10-16 22:00 | 20.5 | 39.1
2020-10-16 22:05 | 20.5 | 39.1
2020-10-16 22:10 | 20.5 | 39.1
2020-10-16 22:15 | 20.5 | 39.1
2020-10-16 22:20 | 20.5 | 39.1
2020-10-16 22:25 | 20.5 | 39.1
2020-10-16 22:30 | 20.5 | 39.1
2020-10-16 22:35 | 20.5 | 39.1
2020-10-16 22:40 | 20.5 | 39
2020-10-16 22:45 | 20.5 | 39
2020-10-16 22:50 | 20.5 | 39
2020-10-16 22:55 | 20.5 | 39
2020-10-16 23:00 | 20.5 | 39
2020-10-16 23:05 | 20.5 | 39
2020-10-16 23:10 | 20.5 | 39
2020-10-16 23:15 | 20.5 | 39.2
2020-10-16 23:20 | 20.5 | 39.2
2020-10-16 23:25 | 20.5 | 39.2
2020-10-16 23:30 | 20.5 | 39.2
2020-10-16 23:35 | 20.5 | 39.3
2020-10-16 23:40 | 20.5 | 39.3
2020-10-16 23:45 | 20.5 | 39.4
2020-10-16 23:50 | 20.5 | 39.4
2020-10-16 23:55 | 20.5 | 39.4
2020-10-17 | 20.5 | 39.5
2020-10-17 00:05 | 20.5 | 39.5
2020-10-17 00:10 | 20.5 | 39.7
2020-10-17 00:15 | 20.5 | 39.8
2020-10-17 00:20 | 20.5 | 39.9
2020-10-17 00:25 | 20.5 | 40
2020-10-17 00:30 | 20.5 | 40.1
2020-10-17 00:35 | 20.5 | 40.1
2020-10-17 00:40 | 20.5 | 40.3
2020-10-17 00:45 | 20.5 | 40.4
2020-10-17 00:50 | 20.5 | 40.5
2020-10-17 00:55 | 20.5 | 40.5
2020-10-17 01:00 | 20.5 | 40.6
2020-10-17 01:05 | 20.5 | 40.6
2020-10-17 01:10 | 20.5 | 40.8
2020-10-17 01:15 | 20.5 | 40.8
2020-10-17 01:20 | 20.5 | 40.9
2020-10-17 01:25 | 20.5 | 40.9
2020-10-17 01:30 | 20.5 | 41
2020-10-17 01:35 | 20.5 | 41
2020-10-17 01:40 | 20.5 | 41.1
2020-10-17 01:45 | 20.5 | 41.1
2020-10-17 01:50 | 20.5 | 41.1
2020-10-17 01:55 | 20.5 | 41.1
2020-10-17 02:00 | 20.5 | 41.2
2020-10-17 02:05 | 20.5 | 41.2
2020-10-17 02:10 | 20.5 | 41.2
2020-10-17 02:15 | 20.5 | 41.2
2020-10-17 02:20 | 20.5 | 41.3
2020-10-17 02:25 | 20.5 | 41.3
2020-10-17 02:30 | 20.5 | 41.3
2020-10-17 02:35 | 20.5 | 41.2
2020-10-17 02:40 | 20.5 | 41.2
2020-10-17 02:45 | 20.5 | 41.1
2020-10-17 02:50 | 20.5 | 41.1
2020-10-17 02:55 | 20.5 | 41.2
2020-10-17 03:00 | 20.5 | 41.2
2020-10-17 03:05 | 20.5 | 41.2
2020-10-17 03:10 | 20.5 | 41.3
2020-10-17 03:15 | 20.5 | 41.3
2020-10-17 03:20 | 20.5 | 41.4
2020-10-17 03:25 | 20.5 | 41.5
2020-10-17 03:30 | 20.5 | 41.6
2020-10-17 03:35 | 20.5 | 41.5
2020-10-17 03:40 | 20.5 | 41.6
2020-10-17 03:45 | 20.5 | 41.6
2020-10-17 03:50 | 20.5 | 41.6
2020-10-17 03:55 | 20.5 | 41.7
2020-10-17 04:00 | 20.5 | 41.7
2020-10-17 04:05 | 20.5 | 41.6
2020-10-17 04:10 | 20.5 | 41.7
2020-10-17 04:15 | 20.5 | 41.6
2020-10-17 04:20 | 20.5 | 41.6
2020-10-17 04:25 | 20.5 | 41.7
2020-10-17 04:30 | 20.5 | 41.6
2020-10-17 04:35 | 20.5 | 41.6
2020-10-17 04:40 | 20.4 | 41.7
2020-10-17 04:45 | 20.5 | 41.6
2020-10-17 04:50 | 20.5 | 41.6
2020-10-17 04:55 | 20.5 | 41.6
2020-10-17 05:00 | 20.5 | 41.5
2020-10-17 05:05 | 20.5 | 41.6
2020-10-17 05:10 | 20.5 | 41.6
2020-10-17 05:15 | 20.5 | 41.6
2020-10-17 05:20 | 20.5 | 41.6
2020-10-17 05:25 | 20.5 | 41.6
2020-10-17 05:30 | 20.5 | 41.6
2020-10-17 05:35 | 20.5 | 41.6
2020-10-17 05:40 | 20.5 | 41.7
2020-10-17 05:45 | 20.5 | 41.6
2020-10-17 05:50 | 20.5 | 41.6
2020-10-17 05:55 | 20.5 | 41.7
2020-10-17 06:00 | 20.5 | 41.7
2020-10-17 06:05 | 20.5 | 41.7
2020-10-17 06:10 | 20.5 | 41.8
2020-10-17 06:15 | 20.5 | 41.8
2020-10-17 06:20 | 20.5 | 41.8
2020-10-17 06:25 | 20.5 | 41.9
2020-10-17 06:30 | 20.5 | 41.9
2020-10-17 06:35 | 20.5 | 42
2020-10-17 06:40 | 20.5 | 42.1
2020-10-17 06:45 | 20.5 | 42.1
2020-10-17 06:50 | 20.4 | 42.2
2020-10-17 06:55 | 20.5 | 42.2
2020-10-17 07:00 | 20.5 | 42.2
2020-10-17 07:05 | 20.5 | 42.3
2020-10-17 07:10 | 20.5 | 42.2
2020-10-17 07:15 | 20.5 | 42.2
2020-10-17 07:20 | 20.5 | 42.3
2020-10-17 07:25 | 20.5 | 42.3
2020-10-17 07:30 | 20.5 | 42.3
2020-10-17 07:35 | 20.5 | 42.4
2020-10-17 07:40 | 20.5 | 42.3
2020-10-17 07:45 | 20.5 | 42.3
2020-10-17 07:50 | 20.5 | 42.4
2020-10-17 07:55 | 20.5 | 42.5
2020-10-17 08:00 | 20.5 | 42.5
2020-10-17 08:05 | 20.5 | 42.5
2020-10-17 08:10 | 20.5 | 42.6
2020-10-17 08:15 | 20.5 | 42.7
2020-10-17 08:20 | 20.5 | 42.7
2020-10-17 08:25 | 20.5 | 42.7
2020-10-17 08:30 | 20.5 | 42.7
2020-10-17 08:35 | 20.4 | 42.8
2020-10-17 08:40 | 20.5 | 42.8
2020-10-17 08:45 | 20.5 | 42.8
2020-10-17 08:50 | 20.5 | 42.9
2020-10-17 08:55 | 20.5 | 42.9
2020-10-17 09:00 | 20.4 | 42.9
2020-10-17 09:05 | 20.4 | 43
2020-10-17 09:10 | 20.5 | 43
2020-10-17 09:15 | 20.5 | 43
2020-10-17 09:20 | 20.5 | 43
2020-10-17 09:25 | 20.5 | 43.1
2020-10-17 09:30 | 20.5 | 43.1
2020-10-17 09:35 | 20.5 | 43.2
2020-10-17 09:40 | 20.5 | 43.3
2020-10-17 09:45 | 20.5 | 43.4
2020-10-17 09:50 | 20.5 | 43.5
2020-10-17 09:55 | 20.5 | 43.6
2020-10-17 10:00 | 20.4 | 43.6
2020-10-17 10:05 | 20.4 | 43.7
2020-10-17 10:10 | 20.4 | 43.7
2020-10-17 10:15 | 20.5 | 43.8
2020-10-17 10:20 | 20.5 | 43.8
2020-10-17 10:25 | 20.5 | 43.9
2020-10-17 10:30 | 20.5 | 43.9
2020-10-17 10:35 | 20.5 | 43.9
2020-10-17 10:40 | 20.5 | 43.9
2020-10-17 10:45 | 20.5 | 43.9
2020-10-17 10:50 | 20.4 | 44
2020-10-17 10:55 | 20.5 | 44
2020-10-17 11:00 | 20.5 | 44.1
2020-10-17 11:05 | 20.5 | 44
2020-10-17 11:10 | 20.5 | 44.1
2020-10-17 11:15 | 20.5 | 44.1
2020-10-17 11:20 | 20.4 | 44.1
2020-10-17 11:25 | 20.4 | 44.2
2020-10-17 11:30 | 20.4 | 44.2
2020-10-17 11:35 | 20.5 | 44.2
2020-10-17 11:40 | 20.5 | 44.2
2020-10-17 11:45 | 20.5 | 44.3
2020-10-17 11:50 | 20.5 | 44.3
2020-10-17 11:55 | 20.5 | 44.3
2020-10-17 12:00 | 20.5 | 44.4
2020-10-17 12:05 | 20.5 | 44.4
2020-10-17 12:10 | 20.4 | 44.5
2020-10-17 12:15 | 20.4 | 44.5
2020-10-17 12:20 | 20.4 | 44.5
2020-10-17 12:25 | 20.5 | 44.5
2020-10-17 12:30 | 20.5 | 44.5
2020-10-17 12:35 | 20.4 | 44.6
2020-10-17 12:40 | 20.4 | 44.6
2020-10-17 12:45 | 20.5 | 44.6
2020-10-17 12:50 | 20.5 | 44.5
2020-10-17 12:55 | 20.4 | 44.6
2020-10-17 13:00 | 20.4 | 44.7
2020-10-17 13:05 | 20.4 | 44.7
2020-10-17 13:10 | 20.5 | 44.7
2020-10-17 13:15 | 20.5 | 44.7
2020-10-17 13:20 | 20.5 | 44.6
2020-10-17 13:25 | 20.5 | 44.7
2020-10-17 13:30 | 20.5 | 44.8
2020-10-17 13:35 | 20.5 | 44.8
2020-10-17 13:40 | 20.5 | 44.8
2020-10-17 13:45 | 20.5 | 44.8
2020-10-17 13:50 | 20.5 | 44.9
2020-10-17 13:55 | 20.5 | 44.9
2020-10-17 14:00 | 20.5 | 45
2020-10-17 14:05 | 20.5 | 45
2020-10-17 14:10 | 20.5 | 45
2020-10-17 14:15 | 20.5 | 45
2020-10-17 14:20 | 20.4 | 45.1
2020-10-17 14:25 | 20.4 | 45.1
2020-10-17 14:30 | 20.4 | 45.1
2020-10-17 14:35 | 20.4 | 45.1
2020-10-17 14:40 | 20.5 | 45.1
2020-10-17 14:45 | 20.5 | 45.2
2020-10-17 14:50 | 20.4 | 45.3
2020-10-17 14:55 | 20.5 | 45.2
2020-10-17 15:00 | 20.5 | 45.2
2020-10-17 15:05 | 20.4 | 45.3
2020-10-17 15:10 | 20.4 | 45.3
2020-10-17 15:15 | 20.5 | 45.4
2020-10-17 15:20 | 20.4 | 45.4
2020-10-17 15:25 | 20.4 | 45.4
2020-10-17 15:30 | 20.5 | 45.4
2020-10-17 15:35 | 20.5 | 45.5
2020-10-17 15:40 | 20.5 | 45.5
2020-10-17 15:45 | 20.5 | 45.5
2020-10-17 15:50 | 20.5 | 45.5
2020-10-17 15:55 | 20.4 | 45.7
2020-10-17 16:00 | 20.4 | 45.7
2020-10-17 16:05 | 20.4 | 45.8
2020-10-17 16:10 | 20.4 | 45.9
2020-10-17 16:15 | 20.5 | 45.9
2020-10-17 16:20 | 20.4 | 45.9
2020-10-17 16:25 | 20.4 | 45.9
2020-10-17 16:30 | 20.5 | 46
2020-10-17 16:35 | 20.4 | 46
2020-10-17 16:40 | 20.4 | 46.1
2020-10-17 16:45 | 20.4 | 46
2020-10-17 16:50 | 20.4 | 46.2
2020-10-17 16:55 | 20.4 | 46.1
2020-10-17 17:00 | 20.4 | 46.2
2020-10-17 17:05 | 20.4 | 46.2
2020-10-17 17:10 | 20.4 | 46.3
2020-10-17 17:15 | 20.4 | 46.3
2020-10-17 17:20 | 20.4 | 46.3
2020-10-17 17:25 | 20.4 | 46.3
2020-10-17 17:30 | 20.4 | 46.4
2020-10-17 17:35 | 20.4 | 46.4
2020-10-17 17:40 | 20.4 | 46.5
2020-10-17 17:45 | 20.5 | 46.4
2020-10-17 17:50 | 20.4 | 46.5
2020-10-17 17:55 | 20.4 | 46.6
2020-10-17 18:00 | 20.4 | 46.6
2020-10-17 18:05 | 20.4 | 46.6
2020-10-17 18:10 | 20.4 | 46.6
2020-10-17 18:15 | 20.4 | 46.7
2020-10-17 18:20 | 20.4 | 46.7
2020-10-17 18:25 | 20.4 | 46.7
2020-10-17 18:30 | 20.4 | 46.7
2020-10-17 18:35 | 20.4 | 46.6
2020-10-17 18:40 | 20.4 | 46.6
2020-10-17 18:45 | 20.5 | 46.6
2020-10-17 18:50 | 20.5 | 46.6
2020-10-17 18:55 | 20.4 | 46.7
2020-10-17 19:00 | 20.4 | 46.7
2020-10-17 19:05 | 20.4 | 46.6
2020-10-17 19:10 | 20.4 | 46.6
2020-10-17 19:15 | 20.4 | 46.6
2020-10-17 19:20 | 20.5 | 46.6
2020-10-17 19:25 | 20.5 | 46.5
2020-10-17 19:30 | 20.5 | 46.5
2020-10-17 19:35 | 20.4 | 46.5
2020-10-17 19:40 | 20.4 | 46.6
2020-10-17 19:45 | 20.4 | 46.4
2020-10-17 19:50 | 20.4 | 46.4
2020-10-17 19:55 | 20.4 | 46.5
2020-10-17 20:00 | 20.5 | 46.4
2020-10-17 20:05 | 20.4 | 46.4
2020-10-17 20:10 | 20.5 | 46.3
2020-10-17 20:15 | 20.5 | 46.3
2020-10-17 20:20 | 20.5 | 46.3
2020-10-17 20:25 | 20.5 | 46.3
2020-10-17 20:30 | 20.5 | 46.2
2020-10-17 20:35 | 20.5 | 46.3
2020-10-17 20:40 | 20.5 | 46.2
2020-10-17 20:45 | 20.5 | 46.2
2020-10-17 20:50 | 20.5 | 46.2
2020-10-17 20:55 | 20.5 | 46.2
2020-10-17 21:00 | 20.5 | 46.1
2020-10-17 21:05 | 20.4 | 46.1
2020-10-17 21:10 | 20.5 | 46
2020-10-17 21:15 | 20.5 | 46
2020-10-17 21:20 | 20.5 | 45.9
2020-10-17 21:25 | 20.5 | 45.9
2020-10-17 21:30 | 20.4 | 45.9
2020-10-17 21:35 | 20.5 | 45.8
2020-10-17 21:40 | 20.5 | 45.7
2020-10-17 21:45 | 20.5 | 45.6
2020-10-17 21:50 | 20.5 | 45.6
2020-10-17 21:55 | 20.5 | 45.5
2020-10-17 22:00 | 20.5 | 45.5
2020-10-17 22:05 | 20.5 | 45.5
2020-10-17 22:10 | 20.5 | 45.5
2020-10-17 22:15 | 20.5 | 45.5
2020-10-17 22:20 | 20.5 | 45.4
2020-10-17 22:25 | 20.5 | 45.3
2020-10-17 22:30 | 20.5 | 45.3
2020-10-17 22:35 | 20.5 | 45.3
2020-10-17 22:40 | 20.5 | 45.2
2020-10-17 22:45 | 20.5 | 45.1
2020-10-17 22:50 | 20.5 | 45.1
2020-10-17 22:55 | 20.5 | 45
2020-10-17 23:00 | 20.5 | 45
2020-10-17 23:05 | 20.5 | 45
2020-10-17 23:10 | 20.5 | 44.9
2020-10-17 23:15 | 20.5 | 44.8
2020-10-17 23:20 | 20.5 | 44.7
2020-10-17 23:25 | 20.6 | 44.6
2020-10-17 23:30 | 20.5 | 44.7
2020-10-17 23:35 | 20.5 | 44.6
2020-10-17 23:40 | 20.5 | 44.6
2020-10-17 23:45 | 20.6 | 44.5
2020-10-17 23:50 | 20.5 | 44.5
2020-10-17 23:55 | 20.5 | 44.5
2020-10-18 | 20.5 | 44.5
2020-10-18 00:05 | 20.5 | 44.5
2020-10-18 00:10 | 20.5 | 44.5
2020-10-18 00:15 | 20.5 | 44.4
2020-10-18 00:20 | 20.5 | 44.3
2020-10-18 00:25 | 20.5 | 44.3
2020-10-18 00:30 | 20.5 | 44.3
2020-10-18 00:35 | 20.5 | 44.3
2020-10-18 00:40 | 20.5 | 44.3
2020-10-18 00:45 | 20.5 | 44.3
2020-10-18 00:50 | 20.5 | 44.2
2020-10-18 00:55 | 20.5 | 44.2
2020-10-18 01:00 | 20.5 | 44.1
2020-10-18 01:05 | 20.5 | 44.1
2020-10-18 01:10 | 20.5 | 44.1
2020-10-18 01:15 | 20.5 | 44.1
2020-10-18 01:20 | 20.5 | 44.1
2020-10-18 01:25 | 20.6 | 44
2020-10-18 01:30 | 20.5 | 44
2020-10-18 01:35 | 20.5 | 44.1
2020-10-18 01:40 | 20.5 | 44
2020-10-18 01:45 | 20.5 | 44
2020-10-18 01:50 | 20.5 | 44
2020-10-18 01:55 | 20.5 | 44
2020-10-18 02:00 | 20.5 | 44
2020-10-18 02:05 | 20.5 | 44
2020-10-18 02:10 | 20.5 | 43.9
2020-10-18 02:15 | 20.5 | 44
2020-10-18 02:20 | 20.5 | 44
2020-10-18 02:25 | 20.5 | 43.9
2020-10-18 02:30 | 20.5 | 43.9
2020-10-18 02:35 | 20.5 | 43.9
2020-10-18 02:40 | 20.5 | 43.8
2020-10-18 02:45 | 20.5 | 43.8
2020-10-18 02:50 | 20.5 | 43.8
2020-10-18 02:55 | 20.5 | 43.8
2020-10-18 03:00 | 20.5 | 43.8
2020-10-18 03:05 | 20.5 | 43.8
2020-10-18 03:10 | 20.5 | 43.9
2020-10-18 03:15 | 20.5 | 43.9
2020-10-18 03:20 | 20.5 | 44
2020-10-18 03:25 | 20.5 | 44.1
2020-10-18 03:30 | 20.5 | 44.2
2020-10-18 03:35 | 20.6 | 44.1
2020-10-18 03:40 | 20.5 | 44.2
2020-10-18 03:45 | 20.6 | 44.1
2020-10-18 03:50 | 20.5 | 44
2020-10-18 03:55 | 20.5 | 44
2020-10-18 04:00 | 20.5 | 43.9
2020-10-18 04:05 | 20.5 | 44
2020-10-18 04:10 | 20.5 | 44
2020-10-18 04:15 | 20.5 | 44
2020-10-18 04:20 | 20.5 | 43.9
2020-10-18 04:25 | 20.5 | 43.9
2020-10-18 04:30 | 20.5 | 43.8
2020-10-18 04:35 | 20.5 | 43.8
2020-10-18 04:40 | 20.5 | 43.8
2020-10-18 04:45 | 20.6 | 43.7
2020-10-18 04:50 | 20.5 | 43.8
2020-10-18 04:55 | 20.5 | 43.8
2020-10-18 05:00 | 20.5 | 43.8
2020-10-18 05:05 | 20.5 | 43.8
2020-10-18 05:10 | 20.5 | 43.9
2020-10-18 05:15 | 20.5 | 44
2020-10-18 05:20 | 20.5 | 44.1
2020-10-18 05:25 | 20.5 | 44.2
2020-10-18 05:30 | 20.5 | 44.1
2020-10-18 05:35 | 20.5 | 44.1
2020-10-18 05:40 | 20.6 | 44
2020-10-18 05:45 | 20.5 | 44
2020-10-18 05:50 | 20.5 | 44
2020-10-18 05:55 | 20.5 | 44
2020-10-18 06:00 | 20.6 | 44
2020-10-18 06:05 | 20.5 | 44.1
2020-10-18 06:10 | 20.5 | 44.2
2020-10-18 06:15 | 20.5 | 44.3
2020-10-18 06:20 | 20.5 | 44.3
2020-10-18 06:25 | 20.5 | 44.3
2020-10-18 06:30 | 20.5 | 44.3
2020-10-18 06:35 | 20.5 | 44.4
2020-10-18 06:40 | 20.5 | 44.4
2020-10-18 06:45 | 20.5 | 44.5
2020-10-18 06:50 | 20.5 | 44.5
2020-10-18 06:55 | 20.5 | 44.5
2020-10-18 07:00 | 20.5 | 44.5
2020-10-18 07:05 | 20.5 | 44.5
2020-10-18 07:10 | 20.5 | 44.7
2020-10-18 07:15 | 20.5 | 44.7
2020-10-18 07:20 | 20.5 | 44.7
2020-10-18 07:25 | 20.5 | 44.8
2020-10-18 07:30 | 20.5 | 44.8
2020-10-18 07:35 | 20.5 | 44.9
2020-10-18 07:40 | 20.5 | 44.8
2020-10-18 07:45 | 20.5 | 44.9
2020-10-18 07:50 | 20.5 | 45
2020-10-18 07:55 | 20.5 | 45
2020-10-18 08:00 | 20.5 | 45
2020-10-18 08:05 | 20.5 | 45
2020-10-18 08:10 | 20.5 | 45
2020-10-18 08:15 | 20.5 | 45
2020-10-18 08:20 | 20.5 | 45
2020-10-18 08:25 | 20.5 | 45
2020-10-18 08:30 | 20.5 | 44.9
2020-10-18 08:35 | 20.5 | 45
2020-10-18 08:40 | 20.5 | 44.9
2020-10-18 08:45 | 20.5 | 45
2020-10-18 08:50 | 20.5 | 45.1
2020-10-18 08:55 | 20.5 | 45.2
2020-10-18 09:00 | 20.5 | 45.2
2020-10-18 09:05 | 20.5 | 45.3
2020-10-18 09:10 | 20.5 | 45.5
2020-10-18 09:15 | 20.5 | 45.5
2020-10-18 09:20 | 20.5 | 45.6
2020-10-18 09:25 | 20.5 | 45.6
2020-10-18 09:30 | 20.5 | 45.7
2020-10-18 09:35 | 20.5 | 45.7
2020-10-18 09:40 | 20.5 | 45.9
2020-10-18 09:45 | 20.5 | 45.9
2020-10-18 09:50 | 20.5 | 46.1
2020-10-18 09:55 | 20.5 | 46.1
2020-10-18 10:00 | 20.5 | 46.2
2020-10-18 10:05 | 20.5 | 46.3
2020-10-18 10:10 | 20.5 | 46.4
2020-10-18 10:15 | 20.5 | 46.6
2020-10-18 10:20 | 20.5 | 46.7
2020-10-18 10:25 | 20.4 | 46.8
2020-10-18 10:30 | 20.5 | 46.8
2020-10-18 10:35 | 20.5 | 46.9
2020-10-18 10:40 | 20.5 | 47
2020-10-18 10:45 | 20.5 | 47.2
2020-10-18 10:50 | 20.5 | 47.2
2020-10-18 10:55 | 20.5 | 47.4
2020-10-18 11:00 | 20.5 | 47.5
2020-10-18 11:05 | 20.5 | 47.7
2020-10-18 11:10 | 20.4 | 48
2020-10-18 11:15 | 20.4 | 48.1
2020-10-18 11:20 | 20.5 | 48.1
2020-10-18 11:25 | 20.5 | 48.2
2020-10-18 11:30 | 20.4 | 48.3
2020-10-18 11:35 | 20.4 | 48.4
2020-10-18 11:40 | 20.4 | 48.5
2020-10-18 11:45 | 20.5 | 48.5
2020-10-18 11:50 | 20.5 | 48.6
2020-10-18 11:55 | 20.4 | 48.6
2020-10-18 12:00 | 20.4 | 48.8
2020-10-18 12:05 | 20.4 | 48.8
2020-10-18 12:10 | 20.4 | 48.7
2020-10-18 12:15 | 20.4 | 48.8
2020-10-18 12:20 | 20.4 | 48.7
2020-10-18 12:25 | 20.5 | 48.8
2020-10-18 12:30 | 20.4 | 48.8
2020-10-18 12:35 | 20.5 | 48.8
2020-10-18 12:40 | 20.4 | 48.8
2020-10-18 12:45 | 20.4 | 48.9
2020-10-18 12:50 | 20.4 | 49
2020-10-18 12:55 | 20.4 | 49
2020-10-18 13:00 | 20.4 | 49
2020-10-18 13:05 | 20.4 | 49.1
2020-10-18 13:10 | 20.4 | 49.1
2020-10-18 13:15 | 20.4 | 49.2
2020-10-18 13:20 | 20.4 | 49.2
2020-10-18 13:25 | 20.4 | 49.2
2020-10-18 13:30 | 20.4 | 49.3
2020-10-18 13:35 | 20.4 | 49.3
2020-10-18 13:40 | 20.4 | 49.4
2020-10-18 13:45 | 20.4 | 49.4
2020-10-18 13:50 | 20.4 | 49.5
2020-10-18 13:55 | 20.4 | 49.5
2020-10-18 14:00 | 20.3 | 49.6
2020-10-18 14:05 | 20.4 | 49.6
2020-10-18 14:10 | 20.4 | 49.5
2020-10-18 14:15 | 20.4 | 49.5
2020-10-18 14:20 | 20.4 | 49.4
2020-10-18 14:25 | 20.4 | 49.4
2020-10-18 14:30 | 20.4 | 49.3
2020-10-18 14:35 | 20.4 | 49.3
2020-10-18 14:40 | 20.4 | 49.3
2020-10-18 14:45 | 20.4 | 49.3
2020-10-18 14:50 | 20.4 | 49.3
2020-10-18 14:55 | 20.4 | 49.3
2020-10-18 15:00 | 20.4 | 49.2
2020-10-18 15:05 | 20.4 | 49.2
2020-10-18 15:10 | 20.4 | 49.2
2020-10-18 15:15 | 20.4 | 49.3
2020-10-18 15:20 | 20.4 | 49.3
2020-10-18 15:25 | 20.4 | 49.4
2020-10-18 15:30 | 20.4 | 49.5
2020-10-18 15:35 | 20.4 | 49.5
2020-10-18 15:40 | 20.4 | 49.6
2020-10-18 15:45 | 20.4 | 49.6
2020-10-18 15:50 | 20.4 | 49.6
2020-10-18 15:55 | 20.4 | 49.7
2020-10-18 16:00 | 20.4 | 49.7
2020-10-18 16:05 | 20.4 | 49.8
2020-10-18 16:10 | 20.4 | 49.9
2020-10-18 16:15 | 20.4 | 50
2020-10-18 16:20 | 20.4 | 50.1
2020-10-18 16:25 | 20.4 | 50.2
2020-10-18 16:30 | 20.4 | 50.3
2020-10-18 16:35 | 20.4 | 50.3
2020-10-18 16:40 | 20.4 | 50.4
2020-10-18 16:45 | 20.4 | 50.6
2020-10-18 16:50 | 20.4 | 50.5
2020-10-18 16:55 | 20.4 | 50.6
2020-10-18 17:00 | 20.4 | 50.7
2020-10-18 17:05 | 20.4 | 50.8
2020-10-18 17:10 | 20.4 | 51
2020-10-18 17:15 | 20.4 | 51.1
2020-10-18 17:20 | 20.4 | 51.1
2020-10-18 17:25 | 20.4 | 51.3
2020-10-18 17:30 | 20.4 | 51.3
2020-10-18 17:35 | 20.4 | 51.3
2020-10-18 17:40 | 20.4 | 51.2
2020-10-18 17:45 | 20.4 | 51.3
2020-10-18 17:50 | 20.4 | 51.2
2020-10-18 17:55 | 20.4 | 51.2
2020-10-18 18:00 | 20.4 | 51.1
2020-10-18 18:05 | 20.4 | 51
2020-10-18 18:10 | 20.4 | 50.9
2020-10-18 18:15 | 20.4 | 50.8
2020-10-18 18:20 | 20.4 | 50.7
2020-10-18 18:25 | 20.4 | 50.7
2020-10-18 18:30 | 20.4 | 50.6
2020-10-18 18:35 | 20.4 | 50.6
2020-10-18 18:40 | 20.4 | 50.5
2020-10-18 18:45 | 20.4 | 50.3
2020-10-18 18:50 | 20.4 | 50.3
2020-10-18 18:55 | 20.4 | 50.2
2020-10-18 19:00 | 20.4 | 50.2
2020-10-18 19:05 | 20.4 | 50.2
2020-10-18 19:10 | 20.4 | 50.2
2020-10-18 19:15 | 20.4 | 50.2
2020-10-18 19:20 | 20.4 | 50.1
2020-10-18 19:25 | 20.4 | 50.1
2020-10-18 19:30 | 20.4 | 50.1
2020-10-18 19:35 | 20.4 | 50
2020-10-18 19:40 | 20.4 | 50.1
2020-10-18 19:45 | 20.4 | 50.2
2020-10-18 19:50 | 20.4 | 50.2
2020-10-18 19:55 | 20.4 | 50.2
2020-10-18 20:00 | 20.4 | 50.3
2020-10-18 20:05 | 20.4 | 50.6
2020-10-18 20:10 | 20.4 | 50.8
2020-10-18 20:15 | 20.4 | 51
2020-10-18 20:20 | 20.4 | 51.2
2020-10-18 20:25 | 20.4 | 51.3
2020-10-18 20:30 | 20.4 | 51.4
2020-10-18 20:35 | 20.4 | 51.5
2020-10-18 20:40 | 20.4 | 51.7
2020-10-18 20:45 | 20.4 | 51.7
2020-10-18 20:50 | 20.4 | 51.6
2020-10-18 20:55 | 20.5 | 51.6
2020-10-18 21:00 | 20.4 | 51.6
2020-10-18 21:05 | 20.4 | 51.7
2020-10-18 21:10 | 20.4 | 51.8
2020-10-18 21:15 | 20.4 | 52
2020-10-18 21:20 | 20.4 | 52.1
2020-10-18 21:25 | 20.4 | 52.3
2020-10-18 21:30 | 20.4 | 52.5
2020-10-18 21:35 | 20.4 | 52.6
2020-10-18 21:40 | 20.5 | 52.6
2020-10-18 21:45 | 20.5 | 52.7
2020-10-18 21:50 | 20.5 | 52.7
2020-10-18 21:55 | 20.5 | 52.7
2020-10-18 22:00 | 20.5 | 52.7
2020-10-18 22:05 | 20.4 | 52.9
2020-10-18 22:10 | 20.4 | 52.9
2020-10-18 22:15 | 20.5 | 52.8
2020-10-18 22:20 | 20.5 | 52.8
2020-10-18 22:25 | 20.5 | 52.7
2020-10-18 22:30 | 20.4 | 52.8
2020-10-18 22:35 | 20.4 | 52.7
2020-10-18 22:40 | 20.5 | 52.5
2020-10-18 22:45 | 20.5 | 52.5
2020-10-18 22:50 | 20.4 | 52.5
2020-10-18 22:55 | 20.4 | 52.5
2020-10-18 23:00 | 20.5 | 52.4
2020-10-18 23:05 | 20.5 | 52.3
2020-10-18 23:10 | 20.5 | 52.2
2020-10-18 23:15 | 20.4 | 52.2
2020-10-18 23:20 | 20.5 | 52.1
2020-10-18 23:25 | 20.5 | 52.1
2020-10-18 23:30 | 20.5 | 52.2
2020-10-18 23:35 | 20.5 | 52.4
2020-10-18 23:40 | 20.4 | 52.5
2020-10-18 23:45 | 20.4 | 52.5
2020-10-18 23:50 | 20.4 | 52.4
2020-10-18 23:55 | 20.4 | 52.4
2020-10-19 | 20.4 | 52.5
2020-10-19 00:05 | 20.5 | 52.5
2020-10-19 00:10 | 20.5 | 52.7
2020-10-19 00:15 | 20.5 | 52.8
2020-10-19 00:20 | 20.5 | 52.9
2020-10-19 00:25 | 20.5 | 52.9
2020-10-19 00:30 | 20.5 | 53
2020-10-19 00:35 | 20.5 | 53.1
2020-10-19 00:40 | 20.5 | 53
2020-10-19 00:45 | 20.5 | 53
2020-10-19 00:50 | 20.5 | 53
2020-10-19 00:55 | 20.5 | 53
2020-10-19 01:00 | 20.5 | 52.9
2020-10-19 01:05 | 20.5 | 52.9
2020-10-19 01:10 | 20.5 | 52.8
2020-10-19 01:15 | 20.5 | 52.8
2020-10-19 01:20 | 20.5 | 52.9
2020-10-19 01:25 | 20.4 | 52.9
2020-10-19 01:30 | 20.5 | 52.9
2020-10-19 01:35 | 20.4 | 52.8
2020-10-19 01:40 | 20.5 | 52.8
2020-10-19 01:45 | 20.5 | 52.7
2020-10-19 01:50 | 20.5 | 52.7
2020-10-19 01:55 | 20.5 | 52.6
2020-10-19 02:00 | 20.5 | 52.5
2020-10-19 02:05 | 20.5 | 52.5
2020-10-19 02:10 | 20.5 | 52.5
2020-10-19 02:15 | 20.5 | 52.4
2020-10-19 02:20 | 20.5 | 52.3
2020-10-19 02:25 | 20.5 | 52.4
2020-10-19 02:30 | 20.5 | 52.4
2020-10-19 02:35 | 20.5 | 52.3
2020-10-19 02:40 | 20.5 | 52.3
2020-10-19 02:45 | 20.5 | 52.3
2020-10-19 02:50 | 20.5 | 52.2
2020-10-19 02:55 | 20.5 | 52.2
2020-10-19 03:00 | 20.5 | 52.4
2020-10-19 03:05 | 20.5 | 52.2
2020-10-19 03:10 | 20.5 | 52.2
2020-10-19 03:15 | 20.5 | 52.2
2020-10-19 03:20 | 20.5 | 52.2
2020-10-19 03:25 | 20.4 | 52.2
2020-10-19 03:30 | 20.4 | 52.2
2020-10-19 03:35 | 20.5 | 52.1
2020-10-19 03:40 | 20.5 | 52
2020-10-19 03:45 | 20.5 | 51.8
2020-10-19 03:50 | 20.5 | 51.9
2020-10-19 03:55 | 20.5 | 52
2020-10-19 04:00 | 20.5 | 51.8
2020-10-19 04:05 | 20.5 | 51.8
2020-10-19 04:10 | 20.5 | 51.9
2020-10-19 04:15 | 20.5 | 51.8
2020-10-19 04:20 | 20.5 | 51.8
2020-10-19 04:25 | 20.4 | 51.9
2020-10-19 04:30 | 20.5 | 51.8
2020-10-19 04:35 | 20.5 | 51.7
2020-10-19 04:40 | 20.5 | 51.7
2020-10-19 04:45 | 20.5 | 51.7
2020-10-19 04:50 | 20.5 | 51.8
2020-10-19 04:55 | 20.5 | 51.9
2020-10-19 05:00 | 20.5 | 51.9
2020-10-19 05:05 | 20.5 | 51.9
2020-10-19 05:10 | 20.5 | 51.9
2020-10-19 05:15 | 20.5 | 52
2020-10-19 05:20 | 20.5 | 51.8
2020-10-19 05:25 | 20.5 | 51.8
2020-10-19 05:30 | 20.5 | 51.8
2020-10-19 05:35 | 20.4 | 51.8
2020-10-19 05:40 | 20.4 | 51.8
2020-10-19 05:45 | 20.5 | 51.7
2020-10-19 05:50 | 20.5 | 51.6
2020-10-19 05:55 | 20.5 | 51.5
2020-10-19 06:00 | 20.5 | 51.5
2020-10-19 06:05 | 20.5 | 51.5
2020-10-19 06:10 | 20.5 | 51.5
2020-10-19 06:15 | 20.5 | 51.5
2020-10-19 06:20 | 20.5 | 51.4
2020-10-19 06:25 | 20.5 | 51.5
2020-10-19 06:30 | 20.5 | 51.5
2020-10-19 06:35 | 20.5 | 51.5
2020-10-19 06:40 | 20.5 | 51.5
2020-10-19 06:45 | 20.5 | 51.4
2020-10-19 06:50 | 20.5 | 51.5
2020-10-19 06:55 | 20.5 | 51.6
2020-10-19 07:00 | 20.5 | 51.6
2020-10-19 07:05 | 20.5 | 51.6
2020-10-19 07:10 | 20.5 | 51.6
2020-10-19 07:15 | 20.5 | 51.6
2020-10-19 07:20 | 20.6 | 51.5
2020-10-19 07:25 | 20.5 | 51.5
2020-10-19 07:30 | 20.6 | 51.5
2020-10-19 07:35 | 20.6 | 51.5
2020-10-19 07:40 | 20.5 | 51.6
2020-10-19 07:45 | 20.5 | 51.7
2020-10-19 07:50 | 20.5 | 51.6
2020-10-19 07:55 | 20.5 | 51.7
2020-10-19 08:00 | 20.5 | 51.6
2020-10-19 08:05 | 20.5 | 51.6
2020-10-19 08:10 | 20.5 | 51.7
2020-10-19 08:15 | 20.5 | 51.7
2020-10-19 08:20 | 20.5 | 51.7
2020-10-19 08:25 | 20.5 | 51.8
2020-10-19 08:30 | 20.5 | 51.9
2020-10-19 08:35 | 20.5 | 51.8
2020-10-19 08:40 | 20.5 | 51.9
2020-10-19 08:45 | 20.5 | 52
2020-10-19 08:50 | 20.5 | 52
2020-10-19 08:55 | 20.5 | 52.2
2020-10-19 09:00 | 20.5 | 52.1
2020-10-19 09:05 | 20.5 | 52.3
2020-10-19 09:10 | 20.5 | 52.4
2020-10-19 09:15 | 20.5 | 52.3
2020-10-19 09:20 | 20.5 | 52.4
2020-10-19 09:25 | 20.5 | 52.5
2020-10-19 09:30 | 20.5 | 52.6
2020-10-19 09:35 | 20.5 | 52.6
2020-10-19 09:40 | 20.5 | 52.8
2020-10-19 09:45 | 20.5 | 52.8
2020-10-19 09:50 | 20.5 | 52.9
2020-10-19 09:55 | 20.5 | 53
2020-10-19 10:00 | 20.5 | 53.1
2020-10-19 10:05 | 20.5 | 53.2
2020-10-19 10:10 | 20.5 | 53.3
2020-10-19 10:15 | 20.5 | 53.4
2020-10-19 10:20 | 20.5 | 53.5
2020-10-19 10:25 | 20.5 | 53.7
2020-10-19 10:30 | 20.5 | 53.9
2020-10-19 10:35 | 20.4 | 54.1
2020-10-19 10:40 | 20.4 | 54.3
2020-10-19 10:45 | 20.5 | 54.3
2020-10-19 10:50 | 20.4 | 54.4
2020-10-19 10:55 | 20.4 | 54.5
2020-10-19 11:00 | 20.5 | 54.7
2020-10-19 11:05 | 20.5 | 54.7
2020-10-19 11:10 | 20.5 | 54.7
2020-10-19 11:15 | 20.4 | 54.8
2020-10-19 11:20 | 20.5 | 54.8
2020-10-19 11:25 | 20.4 | 55
2020-10-19 11:30 | 20.4 | 55
2020-10-19 11:35 | 20.4 | 55
2020-10-19 11:40 | 20.4 | 55.2
2020-10-19 11:45 | 20.4 | 55.1
2020-10-19 11:50 | 20.5 | 55.2
2020-10-19 11:55 | 20.5 | 55.3
2020-10-19 12:00 | 20.4 | 55.5
2020-10-19 12:05 | 20.4 | 55.6
2020-10-19 12:10 | 20.5 | 55.6
2020-10-19 12:15 | 20.4 | 55.9
2020-10-19 12:20 | 20.4 | 55.9
2020-10-19 12:25 | 20.4 | 56.2
2020-10-19 12:30 | 20.4 | 56.2
2020-10-19 12:35 | 20.4 | 56.3
2020-10-19 12:40 | 20.4 | 56.4
2020-10-19 12:45 | 20.4 | 56.6
2020-10-19 12:50 | 20.4 | 56.7
2020-10-19 12:55 | 20.4 | 56.7
2020-10-19 13:00 | 20.4 | 56.7
2020-10-19 13:05 | 20.4 | 56.8
2020-10-19 13:10 | 20.4 | 56.7
2020-10-19 13:15 | 20.4 | 56.8
2020-10-19 13:20 | 20.5 | 56.8
2020-10-19 13:25 | 20.4 | 56.9
2020-10-19 13:30 | 20.5 | 56.8
2020-10-19 13:35 | 20.5 | 56.8
2020-10-19 13:40 | 20.5 | 56.9
2020-10-19 13:45 | 20.4 | 57.1
2020-10-19 13:50 | 20.5 | 57
2020-10-19 13:55 | 20.4 | 57.1
2020-10-19 14:00 | 20.4 | 57.2
2020-10-19 14:05 | 20.5 | 57.3
2020-10-19 14:10 | 20.4 | 57.5
2020-10-19 14:15 | 20.5 | 57.7
2020-10-19 14:20 | 20.4 | 57.8
2020-10-19 14:25 | 20.5 | 57.7
2020-10-19 14:30 | 20.4 | 58
2020-10-19 14:35 | 20.5 | 58.3
2020-10-19 14:40 | 20.4 | 58.6
2020-10-19 14:45 | 20.5 | 58.4
2020-10-19 14:50 | 20.4 | 58.2
2020-10-19 14:55 | 20.4 | 58.3
2020-10-19 15:00 | 20.5 | 58.2
2020-10-19 15:05 | 20.5 | 58.1
2020-10-19 15:10 | 20.4 | 58.1
2020-10-19 15:15 | 20.4 | 58.1
2020-10-19 15:20 | 20.4 | 58.1
2020-10-19 15:25 | 20.4 | 58.1
2020-10-19 15:30 | 20.4 | 58.1
2020-10-19 15:35 | 20.4 | 58.1
2020-10-19 15:40 | 20.4 | 58
2020-10-19 15:45 | 20.5 | 57.7
2020-10-19 15:50 | 20.4 | 57.7
2020-10-19 15:55 | 20.5 | 57.6
2020-10-19 16:00 | 20.5 | 57.6
2020-10-19 16:05 | 20.5 | 57.7
2020-10-19 16:10 | 20.5 | 57.7
2020-10-19 16:15 | 20.5 | 57.9
2020-10-19 16:20 | 20.5 | 58
2020-10-19 16:25 | 20.4 | 58.1
2020-10-19 16:30 | 20.3 | 58.5
2020-10-19 16:35 | 20.4 | 58.3
2020-10-19 16:40 | 20.4 | 58.4
2020-10-19 16:45 | 20.4 | 58.3
2020-10-19 16:50 | 20.4 | 58.1
2020-10-19 16:55 | 20.4 | 58
2020-10-19 17:00 | 20.4 | 58.1
2020-10-19 17:05 | 20.4 | 58.2
2020-10-19 17:10 | 20.4 | 58
2020-10-19 17:15 | 20.4 | 57.9
2020-10-19 17:20 | 20.4 | 57.9
2020-10-19 17:25 | 20.4 | 58
2020-10-19 17:30 | 20.4 | 58.1
2020-10-19 17:35 | 20.5 | 58
2020-10-19 17:40 | 20.4 | 58.6
2020-10-19 17:45 | 20.5 | 58
2020-10-19 17:50 | 20.4 | 58.1
2020-10-19 17:55 | 20.4 | 57.9
2020-10-19 18:00 | 20.4 | 57.8
2020-10-19 18:05 | 20.4 | 57.7
2020-10-19 18:10 | 20.4 | 57.6
2020-10-19 18:15 | 20.4 | 57.6
2020-10-19 18:20 | 20.4 | 57.6
2020-10-19 18:25 | 20.5 | 57.4
2020-10-19 18:30 | 20.5 | 57.4
2020-10-19 18:35 | 20.4 | 57.5
2020-10-19 18:40 | 20.4 | 57.4
2020-10-19 18:45 | 20.5 | 57.4
2020-10-19 18:50 | 20.5 | 57.3
2020-10-19 18:55 | 20.5 | 57.3
2020-10-19 19:00 | 20.4 | 57.3
2020-10-19 19:05 | 20.4 | 57.2
2020-10-19 19:10 | 20.4 | 57.2
2020-10-19 19:15 | 20.4 | 57.2
2020-10-19 19:20 | 20.4 | 57.1
2020-10-19 19:25 | 20.5 | 57
2020-10-19 19:30 | 20.4 | 57.1
2020-10-19 19:35 | 20.5 | 57.1
2020-10-19 19:40 | 20.5 | 56.9
2020-10-19 19:45 | 20.5 | 56.8
2020-10-19 19:50 | 20.5 | 56.8
2020-10-19 19:55 | 20.5 | 56.6
2020-10-19 20:00 | 20.4 | 56.7
2020-10-19 20:05 | 20.5 | 56.6
2020-10-19 20:10 | 20.4 | 56.6
2020-10-19 20:15 | 20.5 | 56.6
2020-10-19 20:20 | 20.5 | 56.4
2020-10-19 20:25 | 20.5 | 56.3
2020-10-19 20:30 | 20.5 | 56.4
2020-10-19 20:35 | 20.5 | 56.3
2020-10-19 20:40 | 20.4 | 56.4
2020-10-19 20:45 | 20.4 | 56.4
2020-10-19 20:50 | 20.4 | 56.4
2020-10-19 20:55 | 20.4 | 56.4
2020-10-19 21:00 | 20.5 | 56.3
2020-10-19 21:05 | 20.5 | 56.3
2020-10-19 21:10 | 20.5 | 56.2
2020-10-19 21:15 | 20.5 | 56.1
2020-10-19 21:20 | 20.5 | 56
2020-10-19 21:25 | 20.5 | 56
2020-10-19 21:30 | 20.5 | 56
2020-10-19 21:35 | 20.5 | 56
2020-10-19 21:40 | 20.5 | 55.9
2020-10-19 21:45 | 20.5 | 55.9
2020-10-19 21:50 | 20.5 | 55.7
2020-10-19 21:55 | 20.5 | 55.8
2020-10-19 22:00 | 20.5 | 55.8
2020-10-19 22:05 | 20.5 | 55.7
2020-10-19 22:10 | 20.5 | 55.6
2020-10-19 22:15 | 20.5 | 55.6
2020-10-19 22:20 | 20.5 | 55.5
2020-10-19 22:25 | 20.5 | 55.6
2020-10-19 22:30 | 20.5 | 55.5
2020-10-19 22:35 | 20.5 | 55.5
2020-10-19 22:40 | 20.5 | 55.5
2020-10-19 22:45 | 20.5 | 55.4
2020-10-19 22:50 | 20.5 | 55.5
2020-10-19 22:55 | 20.5 | 55.4
2020-10-19 23:00 | 20.5 | 55.3
2020-10-19 23:05 | 20.5 | 55.4
2020-10-19 23:10 | 20.5 | 55.3
2020-10-19 23:15 | 20.5 | 55.2
2020-10-19 23:20 | 20.5 | 55.2
2020-10-19 23:25 | 20.5 | 55.2
2020-10-19 23:30 | 20.5 | 55.1
2020-10-19 23:35 | 20.5 | 55.2
2020-10-19 23:40 | 20.5 | 55.2
2020-10-19 23:45 | 20.5 | 55.1
2020-10-19 23:50 | 20.5 | 55.2
2020-10-19 23:55 | 20.5 | 55.1

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | td | utd | ktd | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for G | G | uG | k | multiplier for C | C | uC | k | nom Freq
coax zero | 1591.5 | 3Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 3.2278e-12 | 2.0895 | 0 | 3.2274e-08 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-05 | 2.0895 | 1000000 | 0 | 3.2278e-06 | 2.0895 | 1591.5
box_zero | 1591.5 | 3Y | 49 | 0 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 0 | 0 | 0 | 2.9379e-15 | 2.4469 | 0 | 0 | 0 | 1.2038e-15 | 1.5931e-11 | 0 | 0.00018442 | box_zero1591.5 | 1000 | 0 | 2.9379e-12 | 2.4469 | 1000000 | 0 | 0 | 0 | 1591.5
A1 | 1591.5 | 3Y | 100013 | 40 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.9973e-08 | 1.1976e-11 | 2.0231 | 3.9998e-08 | 5.6572e-08 | 2.0154 | 4.001e-05 | 5.6589e-05 | 2.0154 | 2.2006e-12 | 2.7008e-08 | 5.4421 | 2.0947 | A11591.5 | 1000 | 3.9998e-05 | 5.6572e-05 | 2.0154 | 1000000 | 0.099973 | 1.1976e-05 | 2.0231 | 1591.5
A2 | 1591.5 | 3Y | 200413 | 90 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 2.0038e-07 | 2.2995e-11 | 2.0381 | 8.9997e-08 | 9.2766e-08 | 2.0596 | 4.4914e-05 | 4.6296e-05 | 2.0596 | 4.4096e-12 | 5.4117e-08 | 5.2149 | 1.7142 | A21591.5 | 1000 | 8.9997e-05 | 9.2766e-05 | 2.0596 | 1000000 | 0.20038 | 2.2995e-05 | 2.0381 | 1591.5
A3 | 1591.5 | 3Y | 300376 | 119 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 3.0035e-07 | 3.4163e-11 | 2.0415 | 1.19e-07 | 1.4934e-07 | 2.0927 | 3.962e-05 | 4.9721e-05 | 2.0927 | 6.609e-12 | 8.1109e-08 | 5.1691 | 1.8412 | A31591.5 | 1000 | 0.000119 | 0.00014934 | 2.0927 | 1000000 | 0.30035 | 3.4163e-05 | 2.0415 | 1591.5
A4 | 1591.5 | 3Y | 400767 | 209 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.0075e-07 | 4.5439e-11 | 2.0427 | 2.0899e-07 | 2.2923e-07 | 2.1357 | 5.2151e-05 | 5.7201e-05 | 2.1357 | 8.8178e-12 | 1.0822e-07 | 5.1531 | 2.1182 | A41591.5 | 1000 | 0.00020899 | 0.00022923 | 2.1357 | 1000000 | 0.40075 | 4.5439e-05 | 2.0427 | 1591.5
A5 | 1591.5 | 3Y | 500576 | 346 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 5.0057e-07 | 5.6674e-11 | 2.0433 | 3.4599e-07 | 3.3355e-07 | 2.1832 | 6.912e-05 | 6.6636e-05 | 2.1832 | 1.1014e-11 | 1.3517e-07 | 5.1457 | 2.4677 | A51591.5 | 1000 | 0.00034599 | 0.00033355 | 2.1832 | 1000000 | 0.50057 | 5.6674e-05 | 2.0433 | 1591.5
A6 | 1591.5 | 3Y | 600560 | 84 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 6.0057e-07 | 6.794e-11 | 2.0437 | 8.3997e-08 | 4.6353e-07 | 2.2271 | 1.3987e-05 | 7.7184e-05 | 2.2271 | 1.3214e-11 | 1.6216e-07 | 5.1417 | 2.8584 | A61591.5 | 1000 | 8.3997e-05 | 0.00046353 | 2.2271 | 1000000 | 0.60057 | 6.794e-05 | 2.0437 | 1591.5
A7 | 1591.5 | 3Y | 700940 | 522 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 7.0095e-07 | 7.9259e-11 | 2.0439 | 5.2198e-07 | 6.1964e-07 | 2.2643 | 7.4469e-05 | 8.8402e-05 | 2.2643 | 1.5422e-11 | 1.8927e-07 | 5.1392 | 3.2739 | A71591.5 | 1000 | 0.00052198 | 0.00061964 | 2.2643 | 1000000 | 0.70095 | 7.9259e-05 | 2.0439 | 1591.5
A8 | 1591.5 | 3Y | 800803 | 736 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.0083e-07 | 9.0523e-11 | 2.044 | 7.3597e-07 | 8.0025e-07 | 2.2944 | 9.1904e-05 | 9.9931e-05 | 2.2944 | 1.7619e-11 | 2.1623e-07 | 5.1377 | 3.7009 | A81591.5 | 1000 | 0.00073597 | 0.00080025 | 2.2944 | 1000000 | 0.80083 | 9.0523e-05 | 2.044 | 1591.5
A9 | 1591.5 | 3Y | 901292 | 1013 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.0132e-07 | 1.0186e-10 | 2.0441 | 1.013e-06 | 1.0073e-06 | 2.3187 | 0.00011239 | 0.00011176 | 2.3187 | 1.983e-11 | 2.4337e-07 | 5.1366 | 4.1389 | A91591.5 | 1000 | 0.001013 | 0.0010073 | 2.3187 | 1000000 | 0.90132 | 0.00010186 | 2.0441 | 1591.5
A10 | 1591.5 | 3Y | 1001258 | 1327 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0013e-06 | 1.1314e-10 | 2.0441 | 1.3269e-06 | 1.2382e-06 | 2.3381 | 0.00013253 | 0.00012367 | 2.3381 | 2.203e-11 | 2.7036e-07 | 5.1358 | 4.58 | A101591.5 | 1000 | 0.0013269 | 0.0012382 | 2.3381 | 1000000 | 1.0013 | 0.00011314 | 2.0441 | 1591.5
coax zero | 1591.5 | 4Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 3.2278e-13 | 2.0895 | 0 | 3.2274e-09 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-06 | 2.0895 | 1000000000 | 0 | 0.00032278 | 2.0895 | 1591.5
box_zero | 1591.5 | 4Y | 500 | 10 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 0 | 0 | 0 | 3.059e-15 | 2.4469 | 0 | 0 | 0 | 1.2258e-15 | 1.6201e-11 | 0 | 0.00018882 | box_zero1591.5 | 1000 | 0 | 3.059e-12 | 2.4469 | 1000000000 | 0 | 0 | 0 | 1591.5
B1 | 1591.5 | 4Y | 100618 | 47 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0013e-08 | 1.2216e-12 | 2.0259 | 3.6999e-09 | 5.5802e-09 | 2.0188 | 3.6953e-05 | 5.5733e-05 | 2.0188 | 2.2151e-13 | 2.7195e-09 | 5.5152 | 2.0519 | B11591.5 | 1000 | 3.6999e-06 | 5.5802e-06 | 2.0188 | 1000000000 | 10.0127 | 0.0012216 | 2.0259 | 1591.5
B2 | 1591.5 | 4Y | 200526 | 67 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 2.0004e-08 | 2.3194e-12 | 2.0396 | 5.6998e-09 | 8.4289e-09 | 2.0907 | 2.8494e-05 | 4.2136e-05 | 2.0907 | 4.4132e-13 | 5.4172e-09 | 5.2557 | 1.5559 | B21591.5 | 1000 | 5.6998e-06 | 8.4289e-06 | 2.0907 | 1000000000 | 20.0044 | 0.0023194 | 2.0396 | 1591.5
B3 | 1591.5 | 4Y | 300641 | 70 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 3.0017e-08 | 3.4806e-12 | 2.0444 | 5.9998e-09 | 1.1882e-08 | 2.1432 | 1.9989e-05 | 3.9587e-05 | 2.1432 | 6.616e-13 | 8.1205e-09 | 5.2609 | 1.4632 | B31591.5 | 1000 | 5.9998e-06 | 1.1882e-05 | 2.1432 | 1000000000 | 30.0168 | 0.0034806 | 2.0444 | 1591.5
B4 | 1591.5 | 4Y | 400667 | 64 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.002e-08 | 4.6032e-12 | 2.0449 | 5.3998e-09 | 1.5527e-08 | 2.1698 | 1.3493e-05 | 3.88e-05 | 2.1698 | 8.8167e-13 | 1.0821e-08 | 5.221 | 1.4349 | B41591.5 | 1000 | 5.3998e-06 | 1.5527e-05 | 2.1698 | 1000000000 | 40.0203 | 0.0046032 | 2.0449 | 1591.5
B5 | 1591.5 | 4Y | 500749 | 57 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 5.0029e-08 | 5.773e-12 | 2.0463 | 4.6998e-09 | 1.9262e-08 | 2.1823 | 9.3944e-06 | 3.8503e-05 | 2.1823 | 1.1019e-12 | 1.3524e-08 | 5.2392 | 1.4243 | B51591.5 | 1000 | 4.6998e-06 | 1.9262e-05 | 2.1823 | 1000000000 | 50.0294 | 0.005773 | 2.0463 | 1591.5
B6 | 1591.5 | 4Y | 600867 | -61 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 6.0042e-08 | 6.8891e-12 | 2.0458 | -7.0997e-09 | 2.3056e-08 | 2.1873 | -1.1825e-05 | 3.8401e-05 | 2.1873 | 1.3222e-12 | 1.6227e-08 | 5.2105 | 1.4208 | B61591.5 | 1000 | -7.0997e-06 | 2.3056e-05 | 2.1873 | 1000000000 | 60.0421 | 0.0068891 | 2.0458 | 1591.5
B7 | 1591.5 | 4Y | 700768 | 99 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 7.0033e-08 | 8.0311e-12 | 2.046 | 8.8997e-09 | 2.689e-08 | 2.188 | 1.2708e-05 | 3.8397e-05 | 2.188 | 1.542e-12 | 1.8925e-08 | 5.2084 | 1.4209 | B71591.5 | 1000 | 8.8997e-06 | 2.689e-05 | 2.188 | 1000000000 | 70.0331 | 0.0080311 | 2.046 | 1591.5
B8 | 1591.5 | 4Y | 800879 | 189 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.0045e-08 | 9.2201e-12 | 2.0469 | 1.7899e-08 | 3.0781e-08 | 2.1861 | 2.2362e-05 | 3.8456e-05 | 2.1861 | 1.7622e-12 | 2.1628e-08 | 5.2321 | 1.4232 | B81591.5 | 1000 | 1.7899e-05 | 3.0781e-05 | 2.1861 | 1000000000 | 80.0451 | 0.0092201 | 2.0469 | 1591.5
B9 | 1591.5 | 4Y | 900901 | 329 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.0048e-08 | 1.032e-11 | 2.0462 | 3.1899e-08 | 3.4722e-08 | 2.1826 | 3.5425e-05 | 3.8561e-05 | 2.1826 | 1.9823e-12 | 2.4328e-08 | 5.2061 | 1.4272 | B91591.5 | 1000 | 3.1899e-05 | 3.4722e-05 | 2.1826 | 1000000000 | 90.0482 | 0.01032 | 2.0462 | 1591.5
B10 | 1591.5 | 4Y | 1001011 | 514 | 1591.5 | 19.71 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0006e-07 | 1.152e-11 | 2.047 | 5.0398e-08 | 3.8725e-08 | 2.178 | 5.0369e-05 | 3.8703e-05 | 2.178 | 2.2025e-12 | 2.7032e-08 | 5.2305 | 1.4326 | B101591.5 | 1000 | 5.0398e-05 | 3.8725e-05 | 2.178 | 1000000000 | 100.0601 | 0.01152 | 2.047 | 1591.5
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 3.2278e-14 | 2.0895 | 0 | 3.2274e-10 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-07 | 2.0895 | 1000000000 | 0 | 3.2278e-05 | 2.0895 | 1591.5
box_zero | 1591.5 | 5Y | 5001 | 123 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 0 | 0 | 0 | 3.0602e-15 | 2.4469 | 0 | 0 | 0 | 1.226e-15 | 1.6203e-11 | 0 | 0.00018886 | box_zero1591.5 | 1000 | 0 | 3.0602e-12 | 2.4469 | 1000000000 | 0 | 0 | 0 | 1591.5
C1 | 1591.5 | 5Y | 104502 | 163 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.951e-10 | 1.2363e-13 | 2.0264 | 3.9998e-10 | 5.5656e-10 | 2.0185 | 4.0197e-05 | 5.5932e-05 | 2.0185 | 2.3118e-14 | 2.8487e-10 | 5.3476 | 1.9537 | C11591.5 | 1000 | 3.9998e-07 | 5.5656e-07 | 2.0185 | 1000000000 | 0.9951 | 0.00012363 | 2.0264 | 1591.5
C2 | 1591.5 | 5Y | 204731 | 157 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.9975e-09 | 2.3234e-13 | 2.0387 | 3.3999e-10 | 8.4116e-10 | 2.0912 | 1.7021e-05 | 4.2112e-05 | 2.0912 | 4.5171e-14 | 5.5551e-10 | 5.1436 | 1.5142 | C21591.5 | 1000 | 3.3999e-07 | 8.4116e-07 | 2.0912 | 1000000000 | 1.9975 | 0.00023234 | 2.0387 | 1591.5
C3 | 1591.5 | 5Y | 304202 | 160 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 2.9923e-09 | 3.4299e-13 | 2.0417 | 3.6999e-10 | 1.1824e-09 | 2.1455 | 1.2365e-05 | 3.9517e-05 | 2.1455 | 6.7056e-14 | 8.2409e-10 | 5.115 | 1.4348 | C31591.5 | 1000 | 3.6999e-07 | 1.1824e-06 | 2.1455 | 1000000000 | 2.9923 | 0.00034299 | 2.0417 | 1591.5
C4 | 1591.5 | 5Y | 404451 | 146 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 3.9949e-09 | 4.5705e-13 | 2.0434 | 2.2999e-10 | 1.5445e-09 | 2.1754 | 5.7574e-06 | 3.8664e-05 | 2.1754 | 8.9113e-14 | 1.0948e-09 | 5.1289 | 1.4108 | C41591.5 | 1000 | 2.2999e-07 | 1.5445e-06 | 2.1754 | 1000000000 | 3.9949 | 0.00045705 | 2.0434 | 1591.5
C5 | 1591.5 | 5Y | 504786 | 116 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.9983e-09 | 5.6999e-13 | 2.0439 | -6.9997e-11 | 1.9139e-09 | 2.1922 | -1.4005e-06 | 3.8292e-05 | 2.1922 | 1.1119e-13 | 1.3657e-09 | 5.1263 | 1.4014 | C51591.5 | 1000 | -6.9997e-08 | 1.9139e-06 | 2.1922 | 1000000000 | 4.9983 | 0.00056999 | 2.0439 | 1591.5
C6 | 1591.5 | 5Y | 604205 | 67 | 1591.5 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 5.9926e-09 | 6.8202e-13 | 2.0441 | -5.5998e-10 | 2.2831e-09 | 2.2022 | -9.3448e-06 | 3.8099e-05 | 2.2022 | 1.3306e-13 | 1.6341e-09 | 5.1255 | 1.3971 | C61591.5 | 1000 | -5.5998e-07 | 2.2831e-06 | 2.2022 | 1000000000 | 5.9926 | 0.00068202 | 2.0441 | 1591.5
C7 | 1591.5 | 5Y | 704439 | 137 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 6.995e-09 | 7.9164e-13 | 2.0436 | 1.3999e-10 | 2.657e-09 | 2.2085 | 2.0014e-06 | 3.7985e-05 | 2.2085 | 1.5512e-13 | 1.9048e-09 | 5.1035 | 1.3949 | C71591.5 | 1000 | 1.3999e-07 | 2.657e-06 | 2.2085 | 1000000000 | 6.995 | 0.00079164 | 2.0436 | 1591.5
C8 | 1591.5 | 5Y | 803838 | 238 | 1591.5 | 19.76 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 7.9891e-09 | 9.0328e-13 | 2.0437 | 1.15e-09 | 3.0289e-09 | 2.2128 | 1.4395e-05 | 3.7914e-05 | 2.2128 | 1.7699e-13 | 2.1732e-09 | 5.1037 | 1.3937 | C81591.5 | 1000 | 1.15e-06 | 3.0289e-06 | 2.2128 | 1000000000 | 7.9891 | 0.00090328 | 2.0437 | 1591.5
C9 | 1591.5 | 5Y | 904057 | 329 | 1591.5 | 19.78 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.9914e-09 | 1.0077e-12 | 2.0425 | 2.0599e-09 | 3.4045e-09 | 2.2157 | 2.2911e-05 | 3.7865e-05 | 2.2157 | 1.9904e-13 | 2.4438e-09 | 5.0627 | 1.3931 | C91591.5 | 1000 | 2.0599e-06 | 3.4045e-06 | 2.2157 | 1000000000 | 8.9914 | 0.0010077 | 2.0425 | 1591.5
C10 | 1591.5 | 5Y | 1003463 | 493 | 1591.5 | 19.86 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.9855e-09 | 1.0896e-12 | 2.0388 | 3.6999e-09 | 3.7776e-09 | 2.2178 | 3.7053e-05 | 3.7832e-05 | 2.2178 | 2.2091e-13 | 2.7122e-09 | 4.9324 | 1.3928 | C101591.5 | 1000 | 3.6999e-06 | 3.7776e-06 | 2.2178 | 1000000000 | 9.9855 | 0.0010896 | 2.0388 | 1591.5
coax zero | 1591.5 | 5Y | 0 | 0 | 1591.5 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 3.2278e-14 | 2.0895 | 0 | 3.2274e-10 | 2.0895 | 1.2567e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 0 | 3.2274e-07 | 2.0895 | 1000000000000 | 0 | 0.032278 | 2.0895 | 1591.5
main_zero | 1591.5 | 6Y | 50024 | 1095 | 1591.5 | 19.67 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.9144e-11 | 1.5423e-12 | 2.5705 | 1.08e-09 | 4.7013e-11 | 2.0115 | 0.0021978 | 0.00011021 | 2.0182 | 1.2264e-15 | 1.6208e-11 | 1257.6554 | 2.9006 | main_zero1591.5 | 1000 | 1.08e-06 | 4.7013e-08 | 2.0115 | 1000000000000 | 49.1435 | 1.5423 | 2.5705 | 1591.5
D1 | 1591.5 | 6Y | 146641 | 1150 | 1591.5 | 19.72 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.6632e-11 | 1.6686e-14 | 2.0391 | 5.5001e-11 | 5.5027e-11 | 2.0171 | 5.692e-05 | 5.6947e-05 | 2.0171 | 3.3523e-15 | 4.2298e-11 | 4.9776 | 1.3009 | D11591.5 | 1000 | 5.5001e-08 | 5.5027e-08 | 2.0171 | 1000000000000 | 96.6318 | 0.016686 | 2.0391 | 1591.5
D2 | 1591.5 | 6Y | 245729 | 1138 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.9573e-10 | 2.7318e-14 | 2.0475 | 4.3001e-11 | 8.2816e-11 | 2.0883 | 2.197e-05 | 4.2311e-05 | 2.0883 | 5.5325e-15 | 6.9055e-11 | 4.9377 | 1.1993 | D21591.5 | 1000 | 4.3001e-08 | 8.2816e-08 | 2.0883 | 1000000000000 | 195.735 | 0.027318 | 2.0475 | 1591.5
D3 | 1591.5 | 6Y | 342214 | 1170 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 2.9223e-10 | 3.7891e-14 | 2.0483 | 7.5002e-11 | 1.1576e-10 | 2.1426 | 2.5666e-05 | 3.9614e-05 | 2.1426 | 7.6555e-15 | 9.5109e-11 | 4.9495 | 1.2171 | D31591.5 | 1000 | 7.5002e-08 | 1.1576e-07 | 2.1426 | 1000000000000 | 292.2347 | 0.037891 | 2.0483 | 1591.5
D4 | 1591.5 | 6Y | 442013 | 1123 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 3.9205e-10 | 4.9075e-14 | 2.0485 | 2.8001e-11 | 1.5173e-10 | 2.1738 | 7.1424e-06 | 3.8703e-05 | 2.1738 | 9.8514e-15 | 1.2206e-10 | 4.9815 | 1.2431 | D41591.5 | 1000 | 2.8001e-08 | 1.5173e-07 | 2.1738 | 1000000000000 | 392.049 | 0.049075 | 2.0485 | 1591.5
D5 | 1591.5 | 6Y | 538500 | 1140 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.8855e-10 | 5.9937e-14 | 2.0483 | 4.5001e-11 | 1.8721e-10 | 2.1909 | 9.2114e-06 | 3.832e-05 | 2.1909 | 1.1974e-14 | 1.4811e-10 | 5.0054 | 1.264 | D51591.5 | 1000 | 4.5001e-08 | 1.8721e-07 | 2.1909 | 1000000000000 | 488.5508 | 0.059937 | 2.0483 | 1591.5
D6 | 1591.5 | 6Y | 641458 | 1058 | 1591.5 | 19.73 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 5.9152e-10 | 7.1894e-14 | 2.0488 | -3.7001e-11 | 2.2541e-10 | 2.2017 | -6.2554e-06 | 3.8108e-05 | 2.2017 | 1.424e-14 | 1.7591e-10 | 5.0488 | 1.2814 | D61591.5 | 1000 | -3.7001e-08 | 2.2541e-07 | 2.2017 | 1000000000000 | 591.5245 | 0.071894 | 2.0488 | 1591.5
D7 | 1591.5 | 6Y | 740537 | 1164 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 6.9062e-10 | 8.2752e-14 | 2.0478 | 6.9002e-11 | 2.6236e-10 | 2.2083 | 9.9916e-06 | 3.799e-05 | 2.2083 | 1.642e-14 | 2.0267e-10 | 5.0397 | 1.2945 | D71591.5 | 1000 | 6.9002e-08 | 2.6236e-07 | 2.2083 | 1000000000000 | 690.6187 | 0.082752 | 2.0478 | 1591.5
D8 | 1591.5 | 6Y | 837056 | 1308 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 7.8715e-10 | 9.3669e-14 | 2.0476 | 2.1301e-10 | 2.9845e-10 | 2.2126 | 2.7061e-05 | 3.7916e-05 | 2.2126 | 1.8544e-14 | 2.2873e-10 | 5.0512 | 1.3048 | D81591.5 | 1000 | 2.1301e-07 | 2.9845e-07 | 2.2126 | 1000000000000 | 787.1525 | 0.093669 | 2.0476 | 1591.5
D9 | 1591.5 | 6Y | 936847 | 1426 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.8696e-10 | 1.0496e-13 | 2.0474 | 3.3101e-10 | 3.3583e-10 | 2.2158 | 3.7321e-05 | 3.7864e-05 | 2.2158 | 2.0739e-14 | 2.5568e-10 | 5.061 | 1.3135 | D91591.5 | 1000 | 3.3101e-07 | 3.3583e-07 | 2.2158 | 1000000000000 | 886.9588 | 0.10496 | 2.0474 | 1591.5
D10 | 1591.5 | 6Y | 1033147 | 1534 | 1591.5 | 19.74 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.8327e-10 | 1.1587e-13 | 2.0473 | 4.3901e-10 | 3.7195e-10 | 2.218 | 4.4649e-05 | 3.7829e-05 | 2.218 | 2.2858e-14 | 2.8168e-10 | 5.0689 | 1.3204 | D101591.5 | 1000 | 4.3901e-07 | 3.7195e-07 | 2.218 | 1000000000000 | 983.2735 | 0.11587 | 2.0473 | 1591.5
coax zero | 1591.5 | 6Y | 0 | 17 | 1591.4 | 19.84 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 0 | 3.2281e-15 | 2.0895 | 1.7e-11 | 3.2277e-11 | 2.0895 | 1.2568e-16 | 0 |  |  |  | #DIV/0! | #DIV/0! | coax zero1591.5 | 1000 | 1.7e-08 | 3.2277e-08 | 2.0895 | 1000000000000 | 0 | 0.0032281 | 2.0895 | 1591.5
GR1000A | 1591.5 | 6Y | 999924 | -113 | 1591.4 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0001e-09 | 1.2079e-13 | 2.0551 | -1.13e-10 | 3.7747e-10 | 2.2232 | 2.2127e-14 | 2.727e-10 |  |  |  | #DIV/0! | #DIV/0! | GR1000A1591.5 | 1000 | -1.13e-07 | 3.7747e-07 | 2.2232 | 1000000000000 | 1000.0771 | 0.12079 | 2.0551 | 1591.5
GR1000B | 1591.5 | 6Y | 999907 | -83 | 1591.4 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0001e-09 | 1.2079e-13 | 2.0551 | -8.3002e-11 | 3.7746e-10 | 2.2232 | 2.2127e-14 | 2.7269e-10 |  |  |  | #DIV/0! | #DIV/0! | GR1000B1591.5 | 1000 | -8.3002e-08 | 3.7746e-07 | 2.2232 | 1000000000000 | 1000.0601 | 0.12079 | 2.0551 | 1591.5
E1 | 1591.5 | 6Y | 60487 | 1341 | 1591.7 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0465e-11 | 1.2853e-12 | 2.5705 | 2.4601e-10 | 4.4313e-11 | 2.0219 | 0.0023509 | 0.00048049 | 2.0217 | 1.4566e-15 | 1.9036e-11 | 882.4024 | 2.3279 | E11591.5 | 1000 | 2.4601e-07 | 4.4313e-08 | 2.0219 | 1000000000000 | 10.4646 | 1.2853 | 2.5705 | 1591.5
E2 | 1591.5 | 6Y | 71300 | 1391 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 2.1279e-11 | 1.2853e-12 | 2.5704 | 2.9601e-10 | 4.4712e-11 | 2.0195 | 0.0013911 | 0.00021955 | 2.0131 | 1.6945e-15 | 2.1956e-11 | 758.5036 | 2.0364 | E21591.5 | 1000 | 2.9601e-07 | 4.4712e-08 | 2.0195 | 1000000000000 | 21.2793 | 1.2853 | 2.5704 | 1591.5
E3 | 1591.5 | 6Y | 81517 | 1332 | 1591.7 | 19.6 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 3.1498e-11 | 1.2853e-12 | 2.5704 | 2.3701e-10 | 4.5342e-11 | 2.0165 | 0.00075247 | 0.00014576 | 2.0137 | 1.9193e-15 | 2.4715e-11 | 669.6585 | 1.8346 | E31591.5 | 1000 | 2.3701e-07 | 4.5342e-08 | 2.0165 | 1000000000000 | 31.4978 | 1.2853 | 2.5704 | 1591.5
E4 | 1591.5 | 6Y | 91813 | 1324 | 1591.7 | 19.6 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 4.1795e-11 | 1.2853e-12 | 2.5704 | 2.2901e-10 | 4.6228e-11 | 2.0133 | 0.00054794 | 0.00011131 | 2.0119 | 2.1458e-15 | 2.7496e-11 | 598.9583 | 1.6813 | E41591.5 | 1000 | 2.2901e-07 | 4.6228e-08 | 2.0133 | 1000000000000 | 41.7954 | 1.2853 | 2.5704 | 1591.5
E5 | 1591.5 | 6Y | 101710 | 1334 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 5.1694e-11 | 1.2853e-12 | 2.5703 | 2.3901e-10 | 4.7319e-11 | 2.011 | 0.00046236 | 9.1937e-05 | 2.01 | 2.3636e-15 | 3.0169e-11 | 543.7732 | 1.5685 | E51591.5 | 1000 | 2.3901e-07 | 4.7319e-08 | 2.011 | 1000000000000 | 51.6939 | 1.2853 | 2.5703 | 1591.5
E6 | 1591.5 | 6Y | 111557 | 1338 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 6.1542e-11 | 1.2853e-12 | 2.5703 | 2.4301e-10 | 4.8633e-11 | 2.0097 | 0.00039487 | 7.9263e-05 | 2.0091 | 2.5803e-15 | 3.2828e-11 | 498.1111 | 1.4814 | E61591.5 | 1000 | 2.4301e-07 | 4.8633e-08 | 2.0097 | 1000000000000 | 61.5424 | 1.2853 | 2.5703 | 1591.5
E7 | 1591.5 | 6Y | 120990 | 1319 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 7.0977e-11 | 1.2853e-12 | 2.5703 | 2.2401e-10 | 5.01e-11 | 2.0098 | 0.00031561 | 7.0715e-05 | 2.0094 | 2.7878e-15 | 3.5376e-11 | 461.0249 | 1.4162 | E71591.5 | 1000 | 2.2401e-07 | 5.01e-08 | 2.0098 | 1000000000000 | 70.9769 | 1.2853 | 2.5703 | 1591.5
E8 | 1591.5 | 6Y | 130672 | 1315 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.066e-11 | 1.2853e-12 | 2.5702 | 2.2001e-10 | 5.1806e-11 | 2.0114 | 0.00027276 | 6.4311e-05 | 2.0111 | 3.0009e-15 | 3.799e-11 | 428.2947 | 1.3637 | E81591.5 | 1000 | 2.2001e-07 | 5.1806e-08 | 2.0114 | 1000000000000 | 80.6603 | 1.2853 | 2.5702 | 1591.5
E9 | 1591.5 | 6Y | 140694 | 1310 | 1591.7 | 19.61 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 9.0684e-11 | 1.2853e-12 | 2.5702 | 2.1501e-10 | 5.3775e-11 | 2.0146 | 0.0002371 | 5.9353e-05 | 2.0144 | 3.2214e-15 | 4.0697e-11 | 398.9748 | 1.3214 | E91591.5 | 1000 | 2.1501e-07 | 5.3775e-08 | 2.0146 | 1000000000000 | 90.6839 | 1.2853 | 2.5702 | 1591.5
E10 | 1591.5 | 6Y | 150067 | 1304 | 1591.7 | 19.62 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 1.0006e-10 | 1.2853e-12 | 2.5701 | 2.0901e-10 | 5.5788e-11 | 2.0188 | 0.00020889 | 5.5793e-05 | 2.0186 | 3.4276e-15 | 4.3228e-11 | 374.9676 | 1.2905 | E101591.5 | 1000 | 2.0901e-07 | 5.5788e-08 | 2.0188 | 1000000000000 | 100.0583 | 1.2853 | 2.5701 | 1591.5
CuZero | 1591.5 | 6Y | 888 | 15 | 1591.6 | 19.75 | 0.5 | 10 | 0 | 7e-05 | 10 | 0 | 0.0001 | 10 | 8.8814e-13 | 3.2291e-15 | 2.0892 | 1.5e-11 | 3.2276e-11 | 2.0894 | 1.452e-16 | 2.9399e-12 |  |  |  | #DIV/0! | #DIV/0! | CuZero1591.5 | 1000 | 1.5e-08 | 3.2276e-08 | 2.0894 | 1000000000000 | 0.88814 | 0.0032291 | 2.0892 | 1591.5

## AH2700_data
Capacitor description: | AH2700A Zero with AH2 Leads
2020-10-12 | 15:58
-2.1e-07 | 3.05e-05 | 1591.5 | 15
-2e-07 | 2.9e-05 | 1591.5 | 15
-2e-07 | 2.45e-05 | 1591.5 | 15
-2.1e-07 | 2.75e-05 | 1591.5 | 15
-2e-07 | 2.6e-05 | 1591.5 | 15
1.1e-07 | 2.9e-05 | 1591.5 | 15
-1e-07 | 2.9e-05 | 1591.5 | 15
1e-07 | 2.75e-05 | 1591.5 | 15
-2e-07 | 2.9e-05 | 1591.5 | 15
1e-07 | 2.75e-05 | 1591.5 | 15
-3.5e-07 | 2.3e-05 | 1591.5 | 15
-3.5e-07 | 2.75e-05 | 1591.5 | 15
-1e-07 | 2.9e-05 | 1591.5 | 15
-1e-07 | 2.75e-05 | 1591.5 | 15
-1e-07 | 2.45e-05 | 1591.5 | 15
1e-07 | 2.6e-05 | 1591.5 | 15
-2.1e-07 | 2.75e-05 | 1591.5 | 15
-2.1e-07 | 2.9e-05 | 1591.5 | 15
-2.1e-07 | 2.6e-05 | 1591.5 | 15
-1e-07 | 2.45e-05 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
-1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06
end | of | the | test
Capacitor description: | 1AH11A measure
2020-10-12 | 16:07
10 | -0.0001232 | 1591.5 | 15
10 | -0.000119 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001262 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001202 | 1591.5 | 15
10 | -0.0001202 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001217 | 1591.5 | 15
10 | -0.0001262 | 1591.5 | 15
10 | -0.0001232 | 1591.5 | 15
10 | -0.0001247 | 1591.5 | 15
10 | -0.0001187 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10 | 1.8e-07 | -0.00012269 | 2.05e-06
end | of | the | test
Capacitor description: | 1AH11B measure
2020-10-12 | 16:10
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -5.86e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -7.06e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
10 | -6.46e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.76e-05 | 1591.5 | 15
10 | -6.61e-05 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10 | 2.65e-07 | -6.61e-05 | 2.22e-06
end | of | the | test
Capacitor description: | 1AH11C measure
2020-10-12 | 16:13
Capacitor description: | 1AH11C measure
2020-10-12 | 16:14
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001635 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001634 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001636 | 1591.5 | 15
99.9998 | 0.001637 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
99.9998 | 0.001639 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06
end | of | the | test
Capacitor description: | 1AH11D measure
2020-10-12 | 16:17
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002085 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002077 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002083 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.00208 | 1591.5 | 15
99.9998 | 0.002085 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002079 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
99.9998 | 0.002082 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06
end | of | the | test
Capacitor description: | GR1000A measure
2020-10-12 | 16:22
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17947 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17939 | 1591.5 | 1.5
1000.0317 | 0.1794 | 1591.5 | 1.5
1000.0317 | 0.17937 | 1591.5 | 1.5
1000.0317 | 0.17925 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17935 | 1591.5 | 1.5
1000.0317 | 0.17938 | 1591.5 | 1.5
1000.0317 | 0.1794 | 1591.5 | 1.5
1000.0317 | 0.17949 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17937 | 1591.5 | 1.5
1000.0317 | 0.17934 | 1591.5 | 1.5
1000.0317 | 0.17931 | 1591.5 | 1.5
1000.0317 | 0.17935 | 1591.5 | 1.5
1000.0317 | 0.17933 | 1591.5 | 1.5
1000.0317 | 0.17932 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05
end | of | the | test
Capacitor description: | GR1000B measure
2020-10-12 | 16:28
1000.0164 | 0.18445 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.18445 | 1591.5 | 1.5
1000.0164 | 0.18455 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.18443 | 1591.5 | 1.5
1000.0164 | 0.18449 | 1591.5 | 1.5
1000.0164 | 0.18443 | 1591.5 | 1.5
1000.0164 | 0.1844 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.1845 | 1591.5 | 1.5
1000.0164 | 0.1844 | 1591.5 | 1.5
1000.0164 | 0.18451 | 1591.5 | 1.5
1000.0164 | 0.18441 | 1591.5 | 1.5
1000.0164 | 0.18454 | 1591.5 | 1.5
1000.0164 | 0.18435 | 1591.5 | 1.5
1000.0164 | 0.18448 | 1591.5 | 1.5
1000.0164 | 0.18435 | 1591.5 | 1.5
1000.0164 | 0.18438 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05
end | of | the | test
Capacitor description: | UB1000 measure 20.07deg, 1019.93 mbar, 57.4%RH
2020-10-12 | 16:33
1000.1471 | 0.14212 | 1591.5 | 1.5
1000.1471 | 0.14209 | 1591.5 | 1.5
1000.1471 | 0.14217 | 1591.5 | 1.5
1000.1471 | 0.14217 | 1591.5 | 1.5
1000.1471 | 0.14208 | 1591.5 | 1.5
1000.1471 | 0.14199 | 1591.5 | 1.5
1000.1471 | 0.14211 | 1591.5 | 1.5
1000.1471 | 0.14202 | 1591.5 | 1.5
1000.1471 | 0.14205 | 1591.5 | 1.5
1000.1471 | 0.14189 | 1591.5 | 1.5
1000.1471 | 0.14199 | 1591.5 | 1.5
1000.1471 | 0.14195 | 1591.5 | 1.5
1000.1471 | 0.142 | 1591.5 | 1.5
1000.1471 | 0.14193 | 1591.5 | 1.5
1000.1471 | 0.14197 | 1591.5 | 1.5
1000.1471 | 0.14194 | 1591.5 | 1.5
1000.1471 | 0.14185 | 1591.5 | 1.5
1000.1471 | 0.14192 | 1591.5 | 1.5
1000.1471 | 0.1419 | 1591.5 | 1.5
1000.1471 | 0.14184 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1471 | 5.89e-06 | 0.142 | 9.76e-05
end | of | the | test
Capacitor description: | Zero point of GR Decade Capacitor
2020-10-12 | 16:40
49.221 | 1.3388 | 1591.5 | 15
49.221 | 1.3388 | 1591.5 | 15
49.2213 | 1.3365 | 1591.5 | 15
49.2218 | 1.3297 | 1591.5 | 15
49.221 | 1.3394 | 1591.5 | 15
49.222 | 1.3274 | 1591.5 | 15
49.222 | 1.3271 | 1591.5 | 15
49.222 | 1.3268 | 1591.5 | 15
49.222 | 1.3261 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.222 | 1.3262 | 1591.5 | 15
49.2219 | 1.3288 | 1591.5 | 15
49.2219 | 1.3291 | 1591.5 | 15
49.2219 | 1.3293 | 1591.5 | 15
49.222 | 1.3281 | 1591.5 | 15
49.2215 | 1.3358 | 1591.5 | 15
49.2207 | 1.3411 | 1591.5 | 15
49.2212 | 1.3381 | 1591.5 | 15
49.2214 | 1.3351 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
49.2216 | 0.00043741 | 1.3317 | 0.0053142
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 10 measure
2020-10-12 | 16:44
59.6543 | 1.3374 | 1591.5 | 15
59.6543 | 1.3374 | 1591.5 | 15
59.6544 | 1.3364 | 1591.5 | 15
59.6544 | 1.3364 | 1591.5 | 15
59.6542 | 1.3395 | 1591.5 | 15
59.6542 | 1.3389 | 1591.5 | 15
59.6544 | 1.337 | 1591.5 | 15
59.6542 | 1.3393 | 1591.5 | 15
59.6541 | 1.3409 | 1591.5 | 15
59.6541 | 1.3412 | 1591.5 | 15
59.6543 | 1.3383 | 1591.5 | 15
59.6543 | 1.3382 | 1591.5 | 15
59.6544 | 1.3357 | 1591.5 | 15
59.6544 | 1.3357 | 1591.5 | 15
59.6545 | 1.3349 | 1591.5 | 15
59.6545 | 1.3353 | 1591.5 | 15
59.6545 | 1.3345 | 1591.5 | 15
59.6544 | 1.3371 | 1591.5 | 15
59.6544 | 1.3365 | 1591.5 | 15
59.6539 | 1.3432 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
59.6543 | 0.00014858 | 1.3377 | 0.0022088
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 20 measure
2020-10-12 | 16:55
70.2236 | 1.3732 | 1591.5 | 15
70.2235 | 1.3685 | 1591.5 | 15
70.2235 | 1.3704 | 1591.5 | 15
70.2233 | 1.3735 | 1591.5 | 15
70.2235 | 1.3686 | 1591.5 | 15
70.2234 | 1.3679 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2234 | 1.3677 | 1591.5 | 15
70.2234 | 1.3662 | 1591.5 | 15
70.2234 | 1.3667 | 1591.5 | 15
70.2234 | 1.3664 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2234 | 1.368 | 1591.5 | 15
70.2234 | 1.3671 | 1591.5 | 15
70.2233 | 1.3679 | 1591.5 | 15
70.2234 | 1.3669 | 1591.5 | 15
70.2234 | 1.3713 | 1591.5 | 15
70.2234 | 1.3696 | 1591.5 | 15
70.2235 | 1.3698 | 1591.5 | 15
70.2234 | 1.369 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
70.2234 | 6.02e-05 | 1.3687 | 0.0020591
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 30 measure
2020-10-12 | 16:58
80.8297 | 1.3379 | 1591.5 | 15
80.8297 | 1.338 | 1591.5 | 15
80.8297 | 1.3374 | 1591.5 | 15
80.8298 | 1.3374 | 1591.5 | 15
80.8284 | 1.351 | 1591.5 | 15
80.8284 | 1.3512 | 1591.5 | 15
80.8284 | 1.3513 | 1591.5 | 15
80.8284 | 1.3512 | 1591.5 | 15
80.8284 | 1.351 | 1591.5 | 15
80.8285 | 1.351 | 1591.5 | 15
80.8285 | 1.3509 | 1591.5 | 15
80.8285 | 1.3511 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3502 | 1591.5 | 15
80.8287 | 1.3502 | 1591.5 | 15
80.8287 | 1.3501 | 1591.5 | 15
80.8287 | 1.3505 | 1591.5 | 15
80.8287 | 1.3505 | 1591.5 | 15
80.8288 | 1.3499 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
80.8288 | 0.00048247 | 1.3481 | 0.0052107
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 40 measure
2020-10-12 | 17:01
90.9233 | 1.3518 | 1591.5 | 15
90.9233 | 1.3519 | 1591.5 | 15
90.9233 | 1.3519 | 1591.5 | 15
90.9237 | 1.3498 | 1591.5 | 15
90.9237 | 1.3495 | 1591.5 | 15
90.9237 | 1.3498 | 1591.5 | 15
90.9238 | 1.3494 | 1591.5 | 15
90.9238 | 1.3495 | 1591.5 | 15
90.9238 | 1.3499 | 1591.5 | 15
90.9238 | 1.3495 | 1591.5 | 15
90.9239 | 1.3492 | 1591.5 | 15
90.9239 | 1.3495 | 1591.5 | 15
90.9239 | 1.3497 | 1591.5 | 15
90.9239 | 1.3498 | 1591.5 | 15
90.9239 | 1.3497 | 1591.5 | 15
90.9239 | 1.3498 | 1591.5 | 15
90.9239 | 1.3502 | 1591.5 | 15
90.9239 | 1.3501 | 1591.5 | 15
90.924 | 1.3498 | 1591.5 | 15
90.924 | 1.3497 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
90.9238 | 0.00021909 | 1.35 | 0.00079946
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 50 measure
2020-10-12 | 17:04
100.4765 | 1.3524 | 1591.5 | 15
100.4766 | 1.3524 | 1591.5 | 15
100.4766 | 1.3529 | 1591.5 | 15
100.4773 | 1.3434 | 1591.5 | 15
100.4774 | 1.343 | 1591.5 | 15
100.4774 | 1.3424 | 1591.5 | 15
100.4774 | 1.3426 | 1591.5 | 15
100.4775 | 1.3425 | 1591.5 | 15
100.4775 | 1.343 | 1591.5 | 15
100.4776 | 1.3409 | 1591.5 | 15
100.4776 | 1.3407 | 1591.5 | 15
100.4775 | 1.3439 | 1591.5 | 15
100.4773 | 1.3471 | 1591.5 | 15
100.4775 | 1.3443 | 1591.5 | 15
100.4776 | 1.3431 | 1591.5 | 15
100.4776 | 1.3424 | 1591.5 | 15
100.4777 | 1.3413 | 1591.5 | 15
100.4777 | 1.3413 | 1591.5 | 15
100.4777 | 1.341 | 1591.5 | 15
100.4777 | 1.3409 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100.4774 | 0.00036577 | 1.3441 | 0.0038393
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 60 measure
2020-10-12 | 17:06
110.69 | 1.346 | 1591.5 | 15
110.6902 | 1.3465 | 1591.5 | 15
110.6904 | 1.3467 | 1591.5 | 15
110.6906 | 1.3468 | 1591.5 | 15
110.6907 | 1.3476 | 1591.5 | 15
110.6909 | 1.3457 | 1591.5 | 15
110.691 | 1.345 | 1591.5 | 15
110.6911 | 1.3443 | 1591.5 | 15
110.6912 | 1.345 | 1591.5 | 15
110.6913 | 1.345 | 1591.5 | 15
110.6913 | 1.3452 | 1591.5 | 15
110.6914 | 1.3452 | 1591.5 | 15
110.6915 | 1.3451 | 1591.5 | 15
110.6915 | 1.3453 | 1591.5 | 15
110.6915 | 1.3454 | 1591.5 | 15
110.6916 | 1.3452 | 1591.5 | 15
110.6916 | 1.345 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
110.6917 | 1.3449 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
110.6912 | 0.00050256 | 1.3455 | 0.00079577
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 70 measure
2020-10-12 | 17:11
120.5074 | 1.3464 | 1591.5 | 15
120.5075 | 1.3463 | 1591.5 | 15
120.5076 | 1.346 | 1591.5 | 15
120.5077 | 1.346 | 1591.5 | 15
120.5078 | 1.3461 | 1591.5 | 15
120.5079 | 1.3461 | 1591.5 | 15
120.5079 | 1.3462 | 1591.5 | 15
120.508 | 1.3462 | 1591.5 | 15
120.5081 | 1.3462 | 1591.5 | 15
120.5081 | 1.3461 | 1591.5 | 15
120.5082 | 1.3461 | 1591.5 | 15
120.5082 | 1.3462 | 1591.5 | 15
120.5082 | 1.3464 | 1591.5 | 15
120.5083 | 1.3464 | 1591.5 | 15
120.5083 | 1.3464 | 1591.5 | 15
120.5083 | 1.3467 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3465 | 1591.5 | 15
120.5084 | 1.3464 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
120.5081 | 0.00029987 | 1.3463 | 0.00018653
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 80 measure
2020-10-12 | 17:23
129.9739 | 1.3534 | 1591.5 | 15
129.9739 | 1.3539 | 1591.5 | 15
129.9738 | 1.3542 | 1591.5 | 15
129.9739 | 1.3535 | 1591.5 | 15
129.974 | 1.3516 | 1591.5 | 15
129.974 | 1.3516 | 1591.5 | 15
129.974 | 1.3522 | 1591.5 | 15
129.974 | 1.3527 | 1591.5 | 15
129.9739 | 1.3529 | 1591.5 | 15
129.9739 | 1.3529 | 1591.5 | 15
129.9739 | 1.3532 | 1591.5 | 15
129.974 | 1.3517 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.3524 | 1591.5 | 15
129.974 | 1.3513 | 1591.5 | 15
129.974 | 1.352 | 1591.5 | 15
129.974 | 1.3523 | 1591.5 | 15
129.9741 | 1.3512 | 1591.5 | 15
129.974 | 1.3523 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
129.974 | 5.59e-05 | 1.3525 | 0.00082738
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail 90 measure
2020-10-12 | 17:25
139.7372 | 1.353 | 1591.5 | 7.5
139.7376 | 1.3527 | 1591.5 | 7.5
139.7378 | 1.3535 | 1591.5 | 7.5
139.7381 | 1.3514 | 1591.5 | 7.5
139.7382 | 1.3511 | 1591.5 | 7.5
139.7384 | 1.3512 | 1591.5 | 7.5
139.7384 | 1.3522 | 1591.5 | 7.5
139.7386 | 1.3513 | 1591.5 | 7.5
139.7387 | 1.3515 | 1591.5 | 7.5
139.7387 | 1.3521 | 1591.5 | 7.5
139.7387 | 1.3529 | 1591.5 | 7.5
139.7388 | 1.3529 | 1591.5 | 7.5
139.7389 | 1.3528 | 1591.5 | 7.5
139.739 | 1.3515 | 1591.5 | 7.5
139.739 | 1.3524 | 1591.5 | 7.5
139.739 | 1.3521 | 1591.5 | 7.5
139.7391 | 1.3515 | 1591.5 | 7.5
139.7391 | 1.3523 | 1591.5 | 7.5
139.7392 | 1.3522 | 1591.5 | 7.5
139.7392 | 1.3516 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
139.7386 | 0.0005509 | 1.3521 | 0.0006845
end | of | the | test
Capacitor description: | GR sn:4556 1pF dail '100' measure
2020-10-12 | 17:28
149.5761 | 1.356 | 1591.5 | 7.5
149.5763 | 1.3542 | 1591.5 | 7.5
149.5764 | 1.3538 | 1591.5 | 7.5
149.5764 | 1.3558 | 1591.5 | 7.5
149.5765 | 1.3557 | 1591.5 | 7.5
149.5766 | 1.3552 | 1591.5 | 7.5
149.5765 | 1.3565 | 1591.5 | 7.5
149.5766 | 1.3566 | 1591.5 | 7.5
149.5767 | 1.3559 | 1591.5 | 7.5
149.5768 | 1.3545 | 1591.5 | 7.5
149.5768 | 1.355 | 1591.5 | 7.5
149.5769 | 1.3549 | 1591.5 | 7.5
149.5769 | 1.355 | 1591.5 | 7.5
149.5768 | 1.3557 | 1591.5 | 7.5
149.5769 | 1.3561 | 1591.5 | 7.5
149.5769 | 1.3565 | 1591.5 | 7.5
149.5769 | 1.3562 | 1591.5 | 7.5
149.577 | 1.3555 | 1591.5 | 7.5
149.577 | 1.3556 | 1591.5 | 7.5
149.577 | 1.3556 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
149.5767 | 0.00025146 | 1.3555 | 0.00075452
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 1
2020-10-12 | 17:32
145.8881 | 1.6341 | 1591.5 | 7.5
145.8884 | 1.6282 | 1591.5 | 7.5
145.8882 | 1.6363 | 1591.5 | 7.5
145.8884 | 1.629 | 1591.5 | 7.5
145.8881 | 1.6392 | 1591.5 | 7.5
145.8886 | 1.623 | 1591.5 | 7.5
145.8883 | 1.6321 | 1591.5 | 7.5
145.8884 | 1.6312 | 1591.5 | 7.5
145.888 | 1.6409 | 1591.5 | 7.5
145.8885 | 1.6257 | 1591.5 | 7.5
145.8891 | 1.6075 | 1591.5 | 7.5
145.889 | 1.6085 | 1591.5 | 7.5
145.889 | 1.6094 | 1591.5 | 7.5
145.889 | 1.6081 | 1591.5 | 7.5
145.8893 | 1.5986 | 1591.5 | 7.5
145.8894 | 1.5967 | 1591.5 | 7.5
145.8892 | 1.6014 | 1591.5 | 7.5
145.8891 | 1.606 | 1591.5 | 7.5
145.8893 | 1.5992 | 1591.5 | 7.5
145.8894 | 1.5961 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
145.8887 | 0.00046896 | 1.6176 | 0.015335
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 2
2020-10-12 | 17:34
244.935 | 1.4324 | 1591.5 | 7.5
244.935 | 1.4322 | 1591.5 | 7.5
244.935 | 1.432 | 1591.5 | 7.5
244.9351 | 1.4313 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4317 | 1591.5 | 7.5
244.935 | 1.4328 | 1591.5 | 7.5
244.9348 | 1.4348 | 1591.5 | 7.5
244.9348 | 1.4349 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4347 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4355 | 1591.5 | 7.5
244.9348 | 1.4351 | 1591.5 | 7.5
244.9348 | 1.4352 | 1591.5 | 7.5
244.9348 | 1.435 | 1591.5 | 7.5
244.9348 | 1.4356 | 1591.5 | 7.5
244.9348 | 1.435 | 1591.5 | 7.5
244.9348 | 1.4356 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
244.9349 | 0.00010611 | 1.4339 | 0.0016017
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 3
2020-10-12 | 17:37
341.4798 | 1.517 | 1591.5 | 3
341.4799 | 1.516 | 1591.5 | 3
341.4801 | 1.5131 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.4798 | 1.5161 | 1591.5 | 3
341.48 | 1.5135 | 1591.5 | 3
341.48 | 1.5134 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.48 | 1.5133 | 1591.5 | 3
341.4798 | 1.5159 | 1591.5 | 3
341.48 | 1.5132 | 1591.5 | 3
341.48 | 1.5134 | 1591.5 | 3
341.48 | 1.5131 | 1591.5 | 3
341.48 | 1.5138 | 1591.5 | 3
341.4799 | 1.515 | 1591.5 | 3
341.4789 | 1.523 | 1591.5 | 3
341.479 | 1.5229 | 1591.5 | 3
341.4796 | 1.5184 | 1591.5 | 3
341.4797 | 1.5176 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
341.4798 | 0.00032198 | 1.5154 | 0.0029831
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 4
2020-10-12 | 17:39
441.2625 | 1.5404 | 1591.5 | 3
441.2625 | 1.5404 | 1591.5 | 3
441.2627 | 1.5386 | 1591.5 | 3
441.2626 | 1.5398 | 1591.5 | 3
441.2626 | 1.5397 | 1591.5 | 3
441.2626 | 1.5395 | 1591.5 | 3
441.2626 | 1.5394 | 1591.5 | 3
441.2627 | 1.5391 | 1591.5 | 3
441.2627 | 1.5389 | 1591.5 | 3
441.2626 | 1.5395 | 1591.5 | 3
441.2626 | 1.5397 | 1591.5 | 3
441.2623 | 1.5422 | 1591.5 | 3
441.2625 | 1.541 | 1591.5 | 3
441.2623 | 1.5421 | 1591.5 | 3
441.2626 | 1.5399 | 1591.5 | 3
441.2622 | 1.5429 | 1591.5 | 3
441.2623 | 1.5423 | 1591.5 | 3
441.2623 | 1.5425 | 1591.5 | 3
441.2624 | 1.5412 | 1591.5 | 3
441.2624 | 1.5413 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
441.2625 | 0.0001467 | 1.5405 | 0.0012842
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 5
2020-10-12 | 17:41
537.7662 | 1.6209 | 1591.5 | 3
537.766 | 1.622 | 1591.5 | 3
537.7661 | 1.6214 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.621 | 1591.5 | 3
537.7662 | 1.6204 | 1591.5 | 3
537.7662 | 1.6202 | 1591.5 | 3
537.7662 | 1.6199 | 1591.5 | 3
537.7661 | 1.6211 | 1591.5 | 3
537.7661 | 1.6209 | 1591.5 | 3
537.7661 | 1.6211 | 1591.5 | 3
537.7661 | 1.6214 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.7661 | 1.6212 | 1591.5 | 3
537.766 | 1.6217 | 1591.5 | 3
537.7661 | 1.6215 | 1591.5 | 3
537.7663 | 1.6191 | 1591.5 | 3
537.7662 | 1.6202 | 1591.5 | 3
537.7662 | 1.6203 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
537.7661 | 6.53e-05 | 1.6209 | 0.0006773
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 6
2020-10-12 | 17:43
640.7043 | 1.6659 | 1591.5 | 3
640.7043 | 1.6662 | 1591.5 | 3
640.7042 | 1.6663 | 1591.5 | 3
640.7041 | 1.6672 | 1591.5 | 3
640.7042 | 1.6661 | 1591.5 | 3
640.7042 | 1.6664 | 1591.5 | 3
640.7042 | 1.6659 | 1591.5 | 3
640.7042 | 1.6662 | 1591.5 | 3
640.7042 | 1.666 | 1591.5 | 3
640.7042 | 1.6662 | 1591.5 | 3
640.7041 | 1.6665 | 1591.5 | 3
640.7041 | 1.6666 | 1591.5 | 3
640.7041 | 1.6667 | 1591.5 | 3
640.7041 | 1.6666 | 1591.5 | 3
640.7041 | 1.6665 | 1591.5 | 3
640.704 | 1.6665 | 1591.5 | 3
640.704 | 1.6666 | 1591.5 | 3
640.704 | 1.6668 | 1591.5 | 3
640.7039 | 1.6674 | 1591.5 | 3
640.7039 | 1.6672 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
640.7041 | 0.00010485 | 1.6665 | 0.00040726
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 7
2020-10-12 | 17:45
739.8079 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6858 | 1591.5 | 1.5
739.808 | 1.686 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.808 | 1.6862 | 1591.5 | 1.5
739.808 | 1.6861 | 1591.5 | 1.5
739.8081 | 1.6851 | 1591.5 | 1.5
739.8081 | 1.6852 | 1591.5 | 1.5
739.8081 | 1.6849 | 1591.5 | 1.5
739.8081 | 1.6853 | 1591.5 | 1.5
739.8081 | 1.6852 | 1591.5 | 1.5
739.8074 | 1.6895 | 1591.5 | 1.5
739.8074 | 1.6893 | 1591.5 | 1.5
739.8075 | 1.6893 | 1591.5 | 1.5
739.8074 | 1.6894 | 1591.5 | 1.5
739.8076 | 1.6888 | 1591.5 | 1.5
739.8077 | 1.688 | 1591.5 | 1.5
739.8077 | 1.6881 | 1591.5 | 1.5
739.8076 | 1.6888 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
739.8078 | 0.00025568 | 1.687 | 0.0016552
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 8
2020-10-12 | 17:47
836.3519 | 1.7676 | 1591.5 | 1.5
836.3519 | 1.7677 | 1591.5 | 1.5
836.352 | 1.7677 | 1591.5 | 1.5
836.352 | 1.7674 | 1591.5 | 1.5
836.3521 | 1.767 | 1591.5 | 1.5
836.3521 | 1.7675 | 1591.5 | 1.5
836.3521 | 1.7673 | 1591.5 | 1.5
836.3521 | 1.7675 | 1591.5 | 1.5
836.352 | 1.768 | 1591.5 | 1.5
836.352 | 1.7682 | 1591.5 | 1.5
836.3521 | 1.7678 | 1591.5 | 1.5
836.352 | 1.7682 | 1591.5 | 1.5
836.352 | 1.7684 | 1591.5 | 1.5
836.3522 | 1.7676 | 1591.5 | 1.5
836.3521 | 1.7677 | 1591.5 | 1.5
836.3521 | 1.7677 | 1591.5 | 1.5
836.3522 | 1.7674 | 1591.5 | 1.5
836.3522 | 1.7674 | 1591.5 | 1.5
836.3523 | 1.7671 | 1591.5 | 1.5
836.3523 | 1.7671 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
836.3521 | 0.00010751 | 1.7676 | 0.00037086
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail 9
2020-10-12 | 17:50
936.1407 | 1.7891 | 1591.5 | 1.5
936.1407 | 1.7896 | 1591.5 | 1.5
936.1409 | 1.7891 | 1591.5 | 1.5
936.141 | 1.7888 | 1591.5 | 1.5
936.141 | 1.7888 | 1591.5 | 1.5
936.1411 | 1.7889 | 1591.5 | 1.5
936.1416 | 1.7853 | 1591.5 | 1.5
936.1412 | 1.7889 | 1591.5 | 1.5
936.1412 | 1.7886 | 1591.5 | 1.5
936.1413 | 1.7885 | 1591.5 | 1.5
936.1413 | 1.7885 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7883 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7884 | 1591.5 | 1.5
936.1414 | 1.7888 | 1591.5 | 1.5
936.1415 | 1.7886 | 1591.5 | 1.5
936.1414 | 1.789 | 1591.5 | 1.5
936.1415 | 1.7889 | 1591.5 | 1.5
936.1414 | 1.7894 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
936.1413 | 0.00025167 | 1.7886 | 0.00083413
end | of | the | test
Capacitor description: | GR sn:4556 100pF dail X
2020-10-12 | 17:52
1032.4676 | 1.8824 | 1591.5 | 1.5
1032.4676 | 1.8825 | 1591.5 | 1.5
1032.4678 | 1.8811 | 1591.5 | 1.5
1032.4677 | 1.881 | 1591.5 | 1.5
1032.4677 | 1.8814 | 1591.5 | 1.5
1032.4678 | 1.8803 | 1591.5 | 1.5
1032.4677 | 1.8806 | 1591.5 | 1.5
1032.4677 | 1.8807 | 1591.5 | 1.5
1032.4678 | 1.8809 | 1591.5 | 1.5
1032.4678 | 1.8808 | 1591.5 | 1.5
1032.4678 | 1.8808 | 1591.5 | 1.5
1032.4678 | 1.8805 | 1591.5 | 1.5
1032.4678 | 1.8812 | 1591.5 | 1.5
1032.4677 | 1.8814 | 1591.5 | 1.5
1032.4677 | 1.881 | 1591.5 | 1.5
1032.4678 | 1.8809 | 1591.5 | 1.5
1032.4677 | 1.8803 | 1591.5 | 1.5
1032.4681 | 1.8779 | 1591.5 | 1.5
1032.4681 | 1.8773 | 1591.5 | 1.5
1032.4681 | 1.8769 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1032.4678 | 0.00013984 | 1.8805 | 0.0014275
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 1
2020-10-12 | 17:55
1044.2428 | 1.8791 | 1591.5 | 1.5
1044.2429 | 1.8798 | 1591.5 | 1.5
1044.2429 | 1.8793 | 1591.5 | 1.5
1044.2429 | 1.8797 | 1591.5 | 1.5
1044.243 | 1.8804 | 1591.5 | 1.5
1044.243 | 1.8802 | 1591.5 | 1.5
1044.2431 | 1.8804 | 1591.5 | 1.5
1044.243 | 1.8799 | 1591.5 | 1.5
1044.2431 | 1.8802 | 1591.5 | 1.5
1044.2431 | 1.8804 | 1591.5 | 1.5
1044.2432 | 1.8804 | 1591.5 | 1.5
1044.2432 | 1.8803 | 1591.5 | 1.5
1044.2431 | 1.8795 | 1591.5 | 1.5
1044.2432 | 1.8801 | 1591.5 | 1.5
1044.2432 | 1.8799 | 1591.5 | 1.5
1044.2432 | 1.88 | 1591.5 | 1.5
1044.2434 | 1.8814 | 1591.5 | 1.5
1044.2434 | 1.8812 | 1591.5 | 1.5
1044.2435 | 1.8826 | 1591.5 | 1.5
1044.2435 | 1.882 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1044.2431 | 0.00020482 | 1.8803 | 0.00085081
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 2
2020-10-12 | 17:57
2046.6228 | 2.202 | 1591.5 | 0.75
2046.6228 | 2.202 | 1591.5 | 0.75
2046.6229 | 2.202 | 1591.5 | 0.75
2046.6229 | 2.2016 | 1591.5 | 0.75
2046.6229 | 2.2018 | 1591.5 | 0.75
2046.6229 | 2.2023 | 1591.5 | 0.75
2046.6229 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.2019 | 1591.5 | 0.75
2046.6229 | 2.2019 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.6229 | 2.2018 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.623 | 2.2019 | 1591.5 | 0.75
2046.623 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.202 | 1591.5 | 0.75
2046.623 | 2.2016 | 1591.5 | 0.75
2046.623 | 2.202 | 1591.5 | 0.75
2046.6231 | 2.2017 | 1591.5 | 0.75
2046.623 | 2.2018 | 1591.5 | 0.75
2046.6231 | 2.2017 | 1591.5 | 0.75
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
2046.623 | 7.55e-05 | 2.2018 | 0.00016795
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 3
2020-10-12 | 17:59
3041.324 | 2.7767 | 1591.5 | 0.25
3041.3239 | 2.775 | 1591.5 | 0.25
3041.3241 | 2.7756 | 1591.5 | 0.25
3041.3241 | 2.7771 | 1591.5 | 0.25
3041.3241 | 2.7756 | 1591.5 | 0.25
3041.3241 | 2.7759 | 1591.5 | 0.25
3041.3242 | 2.775 | 1591.5 | 0.25
3041.3243 | 2.7754 | 1591.5 | 0.25
3041.3242 | 2.7756 | 1591.5 | 0.25
3041.3242 | 2.7756 | 1591.5 | 0.25
3041.3243 | 2.7758 | 1591.5 | 0.25
3041.3243 | 2.7756 | 1591.5 | 0.25
3041.3243 | 2.7743 | 1591.5 | 0.25
3041.3245 | 2.7765 | 1591.5 | 0.25
3041.3244 | 2.7756 | 1591.5 | 0.25
3041.3245 | 2.7763 | 1591.5 | 0.25
3041.3246 | 2.7766 | 1591.5 | 0.25
3041.3245 | 2.7753 | 1591.5 | 0.25
3041.3246 | 2.7759 | 1591.5 | 0.25
3041.3246 | 2.7772 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
3041.3243 | 0.0001986 | 2.7758 | 0.00070788
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 4
2020-10-12 | 18:01
4043.9172 | 3.217 | 1591.5 | 0.25
4043.9172 | 3.2155 | 1591.5 | 0.25
4043.9173 | 3.2154 | 1591.5 | 0.25
4043.9173 | 3.216 | 1591.5 | 0.25
4043.9172 | 3.2173 | 1591.5 | 0.25
4043.9172 | 3.2165 | 1591.5 | 0.25
4043.9174 | 3.2151 | 1591.5 | 0.25
4043.9173 | 3.2155 | 1591.5 | 0.25
4043.9175 | 3.2167 | 1591.5 | 0.25
4043.9175 | 3.2156 | 1591.5 | 0.25
4043.9176 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2163 | 1591.5 | 0.25
4043.9175 | 3.215 | 1591.5 | 0.25
4043.9177 | 3.2162 | 1591.5 | 0.25
4043.9177 | 3.2167 | 1591.5 | 0.25
4043.9175 | 3.2159 | 1591.5 | 0.25
4043.9177 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2158 | 1591.5 | 0.25
4043.9177 | 3.2172 | 1591.5 | 0.25
4043.9177 | 3.2156 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
4043.9175 | 0.00020302 | 3.2162 | 0.00073231
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 5
2020-10-12 | 18:03
5047.278 | 3.7849 | 1591.5 | 0.25
5047.2783 | 3.7828 | 1591.5 | 0.25
5047.2782 | 3.7828 | 1591.5 | 0.25
5047.2783 | 3.786 | 1591.5 | 0.25
5047.2785 | 3.7852 | 1591.5 | 0.25
5047.2784 | 3.7849 | 1591.5 | 0.25
5047.2787 | 3.783 | 1591.5 | 0.25
5047.2787 | 3.785 | 1591.5 | 0.25
5047.2787 | 3.7848 | 1591.5 | 0.25
5047.2788 | 3.7848 | 1591.5 | 0.25
5047.2788 | 3.7841 | 1591.5 | 0.25
5047.2787 | 3.7827 | 1591.5 | 0.25
5047.2789 | 3.7848 | 1591.5 | 0.25
5047.2787 | 3.7848 | 1591.5 | 0.25
5047.279 | 3.7856 | 1591.5 | 0.25
5047.2789 | 3.7844 | 1591.5 | 0.25
5047.2792 | 3.7849 | 1591.5 | 0.25
5047.2792 | 3.7858 | 1591.5 | 0.25
5047.2793 | 3.7845 | 1591.5 | 0.25
5047.2791 | 3.7856 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
5047.2787 | 0.00034511 | 3.7846 | 0.00098189
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 6
2020-10-12 | 18:05
6041.457 | 4.3639 | 1591.5 | 0.25
6041.457 | 4.3627 | 1591.5 | 0.25
6041.4575 | 4.3638 | 1591.5 | 0.25
6041.4574 | 4.3633 | 1591.5 | 0.25
6041.4573 | 4.3641 | 1591.5 | 0.25
6041.4575 | 4.3642 | 1591.5 | 0.25
6041.4573 | 4.3651 | 1591.5 | 0.25
6041.4577 | 4.3633 | 1591.5 | 0.25
6041.4576 | 4.3622 | 1591.5 | 0.25
6041.4577 | 4.3636 | 1591.5 | 0.25
6041.4578 | 4.3638 | 1591.5 | 0.25
6041.4577 | 4.3653 | 1591.5 | 0.25
6041.4578 | 4.3661 | 1591.5 | 0.25
6041.4579 | 4.3618 | 1591.5 | 0.25
6041.458 | 4.3634 | 1591.5 | 0.25
6041.4578 | 4.3644 | 1591.5 | 0.25
6041.458 | 4.3636 | 1591.5 | 0.25
6041.4583 | 4.3628 | 1591.5 | 0.25
6041.4582 | 4.3627 | 1591.5 | 0.25
6041.4583 | 4.3618 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
6041.4577 | 0.00036621 | 4.3636 | 0.0010893
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 7
2020-10-12 | 18:08
7043.973 | 4.8055 | 1591.5 | 0.25
7043.9727 | 4.806 | 1591.5 | 0.25
7043.973 | 4.8041 | 1591.5 | 0.25
7043.973 | 4.8091 | 1591.5 | 0.25
7043.9729 | 4.8074 | 1591.5 | 0.25
7043.973 | 4.8036 | 1591.5 | 0.25
7043.9733 | 4.8065 | 1591.5 | 0.25
7043.973 | 4.8035 | 1591.5 | 0.25
7043.9731 | 4.8027 | 1591.5 | 0.25
7043.9735 | 4.8041 | 1591.5 | 0.25
7043.9733 | 4.8057 | 1591.5 | 0.25
7043.9735 | 4.8035 | 1591.5 | 0.25
7043.9733 | 4.8057 | 1591.5 | 0.25
7043.9735 | 4.8037 | 1591.5 | 0.25
7043.9735 | 4.8045 | 1591.5 | 0.25
7043.9738 | 4.805 | 1591.5 | 0.25
7043.9738 | 4.8067 | 1591.5 | 0.25
7043.9738 | 4.8046 | 1591.5 | 0.25
7043.9739 | 4.8079 | 1591.5 | 0.25
7043.9741 | 4.8051 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
7043.9733 | 0.00037834 | 4.8052 | 0.0016222
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 8
2020-10-12 | 18:10
8037.9912 | 5.5 | 1591.5 | 0.25
8037.991 | 5.506 | 1591.5 | 0.25
8037.9916 | 5.49 | 1591.5 | 0.25
8037.9915 | 5.494 | 1591.5 | 0.25
8037.9916 | 5.497 | 1591.5 | 0.25
8037.9911 | 5.504 | 1591.5 | 0.25
8037.9915 | 5.502 | 1591.5 | 0.25
8037.9918 | 5.502 | 1591.5 | 0.25
8037.9909 | 5.512 | 1591.5 | 0.25
8037.9916 | 5.505 | 1591.5 | 0.25
8037.992 | 5.501 | 1591.5 | 0.25
8037.9916 | 5.519 | 1591.5 | 0.25
8037.9922 | 5.505 | 1591.5 | 0.25
8037.9919 | 5.498 | 1591.5 | 0.25
8037.9918 | 5.51 | 1591.5 | 0.25
8037.9923 | 5.504 | 1591.5 | 0.25
8037.992 | 5.514 | 1591.5 | 0.25
8037.9911 | 5.525 | 1591.5 | 0.25
8037.9921 | 5.511 | 1591.5 | 0.25
8037.9922 | 5.51 | 1591.5 | 0.25
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
8037.9916 | 0.00041509 | 5.5054 | 0.0081515
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail 9
2020-10-12 | 18:12
9040.2661 | 5.957 | 1591.5 | 0.1
9040.2664 | 5.948 | 1591.5 | 0.1
9040.2657 | 5.947 | 1591.5 | 0.1
9040.2659 | 5.947 | 1591.5 | 0.1
9040.2662 | 5.954 | 1591.5 | 0.1
9040.266 | 5.95 | 1591.5 | 0.1
9040.2663 | 5.945 | 1591.5 | 0.1
9040.2665 | 5.95 | 1591.5 | 0.1
9040.2656 | 5.954 | 1591.5 | 0.1
9040.2672 | 5.952 | 1591.5 | 0.1
9040.2667 | 5.952 | 1591.5 | 0.1
9040.2667 | 5.942 | 1591.5 | 0.1
9040.2669 | 5.949 | 1591.5 | 0.1
9040.2666 | 5.94 | 1591.5 | 0.1
9040.2667 | 5.95 | 1591.5 | 0.1
9040.2665 | 5.952 | 1591.5 | 0.1
9040.2662 | 5.949 | 1591.5 | 0.1
9040.2667 | 5.946 | 1591.5 | 0.1
9040.2668 | 5.949 | 1591.5 | 0.1
9040.2672 | 5.956 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
9040.2665 | 0.00043352 | 5.9494 | 0.0042009
end | of | the | test
Capacitor description: | GR sn:4556 1000pF Dail X
2020-10-12 | 18:14
10034.2262 | 6.724 | 1591.5 | 0.1
10034.2263 | 6.714 | 1591.5 | 0.1
10034.2256 | 6.722 | 1591.5 | 0.1
10034.2259 | 6.714 | 1591.5 | 0.1
10034.2264 | 6.716 | 1591.5 | 0.1
10034.2269 | 6.707 | 1591.5 | 0.1
10034.2267 | 6.717 | 1591.5 | 0.1
10034.2266 | 6.73 | 1591.5 | 0.1
10034.2257 | 6.717 | 1591.5 | 0.1
10034.2269 | 6.721 | 1591.5 | 0.1
10034.2272 | 6.723 | 1591.5 | 0.1
10034.2272 | 6.716 | 1591.5 | 0.1
10034.2279 | 6.719 | 1591.5 | 0.1
10034.2263 | 6.726 | 1591.5 | 0.1
10034.2275 | 6.723 | 1591.5 | 0.1
10034.2272 | 6.72 | 1591.5 | 0.1
10034.2279 | 6.723 | 1591.5 | 0.1
10034.2283 | 6.708 | 1591.5 | 0.1
10034.2273 | 6.717 | 1591.5 | 0.1
10034.2287 | 6.727 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10034.2269 | 0.00082721 | 6.7192 | 0.0057671
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 1
2020-10-12 | 18:17
10060.9467 | 6.713 | 1591.5 | 0.1
10060.9462 | 6.713 | 1591.5 | 0.1
10060.9466 | 6.713 | 1591.5 | 0.1
10060.9469 | 6.715 | 1591.5 | 0.1
10060.9464 | 6.714 | 1591.5 | 0.1
10060.9477 | 6.718 | 1591.5 | 0.1
10060.9475 | 6.718 | 1591.5 | 0.1
10060.9474 | 6.714 | 1591.5 | 0.1
10060.9479 | 6.72 | 1591.5 | 0.1
10060.9464 | 6.731 | 1591.5 | 0.1
10060.9484 | 6.714 | 1591.5 | 0.1
10060.9474 | 6.714 | 1591.5 | 0.1
10060.9478 | 6.71 | 1591.5 | 0.1
10060.9483 | 6.702 | 1591.5 | 0.1
10060.9473 | 6.719 | 1591.5 | 0.1
10060.9483 | 6.729 | 1591.5 | 0.1
10060.948 | 6.712 | 1591.5 | 0.1
10060.9485 | 6.709 | 1591.5 | 0.1
10060.9485 | 6.72 | 1591.5 | 0.1
10060.9494 | 6.723 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10060.9476 | 0.0008394 | 6.7161 | 0.0064535
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 2
2020-10-12 | 18:19
20052.3176 | 14.119 | 1591.5 | 0.1
20052.3172 | 14.11 | 1591.5 | 0.1
20052.3163 | 14.109 | 1591.5 | 0.1
20052.3184 | 14.111 | 1591.5 | 0.1
20052.3179 | 14.117 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
20052.3178 | 14.113 | 1591.5 | 0.1
20052.3201 | 14.11 | 1591.5 | 0.1
20052.3199 | 14.139 | 1591.5 | 0.1
20052.319 | 14.126 | 1591.5 | 0.1
20052.32 | 14.121 | 1591.5 | 0.1
20052.3183 | 14.108 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
20052.3199 | 14.121 | 1591.5 | 0.1
20052.3209 | 14.114 | 1591.5 | 0.1
20052.3183 | 14.123 | 1591.5 | 0.1
20052.3208 | 14.112 | 1591.5 | 0.1
20052.3215 | 14.137 | 1591.5 | 0.1
20052.3213 | 14.137 | 1591.5 | 0.1
20052.3188 | 14.13 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
20052.3191 | 0.0013909 | 14.1209 | 0.0099412
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 3
2020-10-12 | 18:22
30063.9944 | 21.56 | 1591.5 | 0.03
30063.987 | 21.56 | 1591.5 | 0.03
30063.989 | 21.57 | 1591.5 | 0.03
30063.984 | 21.59 | 1591.5 | 0.03
30063.989 | 21.59 | 1591.5 | 0.03
30063.991 | 21.6 | 1591.5 | 0.03
30063.9861 | 21.6 | 1591.5 | 0.03
30063.999 | 21.63 | 1591.5 | 0.03
30063.991 | 21.51 | 1591.5 | 0.03
30063.993 | 21.57 | 1591.5 | 0.03
30063.9966 | 21.64 | 1591.5 | 0.03
30063.999 | 21.69 | 1591.5 | 0.03
30063.993 | 21.63 | 1591.5 | 0.03
30063.9883 | 21.56 | 1591.5 | 0.03
30063.9997 | 21.56 | 1591.5 | 0.03
30063.991 | 21.61 | 1591.5 | 0.03
30063.9972 | 21.59 | 1591.5 | 0.03
30063.988 | 21.74 | 1591.5 | 0.03
30064.008 | 21.52 | 1591.5 | 0.03
30063.997 | 21.6 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
30063.9931 | 0.0056836 | 21.596 | 0.051904
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 4
2020-10-12 | 18:24
40066.92 | 32.15 | 1591.5 | 0.03
40066.918 | 32.12 | 1591.5 | 0.03
40066.93 | 32.1 | 1591.5 | 0.03
40066.9219 | 32.05 | 1591.5 | 0.03
40066.917 | 32.03 | 1591.5 | 0.03
40066.918 | 32 | 1591.5 | 0.03
40066.926 | 32.04 | 1591.5 | 0.03
40066.9222 | 32.17 | 1591.5 | 0.03
40066.917 | 32.15 | 1591.5 | 0.03
40066.928 | 32.04 | 1591.5 | 0.03
40066.923 | 32.07 | 1591.5 | 0.03
40066.941 | 31.94 | 1591.5 | 0.03
40066.927 | 32.1 | 1591.5 | 0.03
40066.917 | 31.98 | 1591.5 | 0.03
40066.941 | 32.12 | 1591.5 | 0.03
40066.9245 | 32.01 | 1591.5 | 0.03
40066.938 | 31.96 | 1591.5 | 0.03
40066.929 | 32 | 1591.5 | 0.03
40066.929 | 32.07 | 1591.5 | 0.03
40066.926 | 31.96 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
40066.9257 | 0.00735 | 32.053 | 0.067164
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 5
2020-10-12 | 18:27
50072.953 | 44.52 | 1591.5 | 0.03
50072.948 | 44.6 | 1591.5 | 0.03
50072.953 | 44.57 | 1591.5 | 0.03
50072.964 | 44.57 | 1591.5 | 0.03
50072.953 | 44.7 | 1591.5 | 0.03
50072.932 | 44.73 | 1591.5 | 0.03
50072.9412 | 44.8 | 1591.5 | 0.03
50072.95 | 44.57 | 1591.5 | 0.03
50072.952 | 44.67 | 1591.5 | 0.03
50072.942 | 44.72 | 1591.5 | 0.03
50072.955 | 44.64 | 1591.5 | 0.03
50072.943 | 44.69 | 1591.5 | 0.03
50072.943 | 44.63 | 1591.5 | 0.03
50072.955 | 44.66 | 1591.5 | 0.03
50072.959 | 44.69 | 1591.5 | 0.03
50072.949 | 44.6 | 1591.5 | 0.03
50072.955 | 44.6 | 1591.5 | 0.03
50072.95 | 44.6 | 1591.5 | 0.03
50072.952 | 44.67 | 1591.5 | 0.03
50072.9596 | 44.72 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
50072.9504 | 0.007224 | 44.6475 | 0.06774
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 6
2020-10-12 | 18:29
60084.285 | 58.6 | 1591.5 | 0.03
60084.301 | 58.7 | 1591.5 | 0.03
60084.301 | 58.4 | 1591.5 | 0.03
60084.313 | 58.7 | 1591.5 | 0.03
60084.294 | 58.5 | 1591.5 | 0.03
60084.3 | 58.5 | 1591.5 | 0.03
60084.287 | 58.6 | 1591.5 | 0.03
60084.291 | 58.5 | 1591.5 | 0.03
60084.294 | 58.5 | 1591.5 | 0.03
60084.29 | 58.4 | 1591.5 | 0.03
60084.297 | 58.6 | 1591.5 | 0.03
60084.3 | 58.5 | 1591.5 | 0.03
60084.303 | 58.4 | 1591.5 | 0.03
60084.306 | 58.5 | 1591.5 | 0.03
60084.305 | 58.5 | 1591.5 | 0.03
60084.278 | 58.5 | 1591.5 | 0.03
60084.318 | 58.6 | 1591.5 | 0.03
60084.291 | 58.5 | 1591.5 | 0.03
60084.289 | 58.3 | 1591.5 | 0.03
60084.303 | 58.4 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
60084.2973 | 0.009408 | 58.51 | 0.099499
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 7
2020-10-12 | 18:35
70075.952 | 73.1 | 1591.5 | 0.01
70075.902 | 73 | 1591.5 | 0.01
70075.916 | 72.9 | 1591.5 | 0.01
70075.941 | 73 | 1591.5 | 0.01
70075.959 | 73.1 | 1591.5 | 0.01
70075.993 | 72.8 | 1591.5 | 0.01
70075.878 | 73.3 | 1591.5 | 0.01
70075.95 | 72.9 | 1591.5 | 0.01
70075.889 | 72.9 | 1591.5 | 0.01
70075.88 | 72.8 | 1591.5 | 0.01
70075.94 | 73.3 | 1591.5 | 0.01
70075.905 | 73.7 | 1591.5 | 0.01
70075.995 | 73.7 | 1591.5 | 0.01
70075.952 | 73 | 1591.5 | 0.01
70075.96 | 72.8 | 1591.5 | 0.01
70075.95 | 73.2 | 1591.5 | 0.01
70075.952 | 73.2 | 1591.5 | 0.01
70075.947 | 73.2 | 1591.5 | 0.01
70075.889 | 72.9 | 1591.5 | 0.01
70075.959 | 72.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
70075.9355 | 0.033996 | 73.07 | 0.27586
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 8
2020-10-12 | 18:37
80087.32 | 91.8 | 1591.5 | 0.01
80087.275 | 91.6 | 1591.5 | 0.01
80087.293 | 91.3 | 1591.5 | 0.01
80087.279 | 91.4 | 1591.5 | 0.01
80087.304 | 91.7 | 1591.5 | 0.01
80087.3 | 91.7 | 1591.5 | 0.01
80087.25 | 91.5 | 1591.5 | 0.01
80087.349 | 91.6 | 1591.5 | 0.01
80087.34 | 91.8 | 1591.5 | 0.01
80087.327 | 90.6 | 1591.5 | 0.01
80087.349 | 92.1 | 1591.5 | 0.01
80087.32 | 91.5 | 1591.5 | 0.01
80087.24 | 92.2 | 1591.5 | 0.01
80087.361 | 91 | 1591.5 | 0.01
80087.261 | 91.4 | 1591.5 | 0.01
80087.36 | 91.6 | 1591.5 | 0.01
80087.354 | 92 | 1591.5 | 0.01
80087.421 | 91.1 | 1591.5 | 0.01
80087.331 | 91.7 | 1591.5 | 0.01
80087.329 | 91.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
80087.3182 | 0.04288 | 91.56 | 0.36387
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 9
2020-10-12 | 18:39
90089.851 | 111.4 | 1591.5 | 0.01
90089.93 | 111.4 | 1591.5 | 0.01
90089.85 | 110.7 | 1591.5 | 0.01
90089.93 | 111.6 | 1591.5 | 0.01
90089.98 | 111.5 | 1591.5 | 0.01
90089.97 | 111 | 1591.5 | 0.01
90089.95 | 110.5 | 1591.5 | 0.01
90089.982 | 110.8 | 1591.5 | 0.01
90089.915 | 111.4 | 1591.5 | 0.01
90089.95 | 110.8 | 1591.5 | 0.01
90089.964 | 110.6 | 1591.5 | 0.01
90089.885 | 110.6 | 1591.5 | 0.01
90089.88 | 111 | 1591.5 | 0.01
90089.95 | 110.9 | 1591.5 | 0.01
90089.87 | 110.8 | 1591.5 | 0.01
90089.885 | 110.6 | 1591.5 | 0.01
90089.99 | 110.8 | 1591.5 | 0.01
90089.984 | 110.9 | 1591.5 | 0.01
90089.88 | 110.3 | 1591.5 | 0.01
90089.93 | 110.1 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
90089.9263 | 0.045232 | 110.885 | 0.39532
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail X
2020-10-12 | 18:41
100101.41 | 133.6 | 1591.5 | 0.01
100101.376 | 132.9 | 1591.5 | 0.01
100101.365 | 133.3 | 1591.5 | 0.01
100101.378 | 132.9 | 1591.5 | 0.01
100101.426 | 133 | 1591.5 | 0.01
100101.41 | 134 | 1591.5 | 0.01
100101.376 | 132.9 | 1591.5 | 0.01
100101.42 | 133 | 1591.5 | 0.01
100101.426 | 133.8 | 1591.5 | 0.01
100101.398 | 133.3 | 1591.5 | 0.01
100101.43 | 132.3 | 1591.5 | 0.01
100101.33 | 132.2 | 1591.5 | 0.01
100101.34 | 133.4 | 1591.5 | 0.01
100101.412 | 132.1 | 1591.5 | 0.01
100101.419 | 134.1 | 1591.5 | 0.01
100101.322 | 132.9 | 1591.5 | 0.01
100101.354 | 132.9 | 1591.5 | 0.01
100101.36 | 132.9 | 1591.5 | 0.01
100101.356 | 132.4 | 1591.5 | 0.01
100101.37 | 133.6 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100101.3839 | 0.033297 | 133.075 | 0.55486
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 1
2020-10-12 | 18:44
Capacitor description: | GR sn:4556 0.1uF Dail 1
2020-10-12 | 18:45
100016.06 | 136.1 | 1591.5 | 0.01
100015.938 | 136.2 | 1591.5 | 0.01
100015.84 | 135.1 | 1591.5 | 0.01
100016.02 | 134.9 | 1591.5 | 0.01
100015.907 | 135.2 | 1591.5 | 0.01
100015.977 | 135.1 | 1591.5 | 0.01
100016 | 135.3 | 1591.5 | 0.01
100015.91 | 134.8 | 1591.5 | 0.01
100015.8 | 136 | 1591.5 | 0.01
100016.01 | 136.2 | 1591.5 | 0.01
100015.99 | 135.7 | 1591.5 | 0.01
100015.94 | 134.5 | 1591.5 | 0.01
100015.972 | 136 | 1591.5 | 0.01
100015.98 | 136.1 | 1591.5 | 0.01
100015.94 | 136 | 1591.5 | 0.01
100015.93 | 135.3 | 1591.5 | 0.01
100015.968 | 136 | 1591.5 | 0.01
100015.82 | 136.2 | 1591.5 | 0.01
100015.938 | 135 | 1591.5 | 0.01
100016 | 135.2 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100015.947 | 0.065305 | 135.545 | 0.54173
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 2
2020-10-12 | 18:47
200421.6 | 483.7 | 1591.5 | 0.01
200421.49 | 483.7 | 1591.5 | 0.01
200421.76 | 482.8 | 1591.5 | 0.01
200421.8 | 483.7 | 1591.5 | 0.01
200421.71 | 483.5 | 1591.5 | 0.01
200421.86 | 483 | 1591.5 | 0.01
200421.62 | 481.2 | 1591.5 | 0.01
200421.78 | 482.9 | 1591.5 | 0.01
200421.81 | 482.1 | 1591.5 | 0.01
200421.61 | 483.3 | 1591.5 | 0.01
200421.7 | 483.8 | 1591.5 | 0.01
200421.51 | 481.2 | 1591.5 | 0.01
200421.82 | 482.8 | 1591.5 | 0.01
200421.71 | 481.5 | 1591.5 | 0.01
200421.74 | 482.8 | 1591.5 | 0.01
200421.82 | 481.5 | 1591.5 | 0.01
200421.54 | 484.1 | 1591.5 | 0.01
200421.92 | 482.6 | 1591.5 | 0.01
200421.69 | 483.5 | 1591.5 | 0.01
200421.75 | 481.9 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
200421.712 | 0.11548 | 482.78 | 0.90089
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 3
2020-10-12 | 18:49
300392.5 | 941 | 1591.5 | 0.003
300392.1 | 931 | 1591.5 | 0.003
300391.8 | 938 | 1591.5 | 0.003
300392.25 | 932 | 1591.5 | 0.003
300391.8 | 934 | 1591.5 | 0.003
300392.07 | 933 | 1591.5 | 0.003
300392 | 938 | 1591.5 | 0.003
300392.3 | 938 | 1591.5 | 0.003
300392.7 | 939 | 1591.5 | 0.003
300391.5 | 930 | 1591.5 | 0.003
300391.3 | 939 | 1591.5 | 0.003
300392.5 | 937 | 1591.5 | 0.003
300392.6 | 933 | 1591.5 | 0.003
300392.5 | 935 | 1591.5 | 0.003
300392.1 | 944 | 1591.5 | 0.003
300391.6 | 942 | 1591.5 | 0.003
300392.4 | 939 | 1591.5 | 0.003
300391.9 | 930 | 1591.5 | 0.003
300392.9 | 938 | 1591.5 | 0.003
300391.9 | 940 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
300392.136 | 0.41094 | 936.55 | 3.9683
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 4
2020-10-12 | 18:51
400792 | 1669 | 1591.5 | 0.003
400792.28 | 1673 | 1591.5 | 0.003
400792.2 | 1677 | 1591.5 | 0.003
400792.1 | 1668 | 1591.5 | 0.003
400792.42 | 1667 | 1591.5 | 0.003
400790.9 | 1664 | 1591.5 | 0.003
400791.85 | 1668 | 1591.5 | 0.003
400791.1 | 1666 | 1591.5 | 0.003
400791.48 | 1671 | 1591.5 | 0.003
400792.4 | 1671 | 1591.5 | 0.003
400790.8 | 1675 | 1591.5 | 0.003
400790.4 | 1668 | 1591.5 | 0.003
400792.18 | 1665 | 1591.5 | 0.003
400791.2 | 1665 | 1591.5 | 0.003
400792.3 | 1673 | 1591.5 | 0.003
400791.5 | 1666 | 1591.5 | 0.003
400791.3 | 1680 | 1591.5 | 0.003
400791.7 | 1679 | 1591.5 | 0.003
400792.2 | 1668 | 1591.5 | 0.003
400793.05 | 1674 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
400791.768 | 0.64947 | 1670.35 | 4.6505
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 5
2020-10-12 | 18:54
500613.2 | 2636 | 1591.5 | 0.003
500612.7 | 2621 | 1591.5 | 0.003
500611.5 | 2629 | 1591.5 | 0.003
500612.7 | 2630 | 1591.5 | 0.003
500612.2 | 2636 | 1591.5 | 0.003
500612 | 2632 | 1591.5 | 0.003
500612 | 2632 | 1591.5 | 0.003
500613.6 | 2635 | 1591.5 | 0.003
500612.9 | 2635 | 1591.5 | 0.003
500612.6 | 2633 | 1591.5 | 0.003
500612.1 | 2626 | 1591.5 | 0.003
500611.4 | 2621 | 1591.5 | 0.003
500613.1 | 2623 | 1591.5 | 0.003
500611.8 | 2639 | 1591.5 | 0.003
500612.1 | 2635 | 1591.5 | 0.003
500611.4 | 2632 | 1591.5 | 0.003
500613.4 | 2638 | 1591.5 | 0.003
500612.9 | 2625 | 1591.5 | 0.003
500612.5 | 2636 | 1591.5 | 0.003
500611.5 | 2637 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
500612.38 | 0.66378 | 2631.55 | 5.4907
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 6
2020-10-12 | 18:56
600591.9 | 3681 | 1591.5 | 0.003
600591.6 | 3673 | 1591.5 | 0.003
600590.8 | 3670 | 1591.5 | 0.003
600592.8 | 3681 | 1591.5 | 0.003
600592.3 | 3670 | 1591.5 | 0.003
600593.2 | 3673 | 1591.5 | 0.003
600591.8 | 3679 | 1591.5 | 0.003
600591.7 | 3680 | 1591.5 | 0.003
600591.1 | 3691 | 1591.5 | 0.003
600591.9 | 3691 | 1591.5 | 0.003
600592 | 3674 | 1591.5 | 0.003
600592.9 | 3671 | 1591.5 | 0.003
600592 | 3670 | 1591.5 | 0.003
600591.1 | 3656 | 1591.5 | 0.003
600592.6 | 3679 | 1591.5 | 0.003
600592.4 | 3673 | 1591.5 | 0.003
600590.24 | 3666 | 1591.5 | 0.003
600590.8 | 3662 | 1591.5 | 0.003
600591.8 | 3675 | 1591.5 | 0.003
600591.3 | 3670 | 1591.5 | 0.003
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
600591.812 | 0.74936 | 3674.25 | 8.2757
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 7
2020-10-12 | 18:58
701008.3 | 4833 | 1591.5 | 0.001
701004.8 | 4869 | 1591.5 | 0.001
701004 | 4900 | 1591.5 | 0.001
701007.3 | 4870 | 1591.5 | 0.001
701005 | 4880 | 1591.5 | 0.001
701005 | 4890 | 1591.5 | 0.001
701006.8 | 4938 | 1591.5 | 0.001
701006 | 4810 | 1591.5 | 0.001
701007.7 | 4870 | 1591.5 | 0.001
701006.4 | 4846 | 1591.5 | 0.001
701005.3 | 4870 | 1591.5 | 0.001
701010 | 4880 | 1591.5 | 0.001
701004.1 | 4950 | 1591.5 | 0.001
701008.9 | 4855 | 1591.5 | 0.001
701010.9 | 4864 | 1591.5 | 0.001
701003.4 | 4920 | 1591.5 | 0.001
701001.4 | 4922 | 1591.5 | 0.001
701003 | 4880 | 1591.5 | 0.001
701007.7 | 4950 | 1591.5 | 0.001
701009 | 4870 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
701006.25 | 2.4309 | 4883.35 | 36.3803
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 8
2020-10-12 | 19:00
800991.9 | 6460 | 1591.5 | 0.001
800983.6 | 6360 | 1591.5 | 0.001
800991.9 | 6360 | 1591.5 | 0.001
800990 | 6400 | 1591.5 | 0.001
800990.9 | 6370 | 1591.5 | 0.001
800992.4 | 6400 | 1591.5 | 0.001
800988.3 | 6320 | 1591.5 | 0.001
800994.4 | 6340 | 1591.5 | 0.001
801000 | 6410 | 1591.5 | 0.001
800991.9 | 6330 | 1591.5 | 0.001
800989.5 | 6390 | 1591.5 | 0.001
800995.3 | 6420 | 1591.5 | 0.001
800985.8 | 6410 | 1591.5 | 0.001
800982.7 | 6390 | 1591.5 | 0.001
800990 | 6350 | 1591.5 | 0.001
800987.6 | 6350 | 1591.5 | 0.001
800997.6 | 6320 | 1591.5 | 0.001
800993.7 | 6370 | 1591.5 | 0.001
800994.2 | 6390 | 1591.5 | 0.001
800992.8 | 6440 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
800991.225 | 4.1986 | 6379 | 37.8021
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 9
2020-10-12 | 19:02
901406 | 7970 | 1591.5 | 0.001
901403.7 | 8020 | 1591.5 | 0.001
901397 | 8060 | 1591.5 | 0.001
901404 | 7950 | 1591.5 | 0.001
901402 | 8020 | 1591.5 | 0.001
901396.7 | 7990 | 1591.5 | 0.001
901405.7 | 7970 | 1591.5 | 0.001
901405 | 7980 | 1591.5 | 0.001
901406 | 8090 | 1591.5 | 0.001
901399 | 8040 | 1591.5 | 0.001
901405.9 | 7990 | 1591.5 | 0.001
901394 | 7970 | 1591.5 | 0.001
901401 | 8080 | 1591.5 | 0.001
901409.5 | 8040 | 1591.5 | 0.001
901403.2 | 8080 | 1591.5 | 0.001
901399.7 | 8040 | 1591.5 | 0.001
901412 | 8000 | 1591.5 | 0.001
901404.1 | 7980 | 1591.5 | 0.001
901404 | 8090 | 1591.5 | 0.001
901402.9 | 8050 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
901403.07 | 4.2023 | 8020.5 | 43.872
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail X
2020-10-12 | 19:05
1001394 | 9950 | 1591.5 | 0.001
1001381.8 | 9870 | 1591.5 | 0.001
1001392 | 10040 | 1591.5 | 0.001
1001397.2 | 9970 | 1591.5 | 0.001
1001389 | 10010 | 1591.5 | 0.001
1001394.9 | 9990 | 1591.5 | 0.001
1001383 | 9990 | 1591.5 | 0.001
1001388.1 | 10020 | 1591.5 | 0.001
1001394 | 9940 | 1591.5 | 0.001
1001392 | 9930 | 1591.5 | 0.001
1001383.4 | 10000 | 1591.5 | 0.001
1001397 | 9940 | 1591.5 | 0.001
1001389 | 10010 | 1591.5 | 0.001
1001395.8 | 9960 | 1591.5 | 0.001
1001396 | 9940 | 1591.5 | 0.001
1001389.3 | 9970 | 1591.5 | 0.001
1001394 | 9970 | 1591.5 | 0.001
1001397.8 | 10020 | 1591.5 | 0.001
1001391 | 9920 | 1591.5 | 0.001
1001397 | 10000 | 1591.5 | 0.001
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1001391.815 | 4.8028 | 9972 | 40.5709
end | of | the | test
Capacitor description: | UB1000 measure After Tpye 1412BC sn: 4556 Calibration 20.10deg, 1017.52 mbar, 39.6%RH
2020-10-16 | 14:55
1000.1471 | 0.10849 | 1591.5 | 1.5
1000.1471 | 0.10837 | 1591.5 | 1.5
1000.1471 | 0.10829 | 1591.5 | 1.5
1000.1471 | 0.10831 | 1591.5 | 1.5
1000.1471 | 0.10826 | 1591.5 | 1.5
1000.1471 | 0.10855 | 1591.5 | 1.5
1000.1471 | 0.10841 | 1591.5 | 1.5
1000.1471 | 0.10838 | 1591.5 | 1.5
1000.1471 | 0.10837 | 1591.5 | 1.5
1000.1471 | 0.10841 | 1591.5 | 1.5
1000.1471 | 0.10826 | 1591.5 | 1.5
1000.1471 | 0.10814 | 1591.5 | 1.5
1000.1471 | 0.10806 | 1591.5 | 1.5
1000.1471 | 0.10803 | 1591.5 | 1.5
1000.1471 | 0.10809 | 1591.5 | 1.5
1000.1471 | 0.10811 | 1591.5 | 1.5
1000.1471 | 0.10808 | 1591.5 | 1.5
1000.1471 | 0.10816 | 1591.5 | 1.5
1000.1471 | 0.10806 | 1591.5 | 1.5
1000.1471 | 0.10811 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1471 | 1.4353e-05 | 0.10825 | 0.0001556
end | of | the | test
Capacitor description: | U_connector Zero Point Check, 19.87 degree, 39.7% RH, 1017.19mbar
2020-10-16 | 15:08
0.0051953 | 0.0001177 | 1591.5 | 15
0.005195 | 0.0001192 | 1591.5 | 15
0.0051947 | 0.0001162 | 1591.5 | 15
0.0051943 | 0.0001207 | 1591.5 | 15
0.005194 | 0.0001192 | 1591.5 | 15
0.0051934 | 0.0001177 | 1591.5 | 15
0.0051932 | 0.0001192 | 1591.5 | 15
0.0051925 | 0.0001192 | 1591.5 | 15
0.005192 | 0.0001177 | 1591.5 | 15
0.005192 | 0.000119 | 1591.5 | 15
0.0051916 | 0.0001162 | 1591.5 | 15
0.0051911 | 0.0001192 | 1591.5 | 15
0.0051908 | 0.0001192 | 1591.5 | 15
0.0051908 | 0.0001192 | 1591.5 | 15
0.0051904 | 0.0001192 | 1591.5 | 15
0.0051902 | 0.0001207 | 1591.5 | 15
0.0051894 | 0.0001192 | 1591.5 | 15
0.0051894 | 0.0001132 | 1591.5 | 15
0.0051892 | 0.000119 | 1591.5 | 15
0.0051892 | 0.0001162 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
0.0051919 | 1.985e-06 | 0.00011836 | 1.7328e-06
end | of | the | test
Capacitor description: | UB1000 measurement @19/10/2020 20.27deg, 1006.12 mbar, 58.2%RH
2020-10-19 | 17:32
1000.1393 | 0.10428 | 1591.5 | 1.5
1000.1393 | 0.10425 | 1591.5 | 1.5
1000.1393 | 0.10423 | 1591.5 | 1.5
1000.1393 | 0.10419 | 1591.5 | 1.5
1000.1393 | 0.10426 | 1591.5 | 1.5
1000.1393 | 0.10429 | 1591.5 | 1.5
1000.1393 | 0.10432 | 1591.5 | 1.5
1000.1393 | 0.10429 | 1591.5 | 1.5
1000.1393 | 0.10426 | 1591.5 | 1.5
1000.1393 | 0.10422 | 1591.5 | 1.5
1000.1393 | 0.10422 | 1591.5 | 1.5
1000.1393 | 0.10423 | 1591.5 | 1.5
1000.1393 | 0.10416 | 1591.5 | 1.5
1000.1393 | 0.10425 | 1591.5 | 1.5
1000.1393 | 0.10433 | 1591.5 | 1.5
1000.1393 | 0.10417 | 1591.5 | 1.5
1000.1393 | 0.10419 | 1591.5 | 1.5
1000.1393 | 0.10432 | 1591.5 | 1.5
1000.1393 | 0.10425 | 1591.5 | 1.5
1000.1393 | 0.1043 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1393 | 7.4833e-06 | 0.10425 | 4.8629e-05
end | of | the | test
Capacitor description: | UB1000 measurement @20/10/2020 20.44deg, 1012.36 mbar, 53.2%RH
2020-10-20 | 12:16
1000.1439 | 0.13889 | 1591.5 | 1.5
1000.1439 | 0.13882 | 1591.5 | 1.5
1000.1439 | 0.13891 | 1591.5 | 1.5
1000.1439 | 0.13891 | 1591.5 | 1.5
1000.1439 | 0.13889 | 1591.5 | 1.5
1000.1439 | 0.1388 | 1591.5 | 1.5
1000.1439 | 0.13879 | 1591.5 | 1.5
1000.1439 | 0.13882 | 1591.5 | 1.5
1000.1439 | 0.13882 | 1591.5 | 1.5
1000.1439 | 0.13883 | 1591.5 | 1.5
1000.1439 | 0.13879 | 1591.5 | 1.5
1000.1439 | 0.13883 | 1591.5 | 1.5
1000.1439 | 0.13883 | 1591.5 | 1.5
1000.1439 | 0.13879 | 1591.5 | 1.5
1000.1439 | 0.13876 | 1591.5 | 1.5
1000.1439 | 0.13871 | 1591.5 | 1.5
1000.1439 | 0.13882 | 1591.5 | 1.5
1000.1439 | 0.13883 | 1591.5 | 1.5
1000.1439 | 0.13874 | 1591.5 | 1.5
1000.1439 | 0.13874 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1439 | 7.4162e-06 | 0.13882 | 5.3703e-05
end | of | the | test
Capacitor description: | UB1000 measurement @23/10/2020 20.12deg, 1012.18 mbar, 50.3%RH
2020-10-23 | 11:48
1000.1511 | 0.10824 | 1591.5 | 1.5
1000.1511 | 0.10818 | 1591.5 | 1.5
1000.1511 | 0.10822 | 1591.5 | 1.5
1000.1511 | 0.10819 | 1591.5 | 1.5
1000.1511 | 0.10828 | 1591.5 | 1.5
1000.1511 | 0.10822 | 1591.5 | 1.5
1000.1511 | 0.10833 | 1591.5 | 1.5
1000.1511 | 0.10821 | 1591.5 | 1.5
1000.1511 | 0.10824 | 1591.5 | 1.5
1000.1511 | 0.1083 | 1591.5 | 1.5
1000.1511 | 0.10819 | 1591.5 | 1.5
1000.1511 | 0.10818 | 1591.5 | 1.5
1000.1511 | 0.10816 | 1591.5 | 1.5
1000.1511 | 0.10828 | 1591.5 | 1.5
1000.1511 | 0.10818 | 1591.5 | 1.5
1000.1511 | 0.10827 | 1591.5 | 1.5
1000.1511 | 0.10831 | 1591.5 | 1.5
1000.1511 | 0.1083 | 1591.5 | 1.5
1000.1511 | 0.10821 | 1591.5 | 1.5
1000.1511 | 0.10828 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1511 | 3e-06 | 0.10824 | 5.0326e-05
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 0 20.05°C
2020-10-23 | 11:56
49.275 | 1.1969 | 1591.5 | 15
49.275 | 1.1967 | 1591.5 | 15
49.275 | 1.1968 | 1591.5 | 15
49.275 | 1.1972 | 1591.5 | 15
49.275 | 1.1965 | 1591.5 | 15
49.275 | 1.1965 | 1591.5 | 15
49.275 | 1.1963 | 1591.5 | 15
49.275 | 1.1965 | 1591.5 | 15
49.275 | 1.1966 | 1591.5 | 15
49.275 | 1.1964 | 1591.5 | 15
49.275 | 1.1966 | 1591.5 | 15
49.275 | 1.1966 | 1591.5 | 15
49.275 | 1.1963 | 1591.5 | 15
49.275 | 1.1967 | 1591.5 | 15
49.275 | 1.1968 | 1591.5 | 15
49.275 | 1.1965 | 1591.5 | 15
49.275 | 1.1965 | 1591.5 | 15
49.275 | 1.1971 | 1591.5 | 15
49.275 | 1.1977 | 1591.5 | 15
49.275 | 1.1972 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
49.275 | 9.6652e-06 | 1.1967 | 0.00034907
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 1 20.02°C
2020-10-23 | 11:59
145.9576 | 1.2741 | 1591.5 | 7.5
145.9576 | 1.2739 | 1591.5 | 7.5
145.9575 | 1.2739 | 1591.5 | 7.5
145.9575 | 1.2737 | 1591.5 | 7.5
145.9575 | 1.274 | 1591.5 | 7.5
145.9575 | 1.274 | 1591.5 | 7.5
145.9575 | 1.2743 | 1591.5 | 7.5
145.9575 | 1.2746 | 1591.5 | 7.5
145.9575 | 1.2744 | 1591.5 | 7.5
145.9575 | 1.2744 | 1591.5 | 7.5
145.9575 | 1.2744 | 1591.5 | 7.5
145.9575 | 1.2749 | 1591.5 | 7.5
145.9575 | 1.2751 | 1591.5 | 7.5
145.9575 | 1.2745 | 1591.5 | 7.5
145.9575 | 1.2751 | 1591.5 | 7.5
145.9575 | 1.2748 | 1591.5 | 7.5
145.9575 | 1.2747 | 1591.5 | 7.5
145.9575 | 1.2745 | 1591.5 | 7.5
145.9575 | 1.2745 | 1591.5 | 7.5
145.9575 | 1.2742 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
145.9575 | 2.8676e-05 | 1.2744 | 0.00038384
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 2 20.01°C
2020-10-23 | 12:01
244.9965 | 1.2989 | 1591.5 | 7.5
244.9965 | 1.2992 | 1591.5 | 7.5
244.9965 | 1.2993 | 1591.5 | 7.5
244.9965 | 1.299 | 1591.5 | 7.5
244.9965 | 1.2988 | 1591.5 | 7.5
244.9965 | 1.2986 | 1591.5 | 7.5
244.9965 | 1.2986 | 1591.5 | 7.5
244.9965 | 1.2988 | 1591.5 | 7.5
244.9965 | 1.2989 | 1591.5 | 7.5
244.9965 | 1.2984 | 1591.5 | 7.5
244.9965 | 1.2984 | 1591.5 | 7.5
244.9965 | 1.2988 | 1591.5 | 7.5
244.9965 | 1.2988 | 1591.5 | 7.5
244.9965 | 1.2987 | 1591.5 | 7.5
244.9965 | 1.2994 | 1591.5 | 7.5
244.9821 | 1.4448 | 1591.5 | 7.5
244.982 | 1.4448 | 1591.5 | 7.5
244.9821 | 1.445 | 1591.5 | 7.5
244.9826 | 1.4441 | 1591.5 | 7.5
244.9826 | 1.4444 | 1591.5 | 7.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
244.993 | 0.0061638 | 1.3353 | 0.063117
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 3 19.99°C
2020-10-23 | 12:04
341.514 | 1.5207 | 1591.5 | 3
341.5141 | 1.5208 | 1591.5 | 3
341.5138 | 1.5208 | 1591.5 | 3
341.5138 | 1.5208 | 1591.5 | 3
341.5137 | 1.5209 | 1591.5 | 3
341.5135 | 1.5208 | 1591.5 | 3
341.513 | 1.5212 | 1591.5 | 3
341.5132 | 1.5212 | 1591.5 | 3
341.5136 | 1.5208 | 1591.5 | 3
341.5137 | 1.5212 | 1591.5 | 3
341.5132 | 1.5209 | 1591.5 | 3
341.513 | 1.5205 | 1591.5 | 3
341.5123 | 1.5202 | 1591.5 | 3
341.5122 | 1.5202 | 1591.5 | 3
341.5126 | 1.5207 | 1591.5 | 3
341.5127 | 1.5214 | 1591.5 | 3
341.5125 | 1.5207 | 1591.5 | 3
341.5127 | 1.5204 | 1591.5 | 3
341.5123 | 1.5207 | 1591.5 | 3
341.5123 | 1.5208 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
341.5131 | 0.0006186 | 1.5208 | 0.00031844
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 4 20.00°C
2020-10-23 | 12:06
441.3128 | 1.5515 | 1591.5 | 3
441.3129 | 1.5518 | 1591.5 | 3
441.3128 | 1.551 | 1591.5 | 3
441.3126 | 1.5515 | 1591.5 | 3
441.3127 | 1.5511 | 1591.5 | 3
441.3125 | 1.5515 | 1591.5 | 3
441.3123 | 1.5514 | 1591.5 | 3
441.3123 | 1.5514 | 1591.5 | 3
441.3122 | 1.5512 | 1591.5 | 3
441.3121 | 1.5514 | 1591.5 | 3
441.3123 | 1.5515 | 1591.5 | 3
441.3123 | 1.5524 | 1591.5 | 3
441.3124 | 1.5523 | 1591.5 | 3
441.3123 | 1.5517 | 1591.5 | 3
441.3123 | 1.5518 | 1591.5 | 3
441.3124 | 1.5517 | 1591.5 | 3
441.3123 | 1.5519 | 1591.5 | 3
441.3123 | 1.5518 | 1591.5 | 3
441.3124 | 1.5517 | 1591.5 | 3
441.312 | 1.5519 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
441.3124 | 0.00022718 | 1.5516 | 0.00032693
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 5 20.00°C
2020-10-23 | 12:08
537.8137 | 1.63 | 1591.5 | 3
537.8137 | 1.631 | 1591.5 | 3
537.8138 | 1.6311 | 1591.5 | 3
537.814 | 1.6306 | 1591.5 | 3
537.814 | 1.6312 | 1591.5 | 3
537.8139 | 1.6308 | 1591.5 | 3
537.8145 | 1.6301 | 1591.5 | 3
537.8138 | 1.6301 | 1591.5 | 3
537.814 | 1.6304 | 1591.5 | 3
537.8137 | 1.6302 | 1591.5 | 3
537.8141 | 1.6308 | 1591.5 | 3
537.814 | 1.6311 | 1591.5 | 3
537.8139 | 1.6312 | 1591.5 | 3
537.8138 | 1.6316 | 1591.5 | 3
537.8137 | 1.6318 | 1591.5 | 3
537.8135 | 1.6312 | 1591.5 | 3
537.8136 | 1.6311 | 1591.5 | 3
537.8136 | 1.6313 | 1591.5 | 3
537.8136 | 1.6311 | 1591.5 | 3
537.8136 | 1.6312 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
537.8138 | 0.00023011 | 1.6309 | 0.00050685
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 6 19.99°C
2020-10-23 | 12:10
640.7429 | 1.6675 | 1591.5 | 3
640.7432 | 1.6677 | 1591.5 | 3
640.7433 | 1.6676 | 1591.5 | 3
640.7435 | 1.6673 | 1591.5 | 3
640.7435 | 1.6677 | 1591.5 | 3
640.7434 | 1.6677 | 1591.5 | 3
640.7435 | 1.6674 | 1591.5 | 3
640.7434 | 1.6676 | 1591.5 | 3
640.7432 | 1.6677 | 1591.5 | 3
640.743 | 1.6676 | 1591.5 | 3
640.7431 | 1.6674 | 1591.5 | 3
640.7431 | 1.6676 | 1591.5 | 3
640.7431 | 1.6677 | 1591.5 | 3
640.7427 | 1.6674 | 1591.5 | 3
640.7429 | 1.6674 | 1591.5 | 3
640.7428 | 1.6672 | 1591.5 | 3
640.7428 | 1.6679 | 1591.5 | 3
640.7428 | 1.6681 | 1591.5 | 3
640.7427 | 1.6682 | 1591.5 | 3
640.7426 | 1.6676 | 1591.5 | 3
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
640.7431 | 0.00028835 | 1.6676 | 0.00024606
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 7 20.00°C
2020-10-23 | 12:12
739.8533 | 1.692 | 1591.5 | 1.5
739.8532 | 1.6922 | 1591.5 | 1.5
739.854 | 1.6926 | 1591.5 | 1.5
739.8534 | 1.6924 | 1591.5 | 1.5
739.853 | 1.6918 | 1591.5 | 1.5
739.853 | 1.6918 | 1591.5 | 1.5
739.8529 | 1.6912 | 1591.5 | 1.5
739.8531 | 1.6916 | 1591.5 | 1.5
739.8528 | 1.691 | 1591.5 | 1.5
739.8531 | 1.6921 | 1591.5 | 1.5
739.8527 | 1.6914 | 1591.5 | 1.5
739.8526 | 1.6919 | 1591.5 | 1.5
739.8526 | 1.6916 | 1591.5 | 1.5
739.8528 | 1.6916 | 1591.5 | 1.5
739.8533 | 1.6925 | 1591.5 | 1.5
739.8534 | 1.6928 | 1591.5 | 1.5
739.8529 | 1.692 | 1591.5 | 1.5
739.853 | 1.6925 | 1591.5 | 1.5
739.8534 | 1.6928 | 1591.5 | 1.5
739.853 | 1.6923 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
739.8531 | 0.00031824 | 1.692 | 0.00049595
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 8 20.01°C
2020-10-23 | 12:15
836.3874 | 1.7737 | 1591.5 | 1.5
836.387 | 1.7738 | 1591.5 | 1.5
836.3865 | 1.7734 | 1591.5 | 1.5
836.3866 | 1.7731 | 1591.5 | 1.5
836.3864 | 1.7728 | 1591.5 | 1.5
836.3863 | 1.7727 | 1591.5 | 1.5
836.386 | 1.7727 | 1591.5 | 1.5
836.3861 | 1.7728 | 1591.5 | 1.5
836.386 | 1.7729 | 1591.5 | 1.5
836.3861 | 1.7726 | 1591.5 | 1.5
836.3864 | 1.7734 | 1591.5 | 1.5
836.3875 | 1.7742 | 1591.5 | 1.5
836.3877 | 1.7742 | 1591.5 | 1.5
836.3877 | 1.775 | 1591.5 | 1.5
836.3878 | 1.7746 | 1591.5 | 1.5
836.3874 | 1.7735 | 1591.5 | 1.5
836.3875 | 1.7735 | 1591.5 | 1.5
836.3871 | 1.7738 | 1591.5 | 1.5
836.3873 | 1.7738 | 1591.5 | 1.5
836.3873 | 1.7739 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
836.3869 | 0.00061833 | 1.7735 | 0.00064987
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail 9 20.01°C
2020-10-23 | 12:18
936.2004 | 1.8014 | 1591.5 | 1.5
936.2005 | 1.8016 | 1591.5 | 1.5
936.2007 | 1.8016 | 1591.5 | 1.5
936.2002 | 1.8011 | 1591.5 | 1.5
936.2001 | 1.8007 | 1591.5 | 1.5
936.2001 | 1.801 | 1591.5 | 1.5
936.2002 | 1.8013 | 1591.5 | 1.5
936.2002 | 1.8016 | 1591.5 | 1.5
936.1998 | 1.8012 | 1591.5 | 1.5
936.2002 | 1.8012 | 1591.5 | 1.5
936.2004 | 1.8013 | 1591.5 | 1.5
936.2005 | 1.8014 | 1591.5 | 1.5
936.2007 | 1.8018 | 1591.5 | 1.5
936.2007 | 1.8021 | 1591.5 | 1.5
936.2007 | 1.8018 | 1591.5 | 1.5
936.2007 | 1.8017 | 1591.5 | 1.5
936.201 | 1.8022 | 1591.5 | 1.5
936.2011 | 1.8023 | 1591.5 | 1.5
936.2012 | 1.8025 | 1591.5 | 1.5
936.2013 | 1.8022 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
936.2005 | 0.00039721 | 1.8016 | 0.00046248
end | of | the | test
Capacitor description: | GR sn:4556 100pF Dail X 20.01°C
2020-10-23 | 12:20
1032.5035 | 1.8766 | 1591.5 | 1.5
1032.5032 | 1.8767 | 1591.5 | 1.5
1032.5032 | 1.877 | 1591.5 | 1.5
1032.5032 | 1.8765 | 1591.5 | 1.5
1032.5035 | 1.8768 | 1591.5 | 1.5
1032.5033 | 1.8774 | 1591.5 | 1.5
1032.5033 | 1.8768 | 1591.5 | 1.5
1032.5033 | 1.8767 | 1591.5 | 1.5
1032.5034 | 1.877 | 1591.5 | 1.5
1032.5033 | 1.8769 | 1591.5 | 1.5
1032.5031 | 1.8772 | 1591.5 | 1.5
1032.503 | 1.877 | 1591.5 | 1.5
1032.5031 | 1.8772 | 1591.5 | 1.5
1032.5029 | 1.8773 | 1591.5 | 1.5
1032.5027 | 1.8769 | 1591.5 | 1.5
1032.5027 | 1.8773 | 1591.5 | 1.5
1032.5025 | 1.8773 | 1591.5 | 1.5
1032.5027 | 1.8773 | 1591.5 | 1.5
1032.5026 | 1.8768 | 1591.5 | 1.5
1032.5025 | 1.8769 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1032.503 | 0.00031839 | 1.877 | 0.00026451
end | of | the | test
Capacitor description: | GR sn:4556 0.1uF Dail 0 20.01°C
2020-10-23 | 12:23
49.2759 | 1.3453 | 1591.5 | 15
49.2764 | 1.3445 | 1591.5 | 15
49.2766 | 1.3447 | 1591.5 | 15
49.2762 | 1.3453 | 1591.5 | 15
49.2765 | 1.3452 | 1591.5 | 15
49.2763 | 1.3446 | 1591.5 | 15
49.2765 | 1.345 | 1591.5 | 15
49.2764 | 1.3448 | 1591.5 | 15
49.276 | 1.3446 | 1591.5 | 15
49.276 | 1.3448 | 1591.5 | 15
49.276 | 1.3446 | 1591.5 | 15
49.276 | 1.3449 | 1591.5 | 15
49.2767 | 1.3443 | 1591.5 | 15
49.2765 | 1.3447 | 1591.5 | 15
49.2765 | 1.3442 | 1591.5 | 15
49.2766 | 1.3448 | 1591.5 | 15
49.2765 | 1.3445 | 1591.5 | 15
49.2764 | 1.3447 | 1591.5 | 15
49.2767 | 1.3447 | 1591.5 | 15
49.2768 | 1.3447 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
49.2764 | 0.00026341 | 1.3448 | 0.00027408
end | of | the | test
Capacitor description: | GR sn:4556 0.01UuF Dail 20.02°C
2020-10-23 | 12:26
49.2766 | 1.344 | 1591.5 | 15
49.2766 | 1.3439 | 1591.5 | 15
49.2766 | 1.3448 | 1591.5 | 15
49.2764 | 1.3448 | 1591.5 | 15
49.2766 | 1.3449 | 1591.5 | 15
49.2766 | 1.3444 | 1591.5 | 15
49.2767 | 1.3448 | 1591.5 | 15
49.2769 | 1.3445 | 1591.5 | 15
49.2769 | 1.3446 | 1591.5 | 15
49.277 | 1.3439 | 1591.5 | 15
49.2768 | 1.3443 | 1591.5 | 15
49.2765 | 1.3444 | 1591.5 | 15
49.2766 | 1.3447 | 1591.5 | 15
49.2764 | 1.3448 | 1591.5 | 15
49.2762 | 1.345 | 1591.5 | 15
49.2763 | 1.3451 | 1591.5 | 15
49.2763 | 1.3449 | 1591.5 | 15
49.2762 | 1.3451 | 1591.5 | 15
49.2762 | 1.3453 | 1591.5 | 15
49.2762 | 1.3449 | 1591.5 | 15
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
49.2765 | 0.00023971 | 1.3447 | 0.00038761
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 1 20.02°C (last reading is the zero dial, and the one before it range written wrong)
2020-10-23 | 12:30
10061.0137 | 6.752 | 1591.5 | 0.1
10061.0126 | 6.738 | 1591.5 | 0.1
10061.0122 | 6.741 | 1591.5 | 0.1
10061.0121 | 6.748 | 1591.5 | 0.1
10061.0096 | 6.736 | 1591.5 | 0.1
10061.0079 | 6.725 | 1591.5 | 0.1
10061.0084 | 6.726 | 1591.5 | 0.1
10061.0082 | 6.729 | 1591.5 | 0.1
10061.0076 | 6.728 | 1591.5 | 0.1
10061.0066 | 6.736 | 1591.5 | 0.1
10061.0083 | 6.746 | 1591.5 | 0.1
10061.0079 | 6.735 | 1591.5 | 0.1
10061.0075 | 6.735 | 1591.5 | 0.1
10061.007 | 6.726 | 1591.5 | 0.1
10061.0064 | 6.735 | 1591.5 | 0.1
10061.0062 | 6.726 | 1591.5 | 0.1
10061.0062 | 6.727 | 1591.5 | 0.1
10061.01 | 6.754 | 1591.5 | 0.1
10061.0085 | 6.748 | 1591.5 | 0.1
10061.0086 | 6.747 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
10061.0088 | 0.0021906 | 6.7369 | 0.0092731
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 2 20.01°C
2020-10-23 | 12:35
20051.526 | 14.132 | 1591.5 | 0.1
20051.5262 | 14.119 | 1591.5 | 0.1
20051.5247 | 14.13 | 1591.5 | 0.1
20051.5238 | 14.115 | 1591.5 | 0.1
20051.5247 | 14.112 | 1591.5 | 0.1
20051.523 | 14.112 | 1591.5 | 0.1
20051.5224 | 14.124 | 1591.5 | 0.1
20051.5204 | 14.11 | 1591.5 | 0.1
20051.5217 | 14.099 | 1591.5 | 0.1
20051.5212 | 14.115 | 1591.5 | 0.1
20051.5188 | 14.13 | 1591.5 | 0.1
20051.5206 | 14.123 | 1591.5 | 0.1
20051.5206 | 14.109 | 1591.5 | 0.1
20051.517 | 14.139 | 1591.5 | 0.1
20051.5189 | 14.119 | 1591.5 | 0.1
20051.5182 | 14.126 | 1591.5 | 0.1
20051.5181 | 14.118 | 1591.5 | 0.1
20051.5202 | 14.127 | 1591.5 | 0.1
20051.5178 | 14.12 | 1591.5 | 0.1
20051.5148 | 14.112 | 1591.5 | 0.1
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
20051.521 | 0.0030431 | 14.1196 | 0.0092762
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 3 20.01°C
2020-10-23 | 12:37
30063.1641 | 21.51 | 1591.5 | 0.03
30063.16 | 21.5 | 1591.5 | 0.03
30063.158 | 21.48 | 1591.5 | 0.03
30063.162 | 21.52 | 1591.5 | 0.03
30063.152 | 21.5 | 1591.5 | 0.03
30063.1531 | 21.55 | 1591.5 | 0.03
30063.1643 | 21.46 | 1591.5 | 0.03
30063.1594 | 21.53 | 1591.5 | 0.03
30063.154 | 21.45 | 1591.5 | 0.03
30063.155 | 21.6 | 1591.5 | 0.03
30063.153 | 21.55 | 1591.5 | 0.03
30063.1551 | 21.52 | 1591.5 | 0.03
30063.1531 | 21.48 | 1591.5 | 0.03
30063.15 | 21.59 | 1591.5 | 0.03
30063.144 | 21.53 | 1591.5 | 0.03
30063.1463 | 21.59 | 1591.5 | 0.03
30063.146 | 21.52 | 1591.5 | 0.03
30063.145 | 21.57 | 1591.5 | 0.03
30063.15 | 21.48 | 1591.5 | 0.03
30063.1426 | 21.52 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
30063.1534 | 0.0064027 | 21.5225 | 0.041698
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 4 20.01°C
2020-10-23 | 12:40
40065.656 | 32.17 | 1591.5 | 0.03
40065.649 | 32.17 | 1591.5 | 0.03
40065.654 | 32.17 | 1591.5 | 0.03
40065.655 | 32.1 | 1591.5 | 0.03
40065.653 | 32.1 | 1591.5 | 0.03
40065.657 | 32.08 | 1591.5 | 0.03
40065.652 | 32.12 | 1591.5 | 0.03
40065.6549 | 32.13 | 1591.5 | 0.03
40065.655 | 32.03 | 1591.5 | 0.03
40065.6511 | 31.99 | 1591.5 | 0.03
40065.6555 | 32.1 | 1591.5 | 0.03
40065.653 | 32.15 | 1591.5 | 0.03
40065.6399 | 32.12 | 1591.5 | 0.03
40065.654 | 32.22 | 1591.5 | 0.03
40065.651 | 32.23 | 1591.5 | 0.03
40065.635 | 32.12 | 1591.5 | 0.03
40065.643 | 31.96 | 1591.5 | 0.03
40065.645 | 32.01 | 1591.5 | 0.03
40065.646 | 32.08 | 1591.5 | 0.03
40065.6519 | 32.11 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
40065.6506 | 0.0057476 | 32.108 | 0.068891
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 5 20.02°C
2020-10-23 | 12:42
50073.854 | 44.85 | 1591.5 | 0.03
50073.863 | 44.8 | 1591.5 | 0.03
50073.8515 | 44.62 | 1591.5 | 0.03
50073.869 | 44.78 | 1591.5 | 0.03
50073.854 | 44.75 | 1591.5 | 0.03
50073.857 | 44.57 | 1591.5 | 0.03
50073.867 | 44.69 | 1591.5 | 0.03
50073.86 | 44.8 | 1591.5 | 0.03
50073.856 | 44.77 | 1591.5 | 0.03
50073.844 | 44.76 | 1591.5 | 0.03
50073.841 | 44.75 | 1591.5 | 0.03
50073.847 | 44.69 | 1591.5 | 0.03
50073.838 | 44.66 | 1591.5 | 0.03
50073.86 | 44.77 | 1591.5 | 0.03
50073.834 | 44.75 | 1591.5 | 0.03
50073.836 | 44.87 | 1591.5 | 0.03
50073.8508 | 44.72 | 1591.5 | 0.03
50073.842 | 44.82 | 1591.5 | 0.03
50073.833 | 44.72 | 1591.5 | 0.03
50073.841 | 44.69 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
50073.8499 | 0.010712 | 44.7415 | 0.072338
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 6 20.02°C
2020-10-23 | 12:44
50073.82 | 44.7 | 1591.5 | 0.03
50073.818 | 44.7 | 1591.5 | 0.03
50073.827 | 44.83 | 1591.5 | 0.03
50073.828 | 44.87 | 1591.5 | 0.03
50073.8166 | 44.7 | 1591.5 | 0.03
50073.812 | 44.62 | 1591.5 | 0.03
50073.822 | 44.68 | 1591.5 | 0.03
50073.817 | 44.63 | 1591.5 | 0.03
50073.821 | 44.85 | 1591.5 | 0.03
50073.804 | 44.78 | 1591.5 | 0.03
50073.81 | 44.78 | 1591.5 | 0.03
50073.808 | 44.84 | 1591.5 | 0.03
50073.803 | 44.72 | 1591.5 | 0.03
50073.825 | 44.79 | 1591.5 | 0.03
50073.806 | 44.78 | 1591.5 | 0.03
50073.818 | 44.79 | 1591.5 | 0.03
50073.806 | 44.8 | 1591.5 | 0.03
50073.803 | 44.78 | 1591.5 | 0.03
50073.809 | 44.81 | 1591.5 | 0.03
50073.814 | 44.83 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
50073.8144 | 0.007845 | 44.764 | 0.070314
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 6 20.02°C
2020-10-23 | 12:47
60085.043 | 58.7 | 1591.5 | 0.03
60085.052 | 58.8 | 1591.5 | 0.03
60085.038 | 58.6 | 1591.5 | 0.03
60085.055 | 58.7 | 1591.5 | 0.03
60085.049 | 58.8 | 1591.5 | 0.03
60085.052 | 58.4 | 1591.5 | 0.03
60085.048 | 58.8 | 1591.5 | 0.03
60085.052 | 58.7 | 1591.5 | 0.03
60085.044 | 58.7 | 1591.5 | 0.03
60085.033 | 58.7 | 1591.5 | 0.03
60085.043 | 58.7 | 1591.5 | 0.03
60085.033 | 58.7 | 1591.5 | 0.03
60085.013 | 58.8 | 1591.5 | 0.03
60085.029 | 58.6 | 1591.5 | 0.03
60085.031 | 58.7 | 1591.5 | 0.03
60085.036 | 58.8 | 1591.5 | 0.03
60085.023 | 58.6 | 1591.5 | 0.03
60085.033 | 58.6 | 1591.5 | 0.03
60085.022 | 58.8 | 1591.5 | 0.03
60085.021 | 58.9 | 1591.5 | 0.03
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
60085.0375 | 0.011796 | 58.705 | 0.10712
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 7 20.02°C
2020-10-23 | 12:49
70075.763 | 73.6 | 1591.5 | 0.01
70075.73 | 73.3 | 1591.5 | 0.01
70075.776 | 74 | 1591.5 | 0.01
70075.736 | 73 | 1591.5 | 0.01
70075.72 | 73.6 | 1591.5 | 0.01
70075.73 | 73.8 | 1591.5 | 0.01
70075.729 | 73.8 | 1591.5 | 0.01
70075.747 | 73.6 | 1591.5 | 0.01
70075.785 | 73.6 | 1591.5 | 0.01
70075.74 | 73.5 | 1591.5 | 0.01
70075.705 | 73.4 | 1591.5 | 0.01
70075.73 | 73.4 | 1591.5 | 0.01
70075.747 | 73.7 | 1591.5 | 0.01
70075.75 | 73.7 | 1591.5 | 0.01
70075.765 | 73.5 | 1591.5 | 0.01
70075.727 | 73.6 | 1591.5 | 0.01
70075.742 | 73.4 | 1591.5 | 0.01
70075.727 | 73.6 | 1591.5 | 0.01
70075.724 | 73.7 | 1591.5 | 0.01
70075.711 | 73.7 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
70075.7392 | 0.0202 | 73.575 | 0.20706
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 8 20.03°C
2020-10-23 | 12:51
80086.984 | 92.2 | 1591.5 | 0.01
80087.047 | 91.6 | 1591.5 | 0.01
80087.027 | 92.1 | 1591.5 | 0.01
80086.957 | 91.9 | 1591.5 | 0.01
80087.007 | 92.2 | 1591.5 | 0.01
80086.996 | 91.8 | 1591.5 | 0.01
80086.928 | 93.2 | 1591.5 | 0.01
80086.935 | 92 | 1591.5 | 0.01
80086.955 | 92.1 | 1591.5 | 0.01
80086.948 | 92.1 | 1591.5 | 0.01
80087.038 | 92.3 | 1591.5 | 0.01
80087.02 | 92.2 | 1591.5 | 0.01
80086.94 | 92.1 | 1591.5 | 0.01
80086.903 | 92.5 | 1591.5 | 0.01
80086.996 | 92.1 | 1591.5 | 0.01
80087.01 | 92.3 | 1591.5 | 0.01
80086.969 | 92.1 | 1591.5 | 0.01
80086.917 | 92 | 1591.5 | 0.01
80086.98 | 92.4 | 1591.5 | 0.01
80086.829 | 93.1 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
80086.9693 | 0.051408 | 92.215 | 0.3678
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail 9 20.04°C
2020-10-23 | 12:54
90089.07 | 111.9 | 1591.5 | 0.01
90089.07 | 111.6 | 1591.5 | 0.01
90089.126 | 111.8 | 1591.5 | 0.01
90089.12 | 111.9 | 1591.5 | 0.01
90089.059 | 111.3 | 1591.5 | 0.01
90089.045 | 111.5 | 1591.5 | 0.01
90089.11 | 111.1 | 1591.5 | 0.01
90089.08 | 111.2 | 1591.5 | 0.01
90089.06 | 111.8 | 1591.5 | 0.01
90089.04 | 112.1 | 1591.5 | 0.01
90089.04 | 111.7 | 1591.5 | 0.01
90089.06 | 111.4 | 1591.5 | 0.01
90089.09 | 111.4 | 1591.5 | 0.01
90089.07 | 112 | 1591.5 | 0.01
90088.99 | 112 | 1591.5 | 0.01
90089.09 | 110.8 | 1591.5 | 0.01
90089.03 | 112.4 | 1591.5 | 0.01
90089.049 | 111.5 | 1591.5 | 0.01
90089.112 | 111.8 | 1591.5 | 0.01
90089.02 | 111.2 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
90089.0666 | 0.034168 | 111.62 | 0.38158
end | of | the | test
Capacitor description: | GR sn:4556 0.01uF Dail X 20.04°C
2020-10-23 | 12:56
100100.3 | 134 | 1591.5 | 0.01
100100.36 | 134.1 | 1591.5 | 0.01
100100.41 | 134.3 | 1591.5 | 0.01
100100.37 | 133.4 | 1591.5 | 0.01
100100.265 | 134.6 | 1591.5 | 0.01
100100.37 | 133.6 | 1591.5 | 0.01
100100.37 | 134 | 1591.5 | 0.01
100100.3 | 134.2 | 1591.5 | 0.01
100100.3 | 134.4 | 1591.5 | 0.01
100100.328 | 134.2 | 1591.5 | 0.01
100100.35 | 135.1 | 1591.5 | 0.01
100100.27 | 133.5 | 1591.5 | 0.01
100100.299 | 133.9 | 1591.5 | 0.01
100100.35 | 134.8 | 1591.5 | 0.01
100100.4 | 134 | 1591.5 | 0.01
100100.272 | 134.8 | 1591.5 | 0.01
100100.321 | 133.7 | 1591.5 | 0.01
100100.25 | 134 | 1591.5 | 0.01
100100.34 | 134 | 1591.5 | 0.01
100100.337 | 134.3 | 1591.5 | 0.01
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
100100.3281 | 0.044279 | 134.145 | 0.43183
end | of | the | test
Capacitor description:
2020-10-23 | 13:01
Capacitor description: | UB1000 measurement @23/10/2020 19.99°C, 1022.80 mbar, 51.5%RH
2020-10-23 | 13:03
1000.1506 | 0.10962 | 1591.5 | 1.5
1000.1506 | 0.10953 | 1591.5 | 1.5
1000.1506 | 0.10951 | 1591.5 | 1.5
1000.1506 | 0.10954 | 1591.5 | 1.5
1000.1506 | 0.10954 | 1591.5 | 1.5
1000.1506 | 0.10948 | 1591.5 | 1.5
1000.1506 | 0.10954 | 1591.5 | 1.5
1000.1506 | 0.10954 | 1591.5 | 1.5
1000.1506 | 0.10947 | 1591.5 | 1.5
1000.1506 | 0.1095 | 1591.5 | 1.5
1000.1506 | 0.10951 | 1591.5 | 1.5
1000.1506 | 0.10945 | 1591.5 | 1.5
1000.1506 | 0.10945 | 1591.5 | 1.5
1000.1506 | 0.10951 | 1591.5 | 1.5
1000.1506 | 0.10942 | 1591.5 | 1.5
1000.1506 | 0.10945 | 1591.5 | 1.5
1000.1506 | 0.10948 | 1591.5 | 1.5
1000.1506 | 0.10951 | 1591.5 | 1.5
1000.1506 | 0.10945 | 1591.5 | 1.5
1000.1506 | 0.10941 | 1591.5 | 1.5
results | of | mean & | std
Caps Mean | Caps Std | Loss Mean | Loss Std
1000.1506 | 5.3852e-06 | 0.1095 | 4.8936e-05
end | of | the | test

## AH2700_cal
Start with internal cal.
AH2700 will be calibrated 'on the day' against known reference capacitors.
Use these results to have a correction factor adequate for this calibration.
[6 empty] | subtract zero |  |  |  | Value from capacitance scale, out.csv |  |  | measured | NMIA
| Reference | C | sdC | G | sdG | C | sdC | G | sdG | C | mean | stdev | tan delta | G | tan delta
|  | pF | pF | nS | nS | pF | pF | nS | nS | pF | -2.483 | 0.13805 |  |  |  | correction
| zero | -1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06 | 0 | 1.9233e-07 | 0 | 2.7294e-06 |  | ppm correction |  | mrad | nS | mrad | mrad
| AH111A | 10 | 1.8e-07 | -0.00012269 | 2.05e-06 | 10 | 2.256e-07 | -0.00014989 | 2.8156e-06 | 9.9999 | -2.4644 |  | -0.0014989 | 0.00018 | 0.0018 | 0.0032989
| AH111B | 10 | 2.65e-07 | -6.61e-05 | 2.22e-06 | 10 | 2.9786e-07 | -9.33e-05 | 2.9416e-06 | 10 | -2.4315 |  | -0.000933 | 0.00024 | 0.0024 | 0.003333
| AH111C | 99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06 | 99.9998 | 2.1693e-07 | 0.0016097 | 2.4943e-06 | 99.9996 | -2.3617 |  | 0.0016097 | 0.0019 | 0.0019 | 0.0002903
| AH111D | 99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06 | 99.9998 | 2.7412e-07 | 0.0020538 | 2.8595e-06 | 99.9996 | -2.3869 |  | 0.0020538 | 0.0023 | 0.0023 | 0.00024625
| GR1000A | 1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05 | 1000.0317 | 6.4014e-06 | 0.17933 | 5.4534e-05 | 1000.0289 | -2.7431 |  | 0.017932
| GR1000B | 1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05 | 1000.0164 | 5.9216e-06 | 0.18442 | 5.7932e-05 | 1000.0139 | -2.5101 |  | 0.018442
2020-10-12 | UB1000 | 1000.1471 | 5.89e-06 | 0.142 | 9.76e-05 | 1000.1471 | 5.8916e-06 | 0.14197 | 9.7619e-05 | 1000.1444 | -2.6266 |  | 0.014195
2020-10-16 | UB1000 | 1000.1471 | 1.4353e-05 | 0.10825 | 0.0001556 | 1000.1471 | 1.4353e-05 | 0.10822 | 0.00015561 | 1000.1444 | -2.6266 |  | 0.01082
2020-10-19 | UB1000 | 1000.1393 | 7.4833e-06 | 0.10425 | 4.8629e-05 | 1000.1393 | 7.4846e-06 | 0.10422 | 4.8667e-05 | 1000.1367 | -2.6266 |  | 0.010421
2020-10-20 | UB1000 | 1000.1439 | 7.4162e-06 | 0.13882 | 5.3703e-05 | 1000.1439 | 7.4174e-06 | 0.13879 | 5.3738e-05 | 1000.1413 | -2.6266 |  | 0.013877
2020-10-23 | UB1000 | 1000.1511 | 3e-06 | 0.10824 | 5.0326e-05 | 1000.1511 | 3.0031e-06 | 0.10821 | 5.0363e-05 | 1000.1485 | -2.6266 |  | 0.010819
2020-10-23 | UB1000 | 1000.1506 | 5.3852e-06 | 0.1095 | 4.8936e-05 | 1000.1506 | 5.3869e-06 | 0.10947 | 4.8974e-05 | 1000.148 | -2.6266 |  | 0.010945
stdev
ppm
Average | UB1000 | 1000.1465 | 0.0044079 | 0.11851 | 0.017085 | 1000.1465 | 0.0044079 | 0.11848 | 0.017085 | 1000.1439 | -2.6266 | 4.4072 | 0.011846

## pyAH2700readings
Summary of capacitance box measurements extracted from pasted csv file in AH2700_data worksheet.
subtract from relevant zero
| Setting | C | sdC | G | sdG |  | Corrected |  | C | G | tan delta
Row |  | pF | pF | nS | nS |  | pF |  | pF | nS
235 | main_zero | 49.2216 | 0.00043741 | 1.3317 | 0.0053142 |  | 49.2215 |  | 49.2215 | 1.3317 | 0.0027056
261 | E1 | 59.6543 | 0.00014858 | 1.3377 | 0.0022088 |  | 59.6542 |  | 10.4327 | 0.0059616 | 5.7143e-05
287 | E2 | 70.2234 | 6.02e-05 | 1.3687 | 0.0020591 |  | 70.2232 |  | 21.0017 | 0.036932 | 0.00017585
313 | E3 | 80.8288 | 0.00048247 | 1.3481 | 0.0052107 |  | 80.8286 |  | 31.6071 | 0.01633 | 5.1665e-05
339 | E4 | 90.9238 | 0.00021909 | 1.35 | 0.00079946 |  | 90.9235 |  | 41.702 | 0.018306 | 4.3898e-05
365 | E5 | 100.4774 | 0.00036577 | 1.3441 | 0.0038393 |  | 100.4771 |  | 51.2556 | 0.012348 | 2.4091e-05
391 | E6 | 110.6912 | 0.00050256 | 1.3455 | 0.00079577 |  | 110.6909 |  | 61.4694 | 0.013767 | 2.2396e-05
417 | E7 | 120.5081 | 0.00029987 | 1.3463 | 0.00018653 |  | 120.5078 |  | 71.2863 | 0.014561 | 2.0426e-05
443 | E8 | 129.974 | 5.59e-05 | 1.3525 | 0.00082738 |  | 129.9736 |  | 80.7521 | 0.020738 | 2.5682e-05
469 | E9 | 139.7386 | 0.0005509 | 1.3521 | 0.0006845 |  | 139.7382 |  | 90.5167 | 0.020384 | 2.252e-05
495 | E10 | 149.5767 | 0.00025146 | 1.3555 | 0.00075452 |  | 149.5763 |  | 100.3548 | 0.023795 | 2.3711e-05
521 | D1 | 145.8887 | 0.00046896 | 1.6176 | 0.015335 |  | 145.8884 |  | 96.6669 | 0.28585 | 0.0002957
547 | D2 | 244.9349 | 0.00010611 | 1.4339 | 0.0016017 |  | 244.9343 |  | 195.7128 | 0.10215 | 5.2193e-05
573 | D3 | 341.4798 | 0.00032198 | 1.5154 | 0.0029831 |  | 341.479 |  | 292.2575 | 0.18372 | 6.2861e-05
599 | D4 | 441.2625 | 0.0001467 | 1.5405 | 0.0012842 |  | 441.2614 |  | 392.0399 | 0.20879 | 5.3258e-05
625 | D5 | 537.7661 | 6.53e-05 | 1.6209 | 0.0006773 |  | 537.7648 |  | 488.5433 | 0.28917 | 5.9191e-05
651 | D6 | 640.7041 | 0.00010485 | 1.6665 | 0.00040726 |  | 640.7025 |  | 591.481 | 0.33476 | 5.6597e-05
677 | D7 | 739.8078 | 0.00025568 | 1.687 | 0.0016552 |  | 739.806 |  | 690.5845 | 0.35524 | 5.1441e-05
703 | D8 | 836.3521 | 0.00010751 | 1.7676 | 0.00037086 |  | 836.35 |  | 787.1285 | 0.43591 | 5.5379e-05
729 | D9 | 936.1413 | 0.00025167 | 1.7886 | 0.00083413 |  | 936.1389 |  | 886.9174 | 0.45689 | 5.1515e-05
755 | D10 | 1032.4678 | 0.00013984 | 1.8805 | 0.0014275 |  | 1032.4652 |  | 983.2437 | 0.54878 | 5.5814e-05
781 | C1 | 1044.2431 | 0.00020482 | 1.8803 | 0.00085081 |  | 1044.2405 |  | 995.019 | 0.54861 | 5.5136e-05
807 | C2 | 2046.623 | 7.55e-05 | 2.2018 | 0.00016795 |  | 2046.6179 |  | 1997.3964 | 0.87012 | 4.3563e-05
833 | C3 | 3041.3243 | 0.0001986 | 2.7758 | 0.00070788 |  | 3041.3167 |  | 2992.0952 | 1.4441 | 4.8264e-05
859 | C4 | 4043.9175 | 0.00020302 | 3.2162 | 0.00073231 |  | 4043.9074 |  | 3994.6859 | 1.8845 | 4.7174e-05
885 | C5 | 5047.2787 | 0.00034511 | 3.7846 | 0.00098189 |  | 5047.2662 |  | 4998.0447 | 2.4528 | 4.9076e-05
911 | C6 | 6041.4577 | 0.00036621 | 4.3636 | 0.0010893 |  | 6041.4427 |  | 5992.2212 | 3.0319 | 5.0597e-05
937 | C7 | 7043.9733 | 0.00037834 | 4.8052 | 0.0016222 |  | 7043.9558 |  | 6994.7343 | 3.4735 | 4.9659e-05
963 | C8 | 8037.9916 | 0.00041509 | 5.5054 | 0.0081515 |  | 8037.9717 |  | 7988.7502 | 4.1737 | 5.2245e-05
989 | C9 | 9040.2665 | 0.00043352 | 5.9494 | 0.0042009 |  | 9040.244 |  | 8991.0225 | 4.6177 | 5.1359e-05
1015 | C10 | 10034.2269 | 0.00082721 | 6.7192 | 0.0057671 |  | 10034.202 |  | 9984.9805 | 5.3875 | 5.3956e-05
1041 | B1 | 10060.9476 | 0.0008394 | 6.7161 | 0.0064535 |  | 10060.9226 |  | 10011.7011 | 5.3843 | 5.378e-05
1067 | B2 | 20052.3191 | 0.0013909 | 14.1209 | 0.0099412 |  | 20052.2693 |  | 20003.0478 | 12.7891 | 6.3936e-05
1093 | B3 | 30063.9931 | 0.0056836 | 21.596 | 0.051904 |  | 30063.9184 |  | 30014.6969 | 20.2643 | 6.7515e-05
1119 | B4 | 40066.9257 | 0.00735 | 32.053 | 0.067164 |  | 40066.8262 |  | 40017.6047 | 30.7213 | 7.6769e-05
1145 | B5 | 50072.9504 | 0.007224 | 44.6475 | 0.06774 |  | 50072.8261 |  | 50023.6046 | 43.3158 | 8.6591e-05
1171 | B6 | 60084.2973 | 0.009408 | 58.51 | 0.099499 |  | 60084.1481 |  | 60034.9266 | 57.1783 | 9.5242e-05
1197 | B7 | 70075.9355 | 0.033996 | 73.07 | 0.27586 |  | 70075.7615 |  | 70026.5399 | 71.7383 | 0.00010244
1223 | B8 | 80087.3182 | 0.04288 | 91.56 | 0.36387 |  | 80087.1193 |  | 80037.8978 | 90.2283 | 0.00011273
1249 | B9 | 90089.9263 | 0.045232 | 110.885 | 0.39532 |  | 90089.7026 |  | 90040.4811 | 109.5533 | 0.00012167
1275 | B10 | 100101.3839 | 0.033297 | 133.075 | 0.55486 |  | 100101.1354 |  | 100051.9138 | 131.7433 | 0.00013167
1303 | A1 | 100015.947 | 0.065305 | 135.545 | 0.54173 |  | 100015.6987 |  | 99966.4772 | 134.2133 | 0.00013426
1329 | A2 | 200421.712 | 0.11548 | 482.78 | 0.90089 |  | 200421.2144 |  | 200371.9928 | 481.4483 | 0.00024028
1355 | A3 | 300392.136 | 0.41094 | 936.55 | 3.9683 |  | 300391.3901 |  | 300342.1686 | 935.2183 | 0.00031138
1381 | A4 | 400791.768 | 0.64947 | 1670.35 | 4.6505 |  | 400790.7729 |  | 400741.5513 | 1669.0183 | 0.00041648
1407 | A5 | 500612.38 | 0.66378 | 2631.55 | 5.4907 |  | 500611.137 |  | 500561.9155 | 2630.2183 | 0.00052545
1433 | A6 | 600591.812 | 0.74936 | 3674.25 | 8.2757 |  | 600590.3208 |  | 600541.0992 | 3672.9183 | 0.0006116
1459 | A7 | 701006.25 | 2.4309 | 4883.35 | 36.3803 |  | 701004.5094 |  | 700955.2879 | 4882.0183 | 0.00069648
1485 | A8 | 800991.225 | 4.1986 | 6379 | 37.8021 |  | 800989.2362 |  | 800940.0147 | 6377.6683 | 0.00079627
1511 | A9 | 901403.07 | 4.2023 | 8020.5 | 43.872 |  | 901400.8319 |  | 901351.6103 | 8019.1683 | 0.00088968
1537 | A10 | 1001391.815 | 4.8028 | 9972 | 40.5709 |  | 1001389.3286 |  | 1001340.1071 | 9970.6683 | 0.00099573
1563 |  | 1000.1471 | 1.4353e-05 | 0.10825 | 0.0001556 |  | 1000.1446 |  | 1000.1394 | 0.10813 | 1.0811e-05 | % change from above |  | UB1000 measure After Tpye 1412BC sn: 4556 Calibration 20.10deg, 1017.52 mbar, 39.6%RH
1589 |  | 0.0051919 | 1.985e-06 | 0.00011836 | 1.7328e-06 |  | 0.0051919 |  | 0 | 0 | #DIV/0! |  |  | U_connector Zero Point Check, 19.87 degree, 39.7% RH, 1017.19mbar
1615 |  | 1000.1393 | 7.4833e-06 | 0.10425 | 4.8629e-05 |  | 1000.1368 |  | 1000.1317 | 0.10413 | 1.0412e-05 |  |  | UB1000 measurement @19/10/2020 20.27deg, 1006.12 mbar, 58.2%RH
1641 |  | 1000.1439 | 7.4162e-06 | 0.13882 | 5.3703e-05 |  | 1000.1414 |  | 1000.1362 | 0.1387 | 1.3868e-05 |  |  | UB1000 measurement @20/10/2020 20.44deg, 1012.36 mbar, 53.2%RH
1667 |  | 1000.1511 | 3e-06 | 0.10824 | 5.0326e-05 |  | 1000.1486 |  | 1000.1435 | 0.10812 | 1.081e-05 |  |  | UB1000 measurement @23/10/2020 20.12deg, 1012.18 mbar, 50.3%RH
1693 |  | 49.275 | 9.6652e-06 | 1.1967 | 0.00034907 |  | 49.2749 |  | 0 | 0 | #DIV/0! |  |  | GR sn:4556 100pF Dail 0 20.05°C
1719 | D1 | 145.9575 | 2.8676e-05 | 1.2744 | 0.00038384 |  | 145.9571 |  | 96.6823 | 0.077689 | 8.0355e-05 | 0.015919 | -72.8213 | GR sn:4556 100pF Dail 1 20.02°C
1745 | D2 | 244.993 | 0.0061638 | 1.3353 | 0.063117 |  | 244.9924 |  | 195.7175 | 0.13856 | 7.0797e-05 | 0.0024029 | 35.6488 | GR sn:4556 100pF Dail 2 20.01°C
1771 | D3 | 341.5131 | 0.0006186 | 1.5208 | 0.00031844 |  | 341.5123 |  | 292.2374 | 0.32406 | 0.00011089 | -0.0068744 | 76.3921 | GR sn:4556 100pF Dail 3 19.99°C
1797 | D4 | 441.3124 | 0.00022718 | 1.5516 | 0.00032693 |  | 441.3113 |  | 392.0364 | 0.3549 | 9.0528e-05 | -0.00088583 | 69.9773 | GR sn:4556 100pF Dail 4 20.00°C
1823 | D5 | 537.8138 | 0.00023011 | 1.6309 | 0.00050685 |  | 537.8125 |  | 488.5376 | 0.43418 | 8.8873e-05 | -0.0011625 | 50.1443 | GR sn:4556 100pF Dail 5 20.00°C
1849 | D6 | 640.7431 | 0.00028835 | 1.6676 | 0.00024606 |  | 640.7415 |  | 591.4666 | 0.4709 | 7.9615e-05 | -0.0024363 | 40.6673 | GR sn:4556 100pF Dail 6 19.99°C
1875 | D7 | 739.8531 | 0.00031824 | 1.692 | 0.00049595 |  | 739.8512 |  | 690.5764 | 0.49529 | 7.1721e-05 | -0.0011766 | 39.423 | GR sn:4556 100pF Dail 7 20.00°C
1901 | D8 | 836.3869 | 0.00061833 | 1.7735 | 0.00064987 |  | 836.3848 |  | 787.1099 | 0.57681 | 7.3282e-05 | -0.0023578 | 32.3239 | GR sn:4556 100pF Dail 8 20.01°C
1927 | D9 | 936.2005 | 0.00039721 | 1.8016 | 0.00046248 |  | 936.1982 |  | 886.9233 | 0.60487 | 6.8199e-05 | 0.00066579 | 32.3887 | GR sn:4556 100pF Dail 9 20.01°C
1953 | D10 | 1032.503 | 0.00031839 | 1.877 | 0.00026451 |  | 1032.5005 |  | 983.2256 | 0.68027 | 6.9187e-05 | -0.0018437 | 23.9591 | GR sn:4556 100pF Dail X 20.01°C
1979 |  | 49.2764 | 0.00026341 | 1.3448 | 0.00027408 |  | 49.2762 |  | 0.0013574 | 0.14803 | 10.9053 |  |  | GR sn:4556 0.1uF Dail 0 20.01°C
2005 |  | 49.2765 | 0.00023971 | 1.3447 | 0.00038761 |  | 49.2764 |  | 0.0015143 | 0.14795 | 9.7701 |  |  | GR sn:4556 0.01UuF Dail 20.02°C
2031 | B1 | 10061.0088 | 0.0021906 | 6.7369 | 0.0092731 |  | 10060.9838 |  | 10011.7089 | 5.5402 | 5.5337e-05 | 7.8115e-05 | 2.8946 | GR sn:4556 0.01uF Dail 1 20.02°C (last reading is the zero dial, and the one before it range written wrong)
2057 | B2 | 20051.521 | 0.0030431 | 14.1196 | 0.0092762 |  | 20051.4712 |  | 20002.1963 | 12.9228 | 6.4607e-05 | -0.0042568 | 1.0454 | GR sn:4556 0.01uF Dail 2 20.01°C
2083 | B3 | 30063.1534 | 0.0064027 | 21.5225 | 0.041698 |  | 30063.0787 |  | 30013.8038 | 20.3258 | 6.7721e-05 | -0.0029755 | 0.3035 | GR sn:4556 0.01uF Dail 3 20.01°C
2109 | B4 | 40065.6506 | 0.0057476 | 32.108 | 0.068891 |  | 40065.5511 |  | 40016.2762 | 30.9113 | 7.7247e-05 | -0.0033198 | 0.61847 | GR sn:4556 0.01uF Dail 4 20.01°C
2135 | B5 | 50073.8499 | 0.010712 | 44.7415 | 0.072338 |  | 50073.7256 |  | 50024.4507 | 43.5448 | 8.7047e-05 | 0.0016914 | 0.52868 | GR sn:4556 0.01uF Dail 5 20.02°C
2161 | B5 | 50073.8144 | 0.007845 | 44.764 | 0.070314 |  | 50073.69 |  | 50024.4152 | 43.5673 | 8.7092e-05 |  |  | GR sn:4556 0.01uF Dail 6 20.02°C
2187 | B6 | 60085.0375 | 0.011796 | 58.705 | 0.10712 |  | 60084.8883 |  | 60035.6134 | 57.5083 | 9.579e-05 | 0.001144 | 0.57715 | GR sn:4556 0.01uF Dail 6 20.02°C
2213 | B7 | 70075.7392 | 0.0202 | 73.575 | 0.20706 |  | 70075.5652 |  | 70026.2903 | 72.3783 | 0.00010336 | -0.00035647 | 0.89214 | GR sn:4556 0.01uF Dail 7 20.02°C
2239 | B8 | 80086.9693 | 0.051408 | 92.215 | 0.3678 |  | 80086.7704 |  | 80037.4956 | 91.0183 | 0.00011372 | -0.00050254 | 0.87556 | GR sn:4556 0.01uF Dail 8 20.03°C
2265 | B9 | 90089.0666 | 0.034168 | 111.62 | 0.38158 |  | 90088.8429 |  | 90039.568 | 110.4233 | 0.00012264 | -0.0010141 | 0.79414 | GR sn:4556 0.01uF Dail 9 20.04°C
2291 | B10 | 100100.3281 | 0.044279 | 134.145 | 0.43183 |  | 100100.0796 |  | 100050.8047 | 132.9483 | 0.00013288 | -0.0011086 | 0.91466 | GR sn:4556 0.01uF Dail X 20.04°C
2319 |  | 1000.1506 | 5.3852e-06 | 0.1095 | 4.8936e-05 |  | 1000.1481 |  | 1000.1429 | 0.10938 | 1.0936e-05 |  |  | UB1000 measurement @23/10/2020 19.99°C, 1022.80 mbar, 51.5%RH
Readings in KJ Lab Diary2, p.123, 27/10/20
manual repeatability check on D1, dial switched between each set of readings.
[6 empty] | delta C | delta G | tand
| 49.3155 | 1.3345 | 145.9686 | 1.4156 |  | 96.6531 | 0.081059 | 0.083866
| 49.3042 | 1.3358 | 145.9696 | 1.4481 |  | 96.6654 | 0.11228 | 0.11615
| 49.3118 | 1.3347 | 145.9803 | 1.4128 |  | 96.6684 | 0.078093 | 0.080784
| 49.2995 | 1.3292 | 145.9777 | 1.4073 |  | 96.6782 | 0.07803 | 0.080711
| 49.2997 | 1.3297 | 145.9745 | 1.4141 |  | 96.6748 | 0.084424 | 0.087328
[7 empty] | 0.086777 | 0.089768
[7 empty] | 0.014496 | 0.014998

## report
| Selected | [7 empty] | Universal Bridge | [7 empty] | AH2700
| Table 2 | [23 empty] | En | En
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | C | tand
E1 | 1pF | 10 | 10.4646 | 1.2853 | 2.5705 | 0.057143 | 0.48049 | 2.0217 | 1pF | 10 | 10.4646 | 1.2853 | 2.5705 | 2.3509 | 0.48049 | 2.0217 | 1pF | 10 | 10.4327 | 1.2853 | 2.5705 | 0.057143 | 0.48049 | 2.0217 | 0.017575 | 3.3756
E2 | DIV | 20 | 21.2793 | 1.2853 | 2.5704 | 0.17585 | 0.21955 | 2.0131 | DIV | 20 | 21.2793 | 1.2853 | 2.5704 | 1.3911 | 0.21955 | 2.0131 | DIV | 20 | 21.0017 | 1.2853 | 2.5704 | 0.17585 | 0.21955 | 2.0131 | 0.15268 | 3.914
E3 |  | 30 | 31.4978 | 1.2853 | 2.5704 | 0.051665 | 0.14576 | 2.0137 |  | 30 | 31.4978 | 1.2853 | 2.5704 | 0.75247 | 0.14576 | 2.0137 |  | 30 | 31.6071 | 1.2853 | 2.5704 | 0.051665 | 0.14576 | 2.0137 | 0.060114 | 3.3998
E4 |  | 40 | 41.7954 | 1.2853 | 2.5704 | 0.043898 | 0.11131 | 2.0119 |  | 40 | 41.7954 | 1.2853 | 2.5704 | 0.54794 | 0.11131 | 2.0119 |  | 40 | 41.702 | 1.2853 | 2.5704 | 0.043898 | 0.11131 | 2.0119 | 0.051368 | 3.2018
E5 |  | 50 | 51.6939 | 1.2853 | 2.5703 | 0.024091 | 0.091937 | 2.01 |  | 50 | 51.6939 | 1.2853 | 2.5703 | 0.46236 | 0.091937 | 2.01 |  | 50 | 51.2556 | 1.2853 | 2.5703 | 0.024091 | 0.091937 | 2.01 | 0.24114 | 3.3708
E6 |  | 60 | 61.5424 | 1.2853 | 2.5703 | 0.022396 | 0.079263 | 2.0091 |  | 60 | 61.5424 | 1.2853 | 2.5703 | 0.39487 | 0.079263 | 2.0091 |  | 60 | 61.4694 | 1.2853 | 2.5703 | 0.022396 | 0.079263 | 2.0091 | 0.040192 | 3.3229
E7 |  | 70 | 70.9769 | 1.2853 | 2.5703 | 0.020426 | 0.070715 | 2.0094 |  | 70 | 70.9769 | 1.2853 | 2.5703 | 0.31561 | 0.070715 | 2.0094 |  | 70 | 71.2863 | 1.2853 | 2.5703 | 0.020426 | 0.070715 | 2.0094 | 0.17021 | 2.9517
E8 |  | 80 | 80.6603 | 1.2853 | 2.5702 | 0.025682 | 0.064311 | 2.0111 |  | 80 | 80.6603 | 1.2853 | 2.5702 | 0.27276 | 0.064311 | 2.0111 |  | 80 | 80.7521 | 1.2853 | 2.5702 | 0.025682 | 0.064311 | 2.0111 | 0.050497 | 2.7167
E9 |  | 90 | 90.6839 | 1.2853 | 2.5702 | 0.02252 | 0.059353 | 2.0144 |  | 90 | 90.6839 | 1.2853 | 2.5702 | 0.2371 | 0.059353 | 2.0144 |  | 90 | 90.5167 | 1.2853 | 2.5702 | 0.02252 | 0.059353 | 2.0144 | 0.09197 | 2.5564
E10 |  | '100' | 100.0583 | 1.2853 | 2.5701 | 0.023711 | 0.055793 | 2.0186 |  | '100' | 100.0583 | 1.2853 | 2.5701 | 0.20889 | 0.055793 | 2.0186 |  | '100' | 100.3548 | 1.2853 | 2.5701 | 0.023711 | 0.055793 | 2.0186 | 0.16312 | 2.3469
Table 3
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
D1 | 100pF | 1 | 96.6318 | 0.016686 | 2.0391 | 0.080355 | 0.056947 | 2.0171 | 100pF | 1 | 96.6318 | 0.016686 | 2.0391 | 0.2957 | 0.056947 | 2.0171 | 100pF | 1 | 96.6669 | 0.016686 | 2.0391 | 0.080355 | 0.056947 | 2.0171 | 1.4866 | 2.674
D2 | STEPS | 2 | 195.735 | 0.027318 | 2.0475 | 0.052193 | 0.042311 | 2.0883 | STEPS | 2 | 195.735 | 0.027318 | 2.0475 | 0.02197 | 0.042311 | 2.0883 | STEPS | 2 | 195.7128 | 0.027318 | 2.0475 | 0.052193 | 0.042311 | 2.0883 | 0.57437 | 0.50509
D3 |  | 3 | 292.2347 | 0.037891 | 2.0483 | 0.062861 | 0.039614 | 2.1426 |  | 3 | 292.2347 | 0.037891 | 2.0483 | 0.025666 | 0.039614 | 2.1426 |  | 3 | 292.2575 | 0.037891 | 2.0483 | 0.062861 | 0.039614 | 2.1426 | 0.42425 | 0.66393
D4 |  | 4 | 392.049 | 0.049075 | 2.0485 | 0.053258 | 0.038703 | 2.1738 |  | 4 | 392.049 | 0.049075 | 2.0485 | 0.0071424 | 0.038703 | 2.1738 |  | 4 | 392.0399 | 0.049075 | 2.0485 | 0.053258 | 0.038703 | 2.1738 | 0.13139 | 0.84254
D5 |  | 5 | 488.5508 | 0.059937 | 2.0483 | 0.059191 | 0.03832 | 2.1909 |  | 5 | 488.5508 | 0.059937 | 2.0483 | 0.0092114 | 0.03832 | 2.1909 |  | 5 | 488.5433 | 0.059937 | 2.0483 | 0.059191 | 0.03832 | 2.1909 | 0.088605 | 0.92225
D6 |  | 6 | 591.5245 | 0.071894 | 2.0488 | 0.056597 | 0.038108 | 2.2017 |  | 6 | 591.5245 | 0.071894 | 2.0488 | -0.0062554 | 0.038108 | 2.2017 |  | 6 | 591.481 | 0.071894 | 2.0488 | 0.056597 | 0.038108 | 2.2017 | 0.42825 | 1.1662
D7 |  | 7 | 690.6187 | 0.082752 | 2.0478 | 0.051441 | 0.03799 | 2.2083 |  | 7 | 690.6187 | 0.082752 | 2.0478 | 0.0099916 | 0.03799 | 2.2083 |  | 7 | 690.5845 | 0.082752 | 2.0478 | 0.051441 | 0.03799 | 2.2083 | 0.2926 | 0.7715
D8 |  | 8 | 787.1525 | 0.093669 | 2.0476 | 0.055379 | 0.037916 | 2.2126 |  | 8 | 787.1525 | 0.093669 | 2.0476 | 0.027061 | 0.037916 | 2.2126 |  | 8 | 787.1285 | 0.093669 | 2.0476 | 0.055379 | 0.037916 | 2.2126 | 0.18116 | 0.52812
D9 |  | 9 | 886.9588 | 0.10496 | 2.0474 | 0.051515 | 0.037864 | 2.2158 |  | 9 | 886.9588 | 0.10496 | 2.0474 | 0.037321 | 0.037864 | 2.2158 |  | 9 | 886.9174 | 0.10496 | 2.0474 | 0.051515 | 0.037864 | 2.2158 | 0.27864 | 0.26507
D10 |  | 10 | 983.2735 | 0.11587 | 2.0473 | 0.055814 | 0.037829 | 2.218 |  | 10 | 983.2735 | 0.11587 | 2.0473 | 0.044649 | 0.037829 | 2.218 |  | 10 | 983.2437 | 0.11587 | 2.0473 | 0.055814 | 0.037829 | 2.218 | 0.18184 | 0.20869
Table 4
| Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
C1 | 1000pF | 1 | 0.9951 | 0.00012363 | 2.0264 | 0.040197 | 0.055932 | 2.0185 | 1000pF | 1 | 0.9951 | 0.00012363 | 2.0264 | 0.040197 | 0.055932 | 2.0185 | 1000pF | 1 | 0.99502 | 0.00012363 | 2.0264 | 0.055136 | 0.055932 | 2.0185 | 0.46113 | 0.18887
C2 | STEPS | 2 | 1.9975 | 0.00023234 | 2.0387 | 0.017021 | 0.042112 | 2.0912 | STEPS | 2 | 1.9975 | 0.00023234 | 2.0387 | 0.017021 | 0.042112 | 2.0912 | STEPS | 2 | 1.9974 | 0.00023234 | 2.0387 | 0.043563 | 0.042112 | 2.0912 | 0.2544 | 0.44565
C3 |  | 3 | 2.9923 | 0.00034299 | 2.0417 | 0.012365 | 0.039517 | 2.1455 |  | 3 | 2.9923 | 0.00034299 | 2.0417 | 0.012365 | 0.039517 | 2.1455 |  | 3 | 2.9921 | 0.00034299 | 2.0417 | 0.048264 | 0.039517 | 2.1455 | 0.38003 | 0.64237
C4 |  | 4 | 3.9949 | 0.00045705 | 2.0434 | 0.0057574 | 0.038664 | 2.1754 |  | 4 | 3.9949 | 0.00045705 | 2.0434 | 0.0057574 | 0.038664 | 2.1754 |  | 4 | 3.9947 | 0.00045705 | 2.0434 | 0.047174 | 0.038664 | 2.1754 | 0.26914 | 0.75745
C5 |  | 5 | 4.9983 | 0.00056999 | 2.0439 | -0.0014005 | 0.038292 | 2.1922 |  | 5 | 4.9983 | 0.00056999 | 2.0439 | -0.0014005 | 0.038292 | 2.1922 |  | 5 | 4.998 | 0.00056999 | 2.0439 | 0.049076 | 0.038292 | 2.1922 | 0.31711 | 0.93211
C6 |  | 6 | 5.9926 | 0.00068202 | 2.0441 | -0.0093448 | 0.038099 | 2.2022 |  | 6 | 5.9926 | 0.00068202 | 2.0441 | -0.0093448 | 0.038099 | 2.2022 |  | 6 | 5.9922 | 0.00068202 | 2.0441 | 0.050597 | 0.038099 | 2.2022 | 0.37185 | 1.1125
C7 |  | 7 | 6.995 | 0.00079164 | 2.0436 | 0.0020014 | 0.037985 | 2.2085 |  | 7 | 6.995 | 0.00079164 | 2.0436 | 0.0020014 | 0.037985 | 2.2085 |  | 7 | 6.9947 | 0.00079164 | 2.0436 | 0.049659 | 0.037985 | 2.2085 | 0.24637 | 0.88716
C8 |  | 8 | 7.9891 | 0.00090328 | 2.0437 | 0.014395 | 0.037914 | 2.2128 |  | 8 | 7.9891 | 0.00090328 | 2.0437 | 0.014395 | 0.037914 | 2.2128 |  | 8 | 7.9888 | 0.00090328 | 2.0437 | 0.052245 | 0.037914 | 2.2128 | 0.2658 | 0.70593
C9 |  | 9 | 8.9914 | 0.0010077 | 2.0425 | 0.022911 | 0.037865 | 2.2157 |  | 9 | 8.9914 | 0.0010077 | 2.0425 | 0.022911 | 0.037865 | 2.2157 |  | 9 | 8.991 | 0.0010077 | 2.0425 | 0.051359 | 0.037865 | 2.2157 | 0.24386 | 0.53126
C10 |  | 10 | 9.9855 | 0.0010896 | 2.0388 | 0.037053 | 0.037832 | 2.2178 |  | 10 | 9.9855 | 0.0010896 | 2.0388 | 0.037053 | 0.037832 | 2.2178 |  | 10 | 9.985 | 0.0010896 | 2.0388 | 0.053956 | 0.037832 | 2.2178 | 0.34982 | 0.31592
Table 5
| Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (nF) | Expanded Uncertainty (nF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
B1 | 0.01µF | 1 | 10.0127 | 0.0012216 | 2.0259 | 0.036953 | 0.055733 | 2.0188 | 0.01µF | 1 | 10.0127 | 0.0012216 | 2.0259 | 0.036953 | 0.055733 | 2.0188 | 0.01µF | 1 | 10.0117 | 0.0012216 | 2.0259 | 0.05378 | 0.055733 | 2.0188 | 0.57935 | 0.2135
B2 | STEPS | 2 | 20.0044 | 0.0023194 | 2.0396 | 0.028494 | 0.042136 | 2.0907 | STEPS | 2 | 20.0044 | 0.0023194 | 2.0396 | 0.028494 | 0.042136 | 2.0907 | STEPS | 2 | 20.003 | 0.0023194 | 2.0396 | 0.063936 | 0.042136 | 2.0907 | 0.41289 | 0.59477
B3 |  | 3 | 30.0168 | 0.0034806 | 2.0444 | 0.019989 | 0.039587 | 2.1432 |  | 3 | 30.0168 | 0.0034806 | 2.0444 | 0.019989 | 0.039587 | 2.1432 |  | 3 | 30.0147 | 0.0034806 | 2.0444 | 0.067515 | 0.039587 | 2.1432 | 0.42809 | 0.84892
B4 |  | 4 | 40.0203 | 0.0046032 | 2.0449 | 0.013493 | 0.0388 | 2.1698 |  | 4 | 40.0203 | 0.0046032 | 2.0449 | 0.013493 | 0.0388 | 2.1698 |  | 4 | 40.0176 | 0.0046032 | 2.0449 | 0.076769 | 0.0388 | 2.1698 | 0.41485 | 1.1532
B5 |  | 5 | 50.0294 | 0.005773 | 2.0463 | 0.0093944 | 0.038503 | 2.1823 |  | 5 | 50.0294 | 0.005773 | 2.0463 | 0.0093944 | 0.038503 | 2.1823 |  | 5 | 50.0236 | 0.005773 | 2.0463 | 0.086591 | 0.038503 | 2.1823 | 0.71071 | 1.4177
B6 |  | 6 | 60.0421 | 0.0068891 | 2.0458 | -0.011825 | 0.038401 | 2.1873 |  | 6 | 60.0421 | 0.0068891 | 2.0458 | -0.011825 | 0.038401 | 2.1873 |  | 6 | 60.0349 | 0.0068891 | 2.0458 | 0.095242 | 0.038401 | 2.1873 | 0.73721 | 1.9715
B7 |  | 7 | 70.0331 | 0.0080311 | 2.046 | 0.012708 | 0.038397 | 2.188 |  | 7 | 70.0331 | 0.0080311 | 2.046 | 0.012708 | 0.038397 | 2.188 |  | 7 | 70.0265 | 0.0080311 | 2.046 | 0.10244 | 0.038397 | 2.188 | 0.57839 | 1.6525
B8 |  | 8 | 80.0451 | 0.0092201 | 2.0469 | 0.022362 | 0.038456 | 2.1861 |  | 8 | 80.0451 | 0.0092201 | 2.0469 | 0.022362 | 0.038456 | 2.1861 |  | 8 | 80.0379 | 0.0092201 | 2.0469 | 0.11273 | 0.038456 | 2.1861 | 0.5532 | 1.6617
B9 |  | 9 | 90.0482 | 0.01032 | 2.0462 | 0.035425 | 0.038561 | 2.1826 |  | 9 | 90.0482 | 0.01032 | 2.0462 | 0.035425 | 0.038561 | 2.1826 |  | 9 | 90.0405 | 0.01032 | 2.0462 | 0.12167 | 0.038561 | 2.1826 | 0.52972 | 1.5815
B10 |  | 10 | 100.0601 | 0.01152 | 2.047 | 0.050369 | 0.038703 | 2.178 |  | 10 | 100.0601 | 0.01152 | 2.047 | 0.050369 | 0.038703 | 2.178 |  | 10 | 100.0519 | 0.01152 | 2.047 | 0.13167 | 0.038703 | 2.178 | 0.50332 | 1.4854
Table 6
| Dial Label | Dial Setting | Capacitance Change (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (µF) | Expanded Uncertainty (µF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
A1 | 0.1µF | 1 | 0.099973 | 1.1976e-05 | 2.0231 | 0.04001 | 0.056589 | 2.0154 | 0.1µF | 1 | 0.099973 | 1.1976e-05 | 2.0231 | 0.04001 | 0.056589 | 2.0154 | 0.1µF | 1 | 0.099966 | 1.1976e-05 | 2.0231 | 0.13426 | 0.056589 | 2.0154 | 0.3855 | 1.1777
A2 | STEPS | 2 | 0.20038 | 2.2995e-05 | 2.0381 | 0.044914 | 0.046296 | 2.0596 | STEPS | 2 | 0.20038 | 2.2995e-05 | 2.0381 | 0.044914 | 0.046296 | 2.0596 | STEPS | 2 | 0.20037 | 2.2995e-05 | 2.0381 | 0.24028 | 0.046296 | 2.0596 | 0.30931 | 2.9839
A3 |  | 3 | 0.30035 | 3.4163e-05 | 2.0415 | 0.03962 | 0.049721 | 2.0927 |  | 3 | 0.30035 | 3.4163e-05 | 2.0415 | 0.03962 | 0.049721 | 2.0927 |  | 3 | 0.30034 | 3.4163e-05 | 2.0415 | 0.31138 | 0.049721 | 2.0927 | 0.24609 | 3.8649
A4 |  | 4 | 0.40075 | 4.5439e-05 | 2.0427 | 0.052151 | 0.057201 | 2.1357 |  | 4 | 0.40075 | 4.5439e-05 | 2.0427 | 0.052151 | 0.057201 | 2.1357 |  | 4 | 0.40074 | 4.5439e-05 | 2.0427 | 0.41648 | 0.057201 | 2.1357 | 0.19532 | 4.5038
A5 |  | 5 | 0.50057 | 5.6674e-05 | 2.0433 | 0.06912 | 0.066636 | 2.1832 |  | 5 | 0.50057 | 5.6674e-05 | 2.0433 | 0.06912 | 0.066636 | 2.1832 |  | 5 | 0.50056 | 5.6674e-05 | 2.0433 | 0.52545 | 0.066636 | 2.1832 | 0.12701 | 4.8424
A6 |  | 6 | 0.60057 | 6.794e-05 | 2.0437 | 0.013987 | 0.077184 | 2.2271 |  | 6 | 0.60057 | 6.794e-05 | 2.0437 | 0.013987 | 0.077184 | 2.2271 |  | 6 | 0.60054 | 6.794e-05 | 2.0437 | 0.6116 | 0.077184 | 2.2271 | 0.24983 | 5.4749
A7 |  | 7 | 0.70095 | 7.9259e-05 | 2.0439 | 0.074469 | 0.088402 | 2.2643 |  | 7 | 0.70095 | 7.9259e-05 | 2.0439 | 0.074469 | 0.088402 | 2.2643 |  | 7 | 0.70096 | 7.9259e-05 | 2.0439 | 0.69648 | 0.088402 | 2.2643 | 0.010179 | 4.9753
A8 |  | 8 | 0.80083 | 9.0523e-05 | 2.044 | 0.091904 | 0.099931 | 2.2944 |  | 8 | 0.80083 | 9.0523e-05 | 2.044 | 0.091904 | 0.099931 | 2.2944 |  | 8 | 0.80094 | 9.0523e-05 | 2.044 | 0.79627 | 0.099931 | 2.2944 | 0.88948 | 4.9841
A9 |  | 9 | 0.90132 | 0.00010186 | 2.0441 | 0.11239 | 0.11176 | 2.3187 |  | 9 | 0.90132 | 0.00010186 | 2.0441 | 0.11239 | 0.11176 | 2.3187 |  | 9 | 0.90135 | 0.00010186 | 2.0441 | 0.88968 | 0.11176 | 2.3187 | 0.1903 | 4.918
A10 |  | 10 | 1.0013 | 0.00011314 | 2.0441 | 0.13253 | 0.12367 | 2.3381 |  | 10 | 1.0013 | 0.00011314 | 2.0441 | 0.13253 | 0.12367 | 2.3381 |  | 10 | 1.0013 | 0.00011314 | 2.0441 | 0.99573 | 0.12367 | 2.3381 | 0.25563 | 4.9357
Table 1
| Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor | Dial Label | Dial Setting | Capacitance Change (pF) | Expanded Uncertainty (pF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
main_zero | All dials | 0 | 49.1435 | 1.5423 | 2.5705 | 2.7056 | 0.11021 | 2.0182 | All dials | 0 | 49.1435 | 1.5423 | 2.5705 | 2.1978 | 0.11021 | 2.0182 | All dials | 0 | 49.2215 | 1.5423 | 2.5705 | 2.7056 | 0.11021 | 2.0182 | 0.035756 | 3.2581
Note on AA17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: taken from repeat run and confirmed by manual measurements on 27/10/20

## older
1999 report | 1999-09-23 | S21341 | E73015 | [5 empty] | 2004 report | 2004-09-10 | S21567 | E7140017 | [5 empty] | 2009 report | 2009-12-01 | S21745 | E36040128 | [5 empty] | 2015 report | 2015-04-30 | S21959 | E96040236 | [5 empty] | 2020 report | 2020-10-23
| Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k
|  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty |  |  |  |  | change | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian)
E1 | 1pF | 10 | 10.2 | 1.2 | 2.4696 | 0.088527 | 0.0346 | 2.1477 |  | 1pF | 10 | 10.6 | 1.2 | 2.4696 | 0.094467 | 0.034599 | 2.1477 |  | 1pF | 10 | 10.3 | 1.2 | 2.4696 | 0.21829 | 0.2037 | 2.1404 |  | 1pF | 10 | 10.3256 | 1.2348 | 2.4695 | 0.092024 | 0.34495 | 2.082 |  | 1pF | 10 | 10.4646 | 1.2853 | 2.5704 | 2.3509 | 0.48049 | 2.0217
E2 | DIV | 20 | 20.4647 | 1.2348 | 2.4696 | 0.13193 | 0.034559 | 2.1491 |  | DIV | 20 | 20.9974 | 1.2348 | 2.4696 | 0.1762 | 0.034559 | 2.1492 |  | DIV | 20 | 21.0168 | 1.2348 | 2.4696 | 0.29118 | 0.11729 | 2.1438 |  | DIV | 20 | 21.0663 | 1.2348 | 2.4695 | 0.13261 | 0.17817 | 2.0702 |  | DIV | 20 | 21.2793 | 1.2853 | 2.5704 | 1.3911 | 0.21955 | 2.0131
E3 |  | 30 | 31.0341 | 1.2348 | 2.4696 | 0.12888 | 0.034533 | 2.1501 |  |  | 30 | 31.4918 | 1.2348 | 2.4696 | 0.16511 | 0.034533 | 2.1501 |  |  | 30 | 31.6982 | 1.2348 | 2.4696 | 0.27445 | 0.089296 | 2.146 |  |  | 30 | 31.7019 | 1.2348 | 2.4694 | 0.052217 | 0.12525 | 2.0637 |  |  | 30 | 31.4978 | 1.2853 | 2.5704 | 0.75247 | 0.14576 | 2.0137
E4 |  | 40 | 41.5489 | 1.2348 | 2.4696 | 0.15403 | 0.034515 | 2.1508 |  |  | 40 | 41.4958 | 1.2348 | 2.4696 | 0.18314 | 0.034516 | 2.1508 |  |  | 40 | 41.8126 | 1.2348 | 2.4696 | 0.29798 | 0.075987 | 2.1474 |  |  | 40 | 41.8955 | 1.2348 | 2.4694 | 0.049209 | 0.10028 | 2.0612 |  |  | 40 | 41.7954 | 1.2853 | 2.5703 | 0.54794 | 0.11131 | 2.0119
E5 |  | 50 | 51.4655 | 1.2348 | 2.4695 | 0.023317 | 0.040484 | 2.0755 |  |  | 50 | 51.1505 | 1.2348 | 2.4695 | 0.01955 | 0.040558 | 2.0753 |  |  | 50 | 51.557 | 1.2348 | 2.4695 | 0.031034 | 0.095176 | 2.0672 |  |  | 50 | 51.636 | 1.2348 | 2.4693 | 0.038303 | 0.085959 | 2.0614 |  |  | 50 | 51.6939 | 1.2853 | 2.5703 | 0.46236 | 0.091936 | 2.01
E6 |  | 60 | 61.3009 | 1.2348 | 2.4695 | 0.014682 | 0.039422 | 2.079 |  |  | 60 | 60.8898 | 1.2348 | 2.4695 | 0.0082117 | 0.039491 | 2.0786 |  |  | 60 | 61.6953 | 1.2348 | 2.4695 | 0.022693 | 0.083241 | 2.0646 |  |  | 60 | 61.1184 | 1.2348 | 2.4693 | 0.032946 | 0.076623 | 2.0633 |  |  | 60 | 61.5424 | 1.2853 | 2.5702 | 0.39487 | 0.079262 | 2.0091
E7 |  | 70 | 70.4282 | 1.2348 | 2.4695 | 0.012779 | 0.038662 | 2.0829 |  |  | 70 | 70.2149 | 1.2348 | 2.4695 | 0 | 0.038694 | 2.0827 |  |  | 70 | 70.4245 | 1.2348 | 2.4695 | 0.0142 | 0.075891 | 2.0639 |  |  | 70 | 70.6509 | 1.2348 | 2.4692 | 0.048452 | 0.069923 | 2.0663 |  |  | 70 | 70.9769 | 1.2853 | 2.5702 | 0.31561 | 0.070715 | 2.0094
E8 |  | 80 | 80.4827 | 1.2348 | 2.4694 | 0.0074552 | 0.038012 | 2.0873 |  |  | 80 | 80.3052 | 1.2348 | 2.4694 | 0 | 0.038026 | 2.0873 |  |  | 80 | 80.2848 | 1.2348 | 2.4694 | 0.0062279 | 0.069657 | 2.0645 |  |  | 80 | 80.2034 | 1.2348 | 2.4691 | 0.04799 | 0.064926 | 2.0699 |  |  | 80 | 80.6603 | 1.2853 | 2.5701 | 0.27276 | 0.06431 | 2.0111
E9 |  | 90 | 90.4851 | 1.2348 | 2.4694 | 0.0022104 | 0.037628 | 2.0893 |  |  | 90 | 90.0535 | 1.2348 | 2.4694 | 0 | 0.037517 | 2.0917 |  |  | 90 | 90.4232 | 1.2348 | 2.4694 | 0 | 0.064791 | 2.0662 |  |  | 90 | 90.3229 | 1.2348 | 2.4691 | 0.041556 | 0.060881 | 2.0742 |  |  | 90 | 90.6839 | 1.2852 | 2.57 | 0.2371 | 0.059353 | 2.0144
E10 |  | '100' | 100.2635 | 1.2348 | 2.4693 | 0 | 0.03708 | 2.0963 |  |  | '100' | 99.8057 | 1.2348 | 2.4693 | -0.01002 | 0.037114 | 2.0958 |  |  | '100' | 100.2685 | 1.2348 | 2.4693 | -0.010971 | 0.061102 | 2.0686 |  |  | '100' | 100.1625 | 1.2348 | 2.469 | 0.025056 | 0.057803 | 2.0785 |  |  | '100' | 100.0583 | 1.2852 | 2.57 | 0.20889 | 0.055792 | 2.0186
D1 | 100pF | 1 | 97.49 | 0.016 | 2.1075 | 0.074881 | 0.037187 | 2.0951 |  | 100pF | 1 | 97.145 | 0.016 | 2.1039 | 0.05147 | 0.037211 | 2.0948 |  | 100pF | 1 | 96.977 | 0.015 | 2.1079 | 0.03506 | 0.062242 | 2.0678 |  | 100pF | 1 | 96.96 | 0.02408 | 2.2338 | 0.32282 | 0.058729 | 2.0771 |  | 100pF | 1 | 96.6318 | 0.018274 | 2.062 | 0.05692 | 0.056946 | 2.0171
D2 | STEPS | 2 | 195.9764 | 0.030256 | 2.1366 | 0.023983 | 0.035383 | 2.1248 |  | STEPS | 2 | 195.8469 | 0.030359 | 2.1321 | 0.020424 | 0.035387 | 2.1247 |  | STEPS | 2 | 195.8394 | 0.030609 | 2.1249 | -0.0010213 | 0.046049 | 2.0993 |  | STEPS | 2 | 195.9392 | 0.040183 | 2.2424 | 0.047648 | 0.045048 | 2.1123 |  | STEPS | 2 | 195.735 | 0.030333 | 2.0683 | 0.02197 | 0.042311 | 2.0883
D3 |  | 3 | 293.3355 | 0.044873 | 2.147 | 0.032046 | 0.034919 | 2.1372 |  |  | 3 | 292.8383 | 0.044978 | 2.1421 | 0.020489 | 0.034923 | 2.1371 |  |  | 3 | 292.6765 | 0.045488 | 2.1304 | 0.0061502 | 0.04163 | 2.1199 |  |  | 3 | 292.7762 | 0.055966 | 2.2449 | 0.13935 | 0.04117 | 2.1293 |  |  | 3 | 292.2347 | 0.042243 | 2.068 | 0.025666 | 0.039614 | 2.1426
D4 |  | 4 | 392.6409 | 0.059837 | 2.1517 | 0.016046 | 0.034735 | 2.143 |  |  | 4 | 392.2299 | 0.060013 | 2.1467 | 0.012748 | 0.034738 | 2.1429 |  |  | 4 | 392.2489 | 0.06082 | 2.133 | -0.0086681 | 0.039572 | 2.1315 |  |  | 4 | 392.4936 | 0.072229 | 2.246 | 0.04977 | 0.039311 | 2.1381 |  |  | 4 | 392.049 | 0.05494 | 2.0678 | 0.0071424 | 0.038703 | 2.1738
D5 |  | 5 | 489.943 | 0.074526 | 2.1542 | 0.02286 | 0.03464 | 2.1462 |  |  | 5 | 489.1993 | 0.07471 | 2.149 | 0.010221 | 0.034651 | 2.1459 |  |  | 5 | 489.044 | 0.075742 | 2.1342 | -0.00593 | 0.038445 | 2.1379 |  |  | 5 | 489.3095 | 0.088022 | 2.2465 | 0.10453 | 0.038276 | 2.1427 |  |  | 5 | 488.5508 | 0.067267 | 2.0675 | 0.0092114 | 0.03832 | 2.1909
D6 |  | 6 | 592.9669 | 0.090096 | 2.1557 | 0.017877 | 0.034587 | 2.1481 |  |  | 6 | 592.3264 | 0.090356 | 2.1505 | 0.0033766 | 0.034701 | 2.1441 |  |  | 6 | 592.0099 | 0.091625 | 2.135 | -0.010642 | 0.037683 | 2.1422 |  |  | 6 | 592.1304 | 0.1048 | 2.2468 | 0.094743 | 0.037568 | 2.1457 |  |  | 6 | 591.5245 | 0.081072 | 2.0683 | -0.0062554 | 0.038108 | 2.2017
D7 |  | 7 | 691.4952 | 0.105 | 2.1566 | 0.023283 | 0.034553 | 2.1494 |  |  | 7 | 691.042 | 0.10534 | 2.1513 | 0.013458 | 0.034558 | 2.1492 |  |  | 7 | 690.8962 | 0.10689 | 2.1355 | -0.0041975 | 0.037181 | 2.1448 |  |  | 7 | 691.1667 | 0.12096 | 2.247 | 0.050106 | 0.037095 | 2.1475 |  |  | 7 | 690.6187 | 0.093154 | 2.0669 | 0.0099916 | 0.03799 | 2.2083
D8 |  | 8 | 788.8593 | 0.11973 | 2.1572 | 0.033594 | 0.03453 | 2.1502 |  |  | 8 | 788.0353 | 0.12008 | 2.1519 | 0.025761 | 0.034531 | 2.1502 |  |  | 8 | 787.7233 | 0.12183 | 2.1358 | 0.0074901 | 0.036819 | 2.1466 |  |  | 8 | 787.9986 | 0.13676 | 2.2471 | 0.083688 | 0.036753 | 2.1487 |  |  | 8 | 787.1525 | 0.10554 | 2.0666 | 0.027061 | 0.037916 | 2.2126
D9 |  | 9 | 888.1807 | 0.13476 | 2.1576 | 0.043573 | 0.034514 | 2.1508 |  |  | 9 | 887.486 | 0.13519 | 2.1523 | 0.038311 | 0.034514 | 2.1508 |  |  | 9 | 887.3457 | 0.13721 | 2.136 | 0.023216 | 0.036535 | 2.1479 |  |  | 9 | 887.734 | 0.15303 | 2.2472 | 0.050688 | 0.036482 | 2.1496 |  |  | 9 | 886.9588 | 0.11835 | 2.0664 | 0.037321 | 0.037864 | 2.2158
D10 |  | X | 985.3288 | 0.14947 | 2.1579 | 0.050137 | 0.034503 | 2.1513 |  |  | X | 984.2543 | 0.14989 | 2.1526 | 0.042673 | 0.034503 | 2.1512 |  |  | X | 983.9547 | 0.15213 | 2.1362 | 0.025916 | 0.036317 | 2.1488 |  |  | X | 984.3529 | 0.1688 | 2.2473 | 0.077226 | 0.036273 | 2.1502 |  |  | 10 | 983.2735 | 0.13072 | 2.0662 | 0.044649 | 0.037829 | 2.218
C1 | 1000pF | 1 | 0.99505 | 0.00017 | 2.0804 | 0.020099 | 0.040169 | 2.0762 |  | 1000pF | 1 | 0.99471 | 0.00017 | 2.078 | 0.023121 | 0.040172 | 2.0763 |  | 1000pF | 1 | 0.9952 | 0.00016 | 2.0906 | 0.042201 | 0.049981 | 2.0662 |  | 1000pF | 1 | 0.99596 | 0.00017247 | 2.222 | 0.024097 | 0.045354 | 2.0619 |  | 1000pF | 1 | 0.9951 | 0.00013731 | 2.0442 | 0.040197 | 0.055932 | 2.0185
C2 | STEPS | 2 | 1.9976 | 0.00031105 | 2.1281 | 0.010011 | 0.03587 | 2.114 |  | STEPS | 2 | 1.9967 | 0.00031216 | 2.1239 | 0.010517 | 0.035847 | 2.1146 |  | STEPS | 2 | 1.9971 | 0.00031376 | 2.1202 | 0.019527 | 0.038708 | 2.0866 |  | STEPS | 2 | 1.9991 | 0.00033526 | 2.2403 | 0.010004 | 0.03751 | 2.1004 |  | STEPS | 2 | 1.9975 | 0.0002605 | 2.0571 | 0.017021 | 0.042112 | 2.0912
C3 |  | 3 | 2.9924 | 0.00045898 | 2.1438 | 0.010025 | 0.035105 | 2.1317 |  |  | 3 | 2.9909 | 0.00046059 | 2.1391 | 0.0080239 | 0.035095 | 2.132 |  |  | 3 | 2.992 | 0.00046576 | 2.1288 | 0.015374 | 0.036484 | 2.1131 |  |  | 3 | 2.9947 | 0.00049746 | 2.2442 | 0.0083478 | 0.035976 | 2.124 |  |  | 3 | 2.9923 | 0.00038567 | 2.0601 | 0.012365 | 0.039517 | 2.1455
C4 |  | 4 | 3.9949 | 0.00060949 | 2.1503 | -0.0025031 | 0.035404 | 2.1236 |  |  | 4 | 3.9927 | 0.00061159 | 2.1453 | 0.0010018 | 0.037183 | 2.0974 |  |  | 4 | 3.9943 | 0.00061976 | 2.1322 | 0.0062586 | 0.035684 | 2.1278 |  |  | 4 | 3.9982 | 0.00066111 | 2.2456 | 0.00075031 | 0.035407 | 2.1354 |  |  | 4 | 3.9949 | 0.00051588 | 2.062 | 0.0057574 | 0.038664 | 2.1754
C5 |  | 5 | 4.9985 | 0.00076076 | 2.1534 | -0.004001 | 0.034876 | 2.1383 |  |  | 5 | 4.9959 | 0.00076339 | 2.1483 | -0.0032025 | 0.03482 | 2.14 |  |  | 5 | 4.9973 | 0.00077423 | 2.1339 | 0.0002001 | 0.035303 | 2.1359 |  |  | 5 | 5.0024 | 0.00082491 | 2.2463 | -0.0047976 | 0.035129 | 2.1413 |  |  | 5 | 4.9983 | 0.00064374 | 2.0625 | -0.0014005 | 0.038292 | 2.1922
C6 |  | 6 | 5.9927 | 0.00091086 | 2.1552 | -0.0066745 | 0.034659 | 2.1455 |  |  | 6 | 5.9896 | 0.00091401 | 2.15 | -0.0066779 | 0.034623 | 2.1468 |  |  | 6 | 5.9916 | 0.00092751 | 2.1348 | -0.0011682 | 0.03509 | 2.1407 |  |  | 6 | 5.9974 | 0.00098726 | 2.2467 | -0.0061691 | 0.034971 | 2.1446 |  |  | 6 | 5.9926 | 0.00077055 | 2.0627 | -0.0093448 | 0.038099 | 2.2022
C7 |  | 7 | 6.9955 | 0.0010624 | 2.1563 | 0 | 0.03456 | 2.1491 |  |  | 7 | 6.9916 | 0.0010661 | 2.1511 | 0.00057208 | 0.042668 | 2.095 |  |  | 7 | 6.9937 | 0.0010821 | 2.1353 | 0.0035745 | 0.034956 | 2.1438 |  |  | 7 | 7.0007 | 0.001151 | 2.2469 | -0.00028568 | 0.034869 | 2.1468 |  |  | 7 | 6.995 | 0.0008923 | 2.062 | 0.0020014 | 0.037985 | 2.2085
C8 |  | 8 | 7.9896 | 0.0012128 | 2.157 | 0.0087611 | 0.034633 | 2.1465 |  |  | 8 | 7.9852 | 0.0012169 | 2.1517 | 0.010519 | 0.034582 | 2.1483 |  |  | 8 | 7.9878 | 0.0012355 | 2.1357 | 0.013395 | 0.034867 | 2.1458 |  |  | 8 | 7.9956 | 0.0013133 | 2.2471 | 0.009755 | 0.0348 | 2.1481 |  |  | 8 | 7.9891 | 0.0010183 | 2.0621 | 0.014395 | 0.037914 | 2.2128
C9 |  | 9 | 8.9917 | 0.0013644 | 2.1575 | 0.02113 | 0.034552 | 2.1494 |  |  | 9 | 8.9867 | 0.0013691 | 2.1522 | 0.022254 | 0.034541 | 2.1498 |  |  | 9 | 8.9899 | 0.0013902 | 2.1359 | 0.024248 | 0.034803 | 2.1472 |  |  | 9 | 8.9988 | 0.001477 | 2.2472 | 0.022002 | 0.034751 | 2.149 |  |  | 9 | 8.9914 | 0.0011305 | 2.0605 | 0.022911 | 0.037865 | 2.2157
C10 |  | X | 9.9859 | 0.0015149 | 2.1579 | 0.034047 | 0.034531 | 2.1502 |  |  | X | 9.9803 | 0.00152 | 2.1526 | 0.036069 | 0.034525 | 2.1504 |  |  | X | 9.984 | 0.0015437 | 2.1361 | 0.039461 | 0.034757 | 2.1481 |  |  | X | 9.9938 | 0.0016394 | 2.2472 | 0.036322 | 0.034714 | 2.1497 |  |  | 10 | 9.9855 | 0.0012021 | 2.0558 | 0.037053 | 0.037832 | 2.2178
B1 | .01mF | 1 | 10.01 | 0.009 | 2.1352 | 0.029969 | 0.2564 | 2.2485 |  | 0.01mF | 1 | 10.0058 | 0.001708 | 2.0752 | 0.036977 | 0.040658 | 2.0749 |  | 0.01mF | 1 | 10.0116 | 0.0017 | 2.0885 | 0.040951 | 0.04884 | 2.0674 |  | 0.01µF | 1 | 10.0212 | 0.0016624 | 2.2202 | 0.040912 | 0.044109 | 2.0612 |  | .01mF | 1 | 10.0127 | 0.0013724 | 2.0447 | 0.036953 | 0.055733 | 2.0188
B2 | STEPS | 2 | 20.0017 | 0.0094115 | 2.1107 | 0.039995 | 0.13041 | 2.2152 |  | STEPS | 2 | 19.9948 | 0.0031296 | 2.1229 | 0.026506 | 0.035945 | 2.1125 |  | STEPS | 2 | 20.0005 | 0.0031445 | 2.1196 | 0.029498 | 0.038051 | 2.0844 |  | STEPS | 2 | 20.0207 | 0.0032845 | 2.24 | 0.029469 | 0.036828 | 2.098 |  | STEPS | 2 | 20.0044 | 0.0026272 | 2.0587 | 0.028494 | 0.042136 | 2.0907
B3 |  | 3 | 30.0116 | 0.0098607 | 2.0845 | 0.026655 | 0.089259 | 2.1739 |  |  | 3 | 30.0004 | 0.0046213 | 2.1387 | 0.017999 | 0.035202 | 2.1288 |  |  | 3 | 30.0119 | 0.0046727 | 2.1286 | 0.020991 | 0.036079 | 2.11 |  |  | 3 | 30.0416 | 0.004917 | 2.2441 | 0.020637 | 0.035565 | 2.1208 |  |  | 3 | 30.0168 | 0.0039728 | 2.0641 | 0.019989 | 0.039587 | 2.1432
B4 |  | 4 | 40.0154 | 0.010499 | 2.0652 | 0.042482 | 0.069653 | 2.1356 |  |  | 4 | 40.001 | 0.0061281 | 2.1451 | 0.011749 | 0.035044 | 2.1331 |  |  | 4 | 40.0129 | 0.006209 | 2.1321 | 0.015494 | 0.035475 | 2.1225 |  |  | 4 | 40.053 | 0.0065495 | 2.2456 | 0.015229 | 0.035196 | 2.1298 |  |  | 4 | 40.0203 | 0.0052439 | 2.0643 | 0.013493 | 0.0388 | 2.1698
B5 |  | 5 | 50.0143 | 0.011311 | 2.0553 | 0.033989 | 0.0587 | 2.1059 |  |  | 5 | 49.9958 | 0.00764 | 2.1482 | 0.0086003 | 0.035078 | 2.1323 |  |  | 5 | 50.0236 | 0.0077504 | 2.1338 | 0.011594 | 0.035265 | 2.1276 |  |  | 5 | 50.0693 | 0.0081834 | 2.2463 | 0.010785 | 0.03509 | 2.1326 |  |  | 5 | 50.0294 | 0.0065972 | 2.0659 | 0.0093944 | 0.038503 | 2.1823
B6 |  | 6 | 60.0241 | 0.01227 | 2.0531 | 0.028321 | 0.052033 | 2.0855 |  |  | 6 | 60.0013 | 0.0091565 | 2.15 | 0 | 0.034601 | 2.1476 |  |  | 6 | 60.0348 | 0.0092937 | 2.1347 | 0.0054965 | 0.03522 | 2.1288 |  |  | 6 | 60.0901 | 0.0098184 | 2.2467 | 0.0048259 | 0.035101 | 2.1324 |  |  | 6 | 60.0421 | 0.0078526 | 2.0651 | -0.011825 | 0.038401 | 2.1873
B7 |  | 7 | 70.016 | 0.013344 | 2.0561 | 0.032848 | 0.047773 | 2.0731 |  |  | 7 | 69.9899 | 0.010672 | 2.151 | 0.011573 | 0.035239 | 2.1286 |  |  | 7 | 70.0234 | 0.010835 | 2.1353 | 0.015851 | 0.035266 | 2.1278 |  |  | 7 | 70.0893 | 0.01145 | 2.2469 | 0.01498 | 0.035179 | 2.1304 |  |  | 7 | 70.0331 | 0.0091554 | 2.0653 | 0.012708 | 0.038397 | 2.188
B8 |  | 8 | 80.0248 | 0.014511 | 2.0621 | 0.024991 | 0.044951 | 2.0664 |  |  | 8 | 79.9952 | 0.012191 | 2.1517 | 0.02125 | 0.035327 | 2.1267 |  |  | 8 | 80.0345 | 0.012379 | 2.1357 | 0.025363 | 0.035373 | 2.1254 |  |  | 8 | 80.1099 | 0.013085 | 2.2471 | 0.024341 | 0.035307 | 2.1273 |  |  | 8 | 80.0451 | 0.010541 | 2.0665 | 0.022362 | 0.038456 | 2.1861
B9 |  | 9 | 90.0277 | 0.015746 | 2.0694 | 0.027768 | 0.043048 | 2.0634 |  |  | 9 | 89.9957 | 0.01371 | 2.1522 | 0.033222 | 0.035482 | 2.1235 |  |  | 9 | 90.0352 | 0.013923 | 2.1359 | 0.03865 | 0.035526 | 2.1223 |  |  | 9 | 90.1209 | 0.014719 | 2.2472 | 0.037948 | 0.035475 | 2.1237 |  |  | 9 | 90.0482 | 0.011766 | 2.0655 | 0.035425 | 0.038561 | 2.1826
B10 |  | X | 100.0375 | 0.017035 | 2.0771 | 0.039983 | 0.041754 | 2.0626 |  |  | X | 100.0012 | 0.015231 | 2.1526 | 0.048897 | 0.035679 | 2.1198 |  |  | X | 100.0457 | 0.015469 | 2.1361 | 0.054672 | 0.035719 | 2.1186 |  |  | X | 100.1418 | 0.016354 | 2.2472 | 0.053323 | 0.035679 | 2.1198 |  |  | 10 | 100.0601 | 0.013172 | 2.0666 | 0.050369 | 0.038703 | 2.178
A1 | 0.1mF | 1 | 0.10003 | 1.7e-05 | 2.0771 | 0.042985 | 0.041755 | 2.0626 |  | 0.1mF | 1 | 0.099994 | 1.7e-05 | 2.0748 | 0.039 | 0.041759 | 2.0626 |  | 0.1mF | 1 | 0.099981 | 1.7e-05 | 2.0882 | 0.039006 | 0.04965 | 2.0592 |  | 0.1µF | 1 | 0.10003 | 1.6523e-05 | 2.2199 | 0.042985 | 0.045009 | 2.052 |  | 0.1mF | 1 | 0.099973 | 1.3323e-05 | 2.0406 | 0.04001 | 0.056589 | 2.0154
A2 | STEPS | 2 | 0.20038 | 3.1239e-05 | 2.1269 | 0.044914 | 0.040858 | 2.073 |  | STEPS | 2 | 0.20036 | 3.1361e-05 | 2.1229 | 0.041922 | 0.040858 | 2.073 |  | STEPS | 2 | 0.20036 | 3.1502e-05 | 2.1196 | 0.046414 | 0.042709 | 2.0568 |  | STEPS | 2 | 0.20049 | 3.2817e-05 | 2.24 | 0.051374 | 0.04162 | 2.0647 |  | STEPS | 2 | 0.20038 | 2.5916e-05 | 2.0567 | 0.044914 | 0.046296 | 2.0596
A3 |  | 3 | 0.3004 | 4.6091e-05 | 2.1434 | 0.036616 | 0.046827 | 2.1029 |  |  | 3 | 0.30035 | 4.6268e-05 | 2.1387 | 0.036622 | 0.046823 | 2.1028 |  |  | 3 | 0.30034 | 4.6763e-05 | 2.1286 | 0.046611 | 0.047486 | 2.0928 |  |  | 3 | 0.30052 | 4.9114e-05 | 2.2441 | 0.05091 | 0.04711 | 2.0985 |  |  | 3 | 0.30035 | 3.8582e-05 | 2.0601 | 0.03962 | 0.049721 | 2.0927
A4 |  | 4 | 0.40073 | 6.1147e-05 | 2.1501 | 0.039926 | 0.055546 | 2.1532 |  |  | 4 | 0.4007 | 6.1386e-05 | 2.1451 | 0.046666 | 0.055542 | 2.1532 |  |  | 4 | 0.40072 | 6.2181e-05 | 2.1321 | 0.058642 | 0.055824 | 2.1469 |  |  | 4 | 0.40097 | 6.5493e-05 | 2.2456 | 0.060602 | 0.055685 | 2.1506 |  |  | 4 | 0.40075 | 5.1356e-05 | 2.0614 | 0.052151 | 0.057201 | 2.1357
A5 |  | 5 | 0.50055 | 7.6188e-05 | 2.1533 | 0.059932 | 0.065718 | 2.2059 |  |  | 5 | 0.50049 | 7.6482e-05 | 2.1482 | 0.067331 | 0.065711 | 2.2059 |  |  | 5 | 0.50055 | 7.7553e-05 | 2.1338 | 0.075913 | 0.065848 | 2.2018 |  |  | 5 | 0.50087 | 8.179e-05 | 2.2463 | 0.079659 | 0.065807 | 2.2044 |  |  | 5 | 0.50057 | 6.4076e-05 | 2.062 | 0.06912 | 0.066636 | 2.1832
A6 |  | 6 | 0.60058 | 9.1289e-05 | 2.1551 | 0.063269 | 0.07676 | 2.2524 |  |  | 6 | 0.60049 | 9.1637e-05 | 2.15 | 0.059948 | 0.076749 | 2.2524 |  |  | 6 | 0.60053 | 9.2966e-05 | 2.1347 | 0.073598 | 0.076818 | 2.2497 |  |  | 6 | 0.60091 | 9.8112e-05 | 2.2467 | 0.084869 | 0.076824 | 2.2514 |  |  | 6 | 0.60057 | 7.6829e-05 | 2.0623 | 0.013987 | 0.077184 | 2.2271
A7 |  | 7 | 0.70093 | 0.00010645 | 2.1563 | 0.068478 | 0.088329 | 2.2906 |  |  | 7 | 0.70085 | 0.00010687 | 2.151 | 0.072765 | 0.08832 | 2.2906 |  |  | 7 | 0.70092 | 0.00010845 | 2.1353 | 0.089307 | 0.088359 | 2.2888 |  |  | 7 | 0.70136 | 0.0001145 | 2.2469 | 0.10123 | 0.088393 | 2.29 |  |  | 7 | 0.70095 | 8.9638e-05 | 2.0625 | 0.074469 | 0.088402 | 2.2643
A8 |  | 8 | 0.80096 | 0.00012158 | 2.157 | 0.079901 | 0.10015 | 2.321 |  |  | 8 | 0.80085 | 0.00012205 | 2.1517 | 0.086529 | 0.10014 | 2.3209 |  |  | 8 | 0.8009 | 0.00012388 | 2.1357 | 0.10712 | 0.10016 | 2.3198 |  |  | 8 | 0.80139 | 0.00013083 | 2.2471 | 0.12079 | 0.10021 | 2.3206 |  |  | 8 | 0.80083 | 0.00010239 | 2.0626 | 0.091904 | 0.099931 | 2.2944
A9 |  | 9 | 0.90129 | 0.00013676 | 2.1575 | 0.092087 | 0.11219 | 2.3451 |  |  | 9 | 0.90119 | 0.00013729 | 2.1522 | 0.10796 | 0.11218 | 2.3451 |  |  | 9 | 0.90128 | 0.00013938 | 2.1359 | 0.13059 | 0.11219 | 2.3443 |  |  | 9 | 0.90184 | 0.00014722 | 2.2472 | 0.1437 | 0.11226 | 2.3449 |  |  | 9 | 0.90133 | 0.00011522 | 2.0627 | 0.11239 | 0.11176 | 2.3187
A10 |  | X | 1.0013 | 0.0001519 | 2.1578 | 0.11384 | 0.1243 | 2.3643 |  |  | X | 1.0012 | 0.00015248 | 2.1526 | 0.12884 | 0.12429 | 2.3642 |  |  | X | 1.0013 | 0.00015481 | 2.1361 | 0.149 | 0.12429 | 2.3636 |  |  | X | 1.0019 | 0.00016354 | 2.2472 | 0.16519 | 0.12437 | 2.3641 |  |  | 10 | 1.0013 | 0.00012798 | 2.0628 | 0.13253 | 0.12367 | 2.3381
| Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k |  | Dial | Setting | Capacitance | Expanded | k | Loss angle | Expanded | k
|  |  |  | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty | [5 empty] | uncertainty |  |  | uncertainty
[6 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian) | [7 empty] | (milliradian) | (milliradian)
E0 | All dials | 0 | 49.4 | 1.5 | 2.4696 | 0 | 0.034666 | 2.1453 |  | All dials | 0 | 49.1 | 1.5 | 2.4696 | 0 | 0.034669 | 2.1453 |  | All dials | 0 | 49.1 | 1.5 | 2.4696 | 3.2825 | 0.035596 | 2.1352 |  | All dials | 0 | 48.9946 | 1.4818 | 2.4696 | 2.9562 | 0.069554 | 2.0988 |  | All dials | 0 | 49.1435 | 1.5423 | 2.5705 | 2.1978 | 0.11021 | 2.0182
e4
| C | uC | D | uD
1999-09-23 | 41.5489 | 1.2348 | 0.15403 | 0.034515
2004-09-10 | 41.4958 | 1.2348 | 0.18314 | 0.034516
2009-12-01 | 41.8126 | 1.2348 | 0.29798 | 0.075987
2015-04-30 | 41.8955 | 1.2348 | 0.049209 | 0.10028
2020-10-23 | 41.7954 | 1.2853 | 0.54794 | 0.11131