## Commentary
2020-07-29
Worksheet 'pyUBreadings' requires the user to enter all the dial readings and settings of the UB as well as estimates of the uncertainty in temperature and in the temperature coefficients for resistance and inductance. In general it is set up not to apply any temperature corrections, just to evaluate the uncertainty.
Use the UniversalBridge3 project to process the input data. This uses the latest ubdict.csv file for corrections to the Universal Bridge. Running 'calculate_imp.py' produces a new 'indResults.xlsx'. Paste these results into the 'pyUBresults' worksheet of this spreadsheet and, if correctly pasted, should immediately produce the report in the'pyReport' worksheet. Note that there are two named blocks in 'pyUBresults' and so the new data should just be pasted over the old data in order to preserve the inductor1 and inductor2 named blocks.
You should check that the input data listed in 'pyUBresults' matches the input data in 'pyUBreadings', just in case the wrong file names are enterd in 'calculate_imp.py'.

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | Muirhead A-517 |  | 4: Client; 10: Enter, say, 5 ppm as 5e-6
Date from | 2019-12-04 |  | 4: Job number; 5: 96330136; 10: 1 sigma uncertainties from specifications
to | 2019-12-05 |  | 4: S Number; 5: S22075
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
L0 | 1591.5 | 4Z | 0 | 62 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 4Z | 101490 | 194135 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 4Z | 199872 | 279075 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 4Z | 292630 | 343493 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 4Z | 393434 | 404033 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 4Z | 490448 | 457475 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 4Z | 590862 | 508465 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 4Z | 689292 | 555325 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 4Z | 788633 | 600890 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 4Z | 889659 | 646103 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 4Z | 1010749 | 696555 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L0 | 1591.5 | 5Z | 0 | -21 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 5Z | 103895 | 167536 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 5Z | 207691 | 242983 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 5Z | 304323 | 300344 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 5Z | 398969 | 348977 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 5Z | 494778 | 394506 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 5Z | 592475 | 437881 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 5Z | 689987 | 478933 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 5Z | 783333 | 517323 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 5Z | 878587 | 555569 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 5Z | 985448 | 596406 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
L0 | 1591.5 | 4Z | 0 | 62 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-07 | 0.0012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 4Z | 101490 | 194135 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.010151 | 9.0105e-07 | 2.0143 | 19.4087 | 0.11512 | 2.084 | 1.9072e-07 | 0.0016385 | 4.7243 | 70.2599 | L11591.5 | 1000 | 19408.6954 | 115.1181 | 2.084 | 1000000 | 10151.0928 | 0.90105 | 2.0143 | 1591.5
L2 | 1591.5 | 4Z | 199872 | 279075 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.019991 | 1.625e-06 | 2.0365 | 27.9033 | 0.1655 | 2.084 | 3.0808e-07 | 0.0023936 | 5.2747 | 69.1405 | L21591.5 | 1000 | 27903.3061 | 165.4962 | 2.084 | 1000000 | 19991.3215 | 1.625 | 2.0365 | 1591.5
L3 | 1591.5 | 4Z | 292630 | 343493 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.029269 | 2.3374e-06 | 2.0432 | 34.3456 | 0.20374 | 2.0838 | 4.2999e-07 | 0.0032215 | 5.4361 | 63.2429 | L31591.5 | 1000 | 34345.5693 | 203.7391 | 2.0838 | 1000000 | 29269.0343 | 2.3374 | 2.0432 | 1591.5
L4 | 1591.5 | 4Z | 393434 | 404033 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.039352 | 3.1205e-06 | 2.0461 | 40.4 | 0.23971 | 2.0837 | 5.6652e-07 | 0.0041696 | 5.5081 | 57.4895 | L41591.5 | 1000 | 40400.0046 | 239.709 | 2.0837 | 1000000 | 39351.5129 | 3.1205 | 2.0461 | 1591.5
L5 | 1591.5 | 4Z | 490448 | 457475 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.049055 | 3.8776e-06 | 2.0475 | 45.7446 | 0.27148 | 2.0835 | 6.9968e-07 | 0.0051043 | 5.5419 | 53.1865 | L51591.5 | 1000 | 45744.5888 | 271.4821 | 2.0835 | 1000000 | 49054.9135 | 3.8776 | 2.0475 | 1591.5
L6 | 1591.5 | 4Z | 590862 | 508465 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.059098 | 4.663e-06 | 2.0483 | 50.844 | 0.30182 | 2.0833 | 8.3842e-07 | 0.0060838 | 5.5617 | 49.6096 | L61591.5 | 1000 | 50843.9555 | 301.8176 | 2.0833 | 1000000 | 59098.3841 | 4.663 | 2.0483 | 1591.5
L7 | 1591.5 | 4Z | 689292 | 555325 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.068943 | 5.434e-06 | 2.0488 | 55.5303 | 0.32971 | 2.0832 | 9.7494e-07 | 0.007051 | 5.5736 | 46.7611 | L71591.5 | 1000 | 55530.2924 | 329.7131 | 2.0832 | 1000000 | 68943.4138 | 5.434 | 2.0488 | 1591.5
L8 | 1591.5 | 4Z | 788633 | 600890 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.07888 | 6.2127e-06 | 2.0491 | 60.0871 | 0.35685 | 2.083 | 1.1131e-06 | 0.0080317 | 5.5816 | 44.4303 | L81591.5 | 1000 | 60087.12 | 356.85 | 2.083 | 1000000 | 78879.5623 | 6.2127 | 2.0491 | 1591.5
L9 | 1591.5 | 4Z | 889659 | 646103 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.088984 | 7.0051e-06 | 2.0494 | 64.6087 | 0.38379 | 2.0829 | 1.2538e-06 | 0.0090322 | 5.5872 | 42.491 | L91591.5 | 1000 | 64608.7451 | 383.7869 | 2.0829 | 1000000 | 88984.2455 | 7.0051 | 2.0494 | 1591.5
L10 | 1591.5 | 4Z | 1010749 | 696555 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.1011 | 7.9552e-06 | 2.0495 | 69.6543 | 0.41387 | 2.0827 | 1.4226e-06 | 0.010234 | 5.5919 | 40.4405 | L101591.5 | 1000 | 69654.3078 | 413.8686 | 2.0827 | 1000000 | 101095.7425 | 7.9552 | 2.0495 | 1591.5
L0 | 1591.5 | 5Z | 0 | -21 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-06 | 0.012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 5Z | 103895 | 167536 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10392 | 9.1712e-06 | 2.0154 | 167.569 | 0.99402 | 2.0837 | 1.93e-06 | 0.016478 | 4.752 | 60.3255 | L11591.5 | 1 | 167.569 | 0.99402 | 2.0837 | 1000000 | 103916.424 | 9.1712 | 2.0154 | 1591.5
L2 | 1591.5 | 5Z | 207691 | 242983 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.20773 | 1.6838e-05 | 2.0375 | 243.0215 | 1.4417 | 2.0837 | 3.1776e-06 | 0.024518 | 5.2991 | 58.8017 | L21591.5 | 1 | 243.0215 | 1.4417 | 2.0837 | 1000000 | 207733.8277 | 16.8385 | 2.0375 | 1591.5
L3 | 1591.5 | 5Z | 304323 | 300344 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.30439 | 2.4273e-05 | 2.0438 | 300.3866 | 1.7825 | 2.0835 | 4.4531e-06 | 0.0332 | 5.4508 | 53.6901 | L31591.5 | 1 | 300.3866 | 1.7825 | 2.0835 | 1000000 | 304385.754 | 24.2729 | 2.0438 | 1591.5
L4 | 1591.5 | 5Z | 398969 | 348977 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.39905 | 3.1629e-05 | 2.0463 | 349.0231 | 2.0717 | 2.0833 | 5.737e-06 | 0.042123 | 5.5131 | 49.1837 | L41591.5 | 1 | 349.0231 | 2.0717 | 2.0833 | 1000000 | 399051.2708 | 31.629 | 2.0463 | 1591.5
L5 | 1591.5 | 5Z | 494778 | 394506 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.49488 | 3.9107e-05 | 2.0476 | 394.5554 | 2.3428 | 2.0831 | 7.0525e-06 | 0.051358 | 5.5452 | 45.6166 | L51591.5 | 1 | 394.5554 | 2.3428 | 2.0831 | 1000000 | 494880.0275 | 39.1071 | 2.0476 | 1591.5
L6 | 1591.5 | 5Z | 592475 | 437881 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.5926 | 4.6749e-05 | 2.0484 | 437.9335 | 2.6011 | 2.0829 | 8.4025e-06 | 0.060888 | 5.5638 | 42.7203 | L61591.5 | 1 | 437.9335 | 2.6011 | 2.0829 | 1000000 | 592597.1734 | 46.7494 | 2.0484 | 1591.5
L7 | 1591.5 | 5Z | 689987 | 478933 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.69013 | 5.4387e-05 | 2.0489 | 478.9884 | 2.8459 | 2.0827 | 9.755e-06 | 0.070468 | 5.5753 | 40.3852 | L71591.5 | 1 | 478.9884 | 2.8459 | 2.0827 | 1000000 | 690129.2812 | 54.387 | 2.0489 | 1591.5
L8 | 1591.5 | 5Z | 783333 | 517323 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.78349 | 6.1704e-05 | 2.0492 | 517.3812 | 3.0748 | 2.0825 | 1.1053e-05 | 0.079681 | 5.5827 | 38.589 | L81591.5 | 1 | 517.3812 | 3.0748 | 2.0825 | 1000000 | 783494.53 | 61.7041 | 2.0492 | 1591.5
L9 | 1591.5 | 5Z | 878587 | 555569 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.87877 | 6.9175e-05 | 2.0494 | 555.6299 | 3.303 | 2.0823 | 1.2379e-05 | 0.089112 | 5.5879 | 37.0661 | L91591.5 | 1 | 555.6299 | 3.303 | 2.0823 | 1000000 | 878768.1722 | 69.1748 | 2.0494 | 1591.5
L10 | 1591.5 | 5Z | 985448 | 596406 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.98565 | 7.7559e-05 | 2.0496 | 596.4699 | 3.5468 | 2.0821 | 1.3869e-05 | 0.099714 | 5.5922 | 35.5703 | L101591.5 | 1 | 596.4699 | 3.5468 | 2.0821 | 1000000 | 985651.2078 | 77.5591 | 2.0496 | 1591.5

## pyReport
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 | 10 | 101.0957 | 0.0079552 | 2.0495 | 69.6543 | 0.41387 | 2.0827
L9 | 9 | 88.9842 | 0.0070051 | 2.0494 | 64.6087 | 0.38379 | 2.0829
L8 | 8 | 78.8796 | 0.0062127 | 2.0491 | 60.0871 | 0.35685 | 2.083
L7 | 7 | 68.9434 | 0.005434 | 2.0488 | 55.5303 | 0.32971 | 2.0832
L6 | 6 | 59.0984 | 0.004663 | 2.0483 | 50.844 | 0.30182 | 2.0833
L5 | 5 | 49.0549 | 0.0038776 | 2.0475 | 45.7446 | 0.27148 | 2.0835
L4 | 4 | 39.3515 | 0.0031205 | 2.0461 | 40.4 | 0.23971 | 2.0837
L3 | 3 | 29.269 | 0.0023374 | 2.0432 | 34.3456 | 0.20374 | 2.0838
L2 | 2 | 19.9913 | 0.001625 | 2.0365 | 27.9033 | 0.1655 | 2.084
L1 | 1 | 10.1511 | 0.00090105 | 2.0143 | 19.4087 | 0.11512 | 2.084
| Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 | 10 | 985.6512 | 0.077559 | 2.0496 | 596.4699 | 3.5468 | 2.0821
L9 | 9 | 878.7682 | 0.069175 | 2.0494 | 555.6299 | 3.303 | 2.0823
L8 | 8 | 783.4945 | 0.061704 | 2.0492 | 517.3812 | 3.0748 | 2.0825
L7 | 7 | 690.1293 | 0.054387 | 2.0489 | 478.9884 | 2.8459 | 2.0827
L6 | 6 | 592.5972 | 0.046749 | 2.0484 | 437.9335 | 2.6011 | 2.0829
L5 | 5 | 494.88 | 0.039107 | 2.0476 | 394.5554 | 2.3428 | 2.0831
L4 | 4 | 399.0513 | 0.031629 | 2.0463 | 349.0231 | 2.0717 | 2.0833
L3 | 3 | 304.3858 | 0.024273 | 2.0438 | 300.3866 | 1.7825 | 2.0835
L2 | 2 | 207.7338 | 0.016838 | 2.0375 | 243.0215 | 1.4417 | 2.0837
L1 | 1 | 103.9164 | 0.0091712 | 2.0154 | 167.569 | 0.99402 | 2.0837

## exampleU
Budget for L8 1591.5
ureal(60.08704790830853,0.1713140491986457,20.451673221652594)
uuttempL8 | 0.12049
temperature | 0.12043
Amp_phase_err | 0.013329
rtempcoL8 | 0.012019
Adialresureal(62.0,10.0,10.0) | 0.0011776
Adialresureal(600890.0,10.0,10.0) | 0.0011776
Adial_lin62 | 0.0010001
Adial_lin600890 | 0.0010001
Amperr | 0.00084121
R4Aacdc | 0.00015022
G1acdc | 0.00015022
G2acdc | 0.00015022
G1alpha | 6.0087e-05
G1stab | 6.0087e-05
R4Aalpha | 6.0087e-05
G2alpha | 6.0087e-05
R4Astab | 6.0087e-05
G2stab | 6.0087e-05
cal_temp | 4.2488e-05
G2err | 3.0043e-05
G1err | 3.0043e-05
R4Aerr | 3.0042e-05
uuttempL0 | 1.2432e-05
rtempcoL0 | 1.2401e-06
C1err | 0
C1alpha | 0
C1stab | 0
Bdial_lin0 | 0
Bdialresureal(0.0,1.0,10.0, label='Bdial_lin0') | 0
Bdial_lin788633 | 0
Bdialresureal(788633.0,1.0,10.0) | 0
ureal(0.0788795386314938,3.031875053076101e-06,27.783955120984952)
uuttempL8 | 1.9769e-06
temperature | 1.9211e-06
Amperr | 1.1043e-06
C1stab | 3.944e-07
R4Aacdc | 1.972e-07
G2acdc | 1.972e-07
xtempcoL8 | 1.972e-07
Bdialres0.0 | 1.1778e-07
Bdialres788633.0 | 1.1778e-07
Amp_phase_err | 1.0161e-07
Bdial_lin0 | 1.0002e-07
Bdial_lin788633 | 1.0002e-07
C1alpha | 7.888e-08
R4Aalpha | 7.888e-08
G2alpha | 7.888e-08
R4Astab | 7.888e-08
G2stab | 7.888e-08
C1err | 7.8868e-08
cal_temp | 5.5776e-08
G2err | 3.9439e-08
R4Aerr | 3.9438e-08
G1err | 0
G1acdc | 0
G1alpha | 0
G1stab | 0
Adial_lin62 | 0
Adialres62.0 | 0
uuttempL0 | 0
xtempcoL0 | 0
Adial_lin600890 | 0
Adialres600890.0 | 0

## Room Conditions
2017-12-04 16:30 | 20.9 | 53.6 |  | RH correction | -2.128 | from Y:\Quality\TECHPROC\E058-Lab Air Cond\loggercalibration_2017.xls
2017-12-04 16:35 | 20.9 | 53.8 | Temperature | RH
2017-12-04 16:40 | 20.9 | 54.8 | #REF! | 52.2805 | 50.1525
2017-12-04 16:45 | 20.9 | 54.4 | #REF! | 1.1136
2017-12-04 16:50 | 20.9 | 54.3
2017-12-04 16:55 | 20.8 | 54.2 | from GMH
2017-12-04 17:00 | 20.8 | 53.6
2017-12-04 17:05 | 20.9 | 53.4
2017-12-04 17:10 | 20.9 | 53.3
2017-12-04 17:15 | 20.8 | 53.7
2017-12-04 17:20 | 20.8 | 53
2017-12-04 17:25 | 20.8 | 52.6
2017-12-04 17:30 | 20.9 | 52.4
2017-12-04 17:35 | 20.9 | 52
2017-12-04 17:40 | 20.9 | 52
2017-12-04 17:45 | 20.9 | 52.1
2017-12-04 17:50 | 20.9 | 52
2017-12-04 17:55 | 20.9 | 51.8
2017-12-04 18:00 | 20.9 | 52.1
2017-12-04 18:05 | 20.9 | 52.1
2017-12-04 18:10 | 20.9 | 52
2017-12-04 18:15 | 20.9 | 52.3
2017-12-04 18:20 | 20.9 | 52.3
2017-12-04 18:25 | 20.9 | 52
2017-12-04 18:30 | 21 | 51.9
2017-12-04 18:35 | 21 | 52
2017-12-04 18:40 | 20.9 | 51.9
2017-12-04 18:45 | 20.9 | 51.9
2017-12-04 18:50 | 20.9 | 51.8
2017-12-04 18:55 | 20.9 | 51.8
2017-12-04 19:00 | 20.9 | 51.8
2017-12-04 19:05 | 20.9 | 51.8
2017-12-04 19:10 | 21 | 51.6
2017-12-04 19:15 | 21 | 51.7
2017-12-04 19:20 | 21 | 51.6
2017-12-04 19:25 | 21 | 51.7
2017-12-04 19:30 | 21 | 51.7
2017-12-04 19:35 | 21 | 51.7
2017-12-04 19:40 | 21 | 51.9
2017-12-04 19:45 | 21 | 51.9
2017-12-04 19:50 | 21 | 51.6
2017-12-04 19:55 | 21 | 51.7
2017-12-04 20:00 | 21 | 51.7
2017-12-04 20:05 | 21 | 51.7
2017-12-04 20:10 | 21 | 51.6
2017-12-04 20:15 | 21 | 51.9
2017-12-04 20:20 | 21 | 51.9
2017-12-04 20:25 | 21 | 51.9
2017-12-04 20:30 | 21 | 51.8
2017-12-04 20:35 | 21 | 51.7
2017-12-04 20:40 | 21.1 | 51.6
2017-12-04 20:45 | 21.1 | 51.8
2017-12-04 20:50 | 21 | 52
2017-12-04 20:55 | 21 | 52.2
2017-12-04 21:00 | 21 | 52.2
2017-12-04 21:05 | 21 | 52
2017-12-04 21:10 | 21 | 51.9
2017-12-04 21:15 | 21 | 52.1
2017-12-04 21:20 | 21.1 | 52
2017-12-04 21:25 | 21 | 52.2
2017-12-04 21:30 | 21 | 51.9
2017-12-04 21:35 | 21 | 52
2017-12-04 21:40 | 21 | 51.8
2017-12-04 21:45 | 21.1 | 51.5
2017-12-04 21:50 | 21 | 51.7
2017-12-04 21:55 | 21.1 | 51.6
2017-12-04 22:00 | 21.1 | 51.5
2017-12-04 22:05 | 21.1 | 51.3
2017-12-04 22:10 | 21.1 | 51.5
2017-12-04 22:15 | 21 | 51.7
2017-12-04 22:20 | 21.1 | 51.6
2017-12-04 22:25 | 21.1 | 52
2017-12-04 22:30 | 21.1 | 51.5
2017-12-04 22:35 | 21.1 | 51.4
2017-12-04 22:40 | 21.1 | 51.2
2017-12-04 22:45 | 21.1 | 51.2
2017-12-04 22:50 | 21.1 | 51.3
2017-12-04 22:55 | 21.1 | 51.4
2017-12-04 23:00 | 21.1 | 51.3
2017-12-04 23:05 | 21.1 | 51.3
2017-12-04 23:10 | 21.2 | 51.4
2017-12-04 23:15 | 21.2 | 51.3
2017-12-04 23:20 | 21.2 | 51.5
2017-12-04 23:25 | 21.1 | 51.7
2017-12-04 23:30 | 21.1 | 51.7
2017-12-04 23:35 | 21.1 | 51.3
2017-12-04 23:40 | 21.1 | 51.5
2017-12-04 23:45 | 21.1 | 51.4
2017-12-04 23:50 | 21.1 | 51.4
2017-12-04 23:55 | 21.1 | 51.3
2017-12-05 | 21.1 | 51.4
2017-12-05 00:05 | 21.1 | 51.3
2017-12-05 00:10 | 21.1 | 51
2017-12-05 00:15 | 21.1 | 51.3
2017-12-05 00:20 | 21.1 | 51.1
2017-12-05 00:25 | 21.2 | 51.2
2017-12-05 00:30 | 21.1 | 51.4
2017-12-05 00:35 | 21.2 | 51.3
2017-12-05 00:40 | 21.1 | 51.5
2017-12-05 00:45 | 21.2 | 51.3
2017-12-05 00:50 | 21.2 | 51.3
2017-12-05 00:55 | 21.1 | 51.3
2017-12-05 01:00 | 21.2 | 51.3
2017-12-05 01:05 | 21.1 | 51.4
2017-12-05 01:10 | 21.2 | 51.2
2017-12-05 01:15 | 21.2 | 51.1
2017-12-05 01:20 | 21.1 | 51.3
2017-12-05 01:25 | 21.2 | 51.9
2017-12-05 01:30 | 21.2 | 51.9
2017-12-05 01:35 | 21.1 | 52
2017-12-05 01:40 | 21.1 | 51.8
2017-12-05 01:45 | 21.2 | 51.8
2017-12-05 01:50 | 21.1 | 51.8
2017-12-05 01:55 | 21.1 | 51.9
2017-12-05 02:00 | 21.1 | 52.1
2017-12-05 02:05 | 21.1 | 51.9
2017-12-05 02:10 | 21.2 | 51.8
2017-12-05 02:15 | 21.2 | 51.7
2017-12-05 02:20 | 21.2 | 51.9
2017-12-05 02:25 | 21.2 | 51.9
2017-12-05 02:30 | 21.2 | 51.8
2017-12-05 02:35 | 21.1 | 51.6
2017-12-05 02:40 | 21.2 | 51.7
2017-12-05 02:45 | 21.2 | 51.6
2017-12-05 02:50 | 21.1 | 51.7
2017-12-05 02:55 | 21.1 | 51.6
2017-12-05 03:00 | 21.1 | 51.6
2017-12-05 03:05 | 21.1 | 51.6
2017-12-05 03:10 | 21.1 | 51.3
2017-12-05 03:15 | 21.2 | 51.1
2017-12-05 03:20 | 21.1 | 51.3
2017-12-05 03:25 | 21.1 | 51.6
2017-12-05 03:30 | 21.1 | 51.7
2017-12-05 03:35 | 21.1 | 51.6
2017-12-05 03:40 | 21.2 | 51.5
2017-12-05 03:45 | 21.1 | 51.3
2017-12-05 03:50 | 21.1 | 51.2
2017-12-05 03:55 | 21.1 | 51.2
2017-12-05 04:00 | 21.2 | 51
2017-12-05 04:05 | 21.2 | 51.1
2017-12-05 04:10 | 21.2 | 51.4
2017-12-05 04:15 | 21.2 | 51.6
2017-12-05 04:20 | 21.1 | 51.5
2017-12-05 04:25 | 21.1 | 51.8
2017-12-05 04:30 | 21.2 | 51.7
2017-12-05 04:35 | 21.1 | 52
2017-12-05 04:40 | 21.1 | 51.9
2017-12-05 04:45 | 21.2 | 52
2017-12-05 04:50 | 21.1 | 52
2017-12-05 04:55 | 21.2 | 51.8
2017-12-05 05:00 | 21.2 | 51.8
2017-12-05 05:05 | 21.1 | 51.9
2017-12-05 05:10 | 21.1 | 51.8
2017-12-05 05:15 | 21.1 | 51.6
2017-12-05 05:20 | 21.2 | 51.7
2017-12-05 05:25 | 21.2 | 51.8
2017-12-05 05:30 | 21.2 | 51.7
2017-12-05 05:35 | 21.1 | 51.8
2017-12-05 05:40 | 21.1 | 51.5
2017-12-05 05:45 | 21.1 | 51.4
2017-12-05 05:50 | 21.1 | 51.5
2017-12-05 05:55 | 21.1 | 51.4
2017-12-05 06:00 | 21.1 | 51.5
2017-12-05 06:05 | 21.1 | 51.4
2017-12-05 06:10 | 21.2 | 51.3
2017-12-05 06:15 | 21.2 | 51.2
2017-12-05 06:20 | 21.2 | 51.2
2017-12-05 06:25 | 21.2 | 51.2
2017-12-05 06:30 | 21.2 | 51.3
2017-12-05 06:35 | 21.1 | 51.3
2017-12-05 06:40 | 21.2 | 51.1
2017-12-05 06:45 | 21.2 | 51.4
2017-12-05 06:50 | 21.2 | 51.6
2017-12-05 06:55 | 21.1 | 51.4
2017-12-05 07:00 | 21.2 | 51.4
2017-12-05 07:05 | 21.1 | 51.3
2017-12-05 07:10 | 21.2 | 51.3
2017-12-05 07:15 | 21.2 | 51.3
2017-12-05 07:20 | 21.1 | 51.4
2017-12-05 07:25 | 21.1 | 51.4
2017-12-05 07:30 | 21.1 | 51.3
2017-12-05 07:35 | 21.1 | 51.3
2017-12-05 07:40 | 21.1 | 51.3
2017-12-05 07:45 | 21.1 | 51.4
2017-12-05 07:50 | 21.2 | 51.3
2017-12-05 07:55 | 21.1 | 51.3
2017-12-05 08:00 | 21.1 | 51.4
2017-12-05 08:05 | 21.1 | 51.4
2017-12-05 08:10 | 21.1 | 51.4
2017-12-05 08:15 | 21.2 | 51.5
2017-12-05 08:20 | 21.2 | 51.6
2017-12-05 08:25 | 21 | 51.8
2017-12-05 08:30 | 21 | 51.8
2017-12-05 08:35 | 21 | 51.8
2017-12-05 08:40 | 21.1 | 51.6
2017-12-05 08:45 | 21 | 51.8
2017-12-05 08:50 | 21 | 51.7
2017-12-05 08:55 | 21 | 51.9
2017-12-05 09:00 | 21.1 | 52.1
2017-12-05 09:05 | 21 | 52.2
2017-12-05 09:10 | 21 | 52.1
2017-12-05 09:15 | 21 | 52
2017-12-05 09:20 | 21 | 52
2017-12-05 09:25 | 21 | 52.1
2017-12-05 09:30 | 21 | 52.1
2017-12-05 09:35 | 21 | 52.3
2017-12-05 09:40 | 20.9 | 52
2017-12-05 09:45 | 21 | 52
2017-12-05 09:50 | 20.9 | 51.8
2017-12-05 09:55 | 21 | 51.6
2017-12-05 10:00 | 20.9 | 51.7
2017-12-05 10:05 | 21 | 51.8
2017-12-05 10:10 | 21 | 51.7
2017-12-05 10:15 | 20.9 | 51.9
2017-12-05 10:20 | 20.9 | 51.9
2017-12-05 10:25 | 20.9 | 51.7
2017-12-05 10:30 | 20.8 | 51.9
2017-12-05 10:35 | 20.9 | 51.9
2017-12-05 10:40 | 20.9 | 52.3
2017-12-05 10:45 | 20.9 | 52.2
2017-12-05 10:50 | 20.8 | 52.4
2017-12-05 10:55 | 20.8 | 52.8
2017-12-05 11:00 | 20.9 | 52.4
2017-12-05 11:05 | 20.8 | 52.5
2017-12-05 11:10 | 20.9 | 52.8
2017-12-05 11:15 | 20.8 | 52.6
2017-12-05 11:20 | 20.8 | 52.7
2017-12-05 11:25 | 20.8 | 52.2
2017-12-05 11:30 | 20.8 | 52.5
2017-12-05 11:35 | 20.8 | 52.4
2017-12-05 11:40 | 20.8 | 52.4
2017-12-05 11:45 | 20.8 | 52.9
2017-12-05 11:50 | 20.8 | 52.7
2017-12-05 11:55 | 20.8 | 52.7
2017-12-05 12:00 | 20.8 | 52.9
2017-12-05 12:05 | 20.9 | 52.9
2017-12-05 12:10 | 20.9 | 53.1
2017-12-05 12:15 | 20.9 | 52.8
2017-12-05 12:20 | 20.9 | 52.9
2017-12-05 12:25 | 20.9 | 53.5
2017-12-05 12:30 | 20.9 | 53.2
2017-12-05 12:35 | 20.9 | 53.2
2017-12-05 12:40 | 20.8 | 53.4
2017-12-05 12:45 | 20.8 | 53.2
2017-12-05 12:50 | 20.9 | 53.7
2017-12-05 12:55 | 20.9 | 53.7
2017-12-05 13:00 | 20.9 | 53.2
2017-12-05 13:05 | 21 | 53.4
2017-12-05 13:10 | 20.9 | 53.8
2017-12-05 13:15 | 20.8 | 54.1
2017-12-05 13:20 | 20.8 | 54
2017-12-05 13:25 | 20.8 | 54
2017-12-05 13:30 | 20.8 | 54.2
2017-12-05 13:35 | 20.8 | 54.2
2017-12-05 13:40 | 20.8 | 54.5
2017-12-05 13:45 | 20.8 | 54.5
2017-12-05 13:50 | 20.8 | 54.4
2017-12-05 13:55 | 20.8 | 54.6
2017-12-05 14:00 | 20.7 | 54.7
2017-12-05 14:05 | 20.8 | 54.9
2017-12-05 14:10 | 20.8 | 54.9
2017-12-05 14:15 | 20.7 | 55.1
2017-12-05 14:20 | 20.7 | 55.6
2017-12-05 14:25 | 20.7 | 55.2
2017-12-05 14:30 | 20.8 | 55.3
2017-12-05 14:35 | 20.8 | 54.6
2017-12-05 14:40 | 20.7 | 54.8
2017-12-05 14:45 | 20.7 | 54.9
2017-12-05 14:50 | 20.7 | 54.7
2017-12-05 14:55 | 20.7 | 54.8
2017-12-05 15:00 | 20.8 | 54.7
2017-12-05 15:05 | 20.8 | 54.4
2017-12-05 15:10 | 20.7 | 54.5
2017-12-05 15:15 | 20.8 | 54.6
2017-12-05 15:20 | 20.8 | 54.6
2017-12-05 15:25 | 20.8 | 54.2
2017-12-05 15:30 | 20.8 | 54.8
2017-12-05 15:35 | 20.8 | 54.2
2017-12-05 15:40 | 20.7 | 54
2017-12-05 15:45 | 20.7 | 54.8
2017-12-05 15:50 | 20.7 | 54.3
2017-12-05 15:55 | 20.7 | 53.8
2017-12-05 16:00 | 20.7 | 53.8
2017-12-05 16:05 | 20.6 | 54.6
2017-12-05 16:10 | 20.6 | 54
2017-12-05 16:15 | 20.7 | 54.2
2017-12-05 16:20 | 20.7 | 54.4
2017-12-05 16:25 | 20.6 | 54.3
2017-12-05 16:30 | 20.7 | 54.2
2017-12-05 16:35 | 20.7 | 54.1
2017-12-05 16:40 | 20.7 | 53.7
2017-12-05 16:45 | 20.7 | 53.9

## Conditions
| Temp |  | RH
All GMH5I readings | #REF! | #REF!
Room | #REF! | #REF! | #REF! | #REF!