## Commentary
2020-07-29
Worksheet 'pyUBreadings' requires the user to enter all the dial readings and settings of the UB as well as estimates of the uncertainty in temperature and in the temperature coefficients for resistance and inductance. In general it is set up not to apply any temperature corrections, just to evaluate the uncertainty.
Use the UniversalBridge3 project to process the input data. This uses the latest ubdict.csv file for corrections to the Universal Bridge. Running 'calculate_imp.py' produces a new 'indResults.xlsx'. Paste these results into the 'pyUBresults' worksheet of this spreadsheet and, if correctly pasted, should immediately produce the report in the'pyReport' worksheet. Note that there are two named blocks in 'pyUBresults' and so the new data should just be pasted over the old data in order to preserve the inductor1 and inductor2 named blocks.
You should check that the input data listed in 'pyUBresults' matches the input data in 'pyUBreadings', just in case the wrong file names are enterd in 'calculate_imp.py'.

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | Muirhead A-517_ TW |  | 4: Client; 5: Babcock; 10: Enter, say, 5 ppm as 5e-6
Date from | 2021-04-30 |  | 4: Job number; 5: 96330149; 10: 1 sigma uncertainties from specifications
to | 2021-04-30 |  | 4: S Number; 5: S22091
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
L0 | 1591.5 | 4Z | 2 | 97 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 4Z | 101664 | 195360 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 4Z | 200788 | 280525 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 4Z | 295216 | 344559 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 4Z | 396262 | 405112 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 4Z | 494255 | 458953 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 4Z | 595223 | 509993 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 4Z | 693271 | 557033 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 4Z | 791139 | 602627 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 4Z | 893586 | 647741 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 4Z | 1013237 | 698740 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L0 | 1591.5 | 5Z | 0 | -20 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L1 | 1591.5 | 5Z | 100956 | 165599 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L2 | 1591.5 | 5Z | 204539 | 241310 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L3 | 1591.5 | 5Z | 302342 | 299040 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L4 | 1591.5 | 5Z | 395720 | 347550 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L5 | 1591.5 | 5Z | 492735 | 393669 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L6 | 1591.5 | 5Z | 592525 | 438249 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L7 | 1591.5 | 5Z | 691041 | 479415 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L8 | 1591.5 | 5Z | 783785 | 517505 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L9 | 1591.5 | 5Z | 876862 | 554555 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10
L10 | 1591.5 | 5Z | 980684 | 594698 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
L0 | 1591.5 | 4Z | 2 | 97 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-07 | 0.0012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 4Z | 101664 | 195360 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.010168 | 9.0609e-07 | 2.0134 | 19.5277 | 0.11583 | 2.084 | 1.9092e-07 | 0.0016398 | 4.7459 | 70.6371 | L11591.5 | 1000 | 19527.704 | 115.8331 | 2.084 | 1000000 | 10168.4275 | 0.90609 | 2.0134 | 1591.5
L2 | 1591.5 | 4Z | 200788 | 280525 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.020083 | 1.64e-06 | 2.0349 | 28.0448 | 0.16635 | 2.084 | 3.0926e-07 | 0.0024016 | 5.3029 | 69.2634 | L21591.5 | 1000 | 28044.8163 | 166.3453 | 2.084 | 1000000 | 20082.9994 | 1.64 | 2.0349 | 1591.5
L3 | 1591.5 | 4Z | 295216 | 344559 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.029528 | 2.3693e-06 | 2.0413 | 34.4487 | 0.20436 | 2.0838 | 4.3346e-07 | 0.0032453 | 5.4661 | 62.9719 | L31591.5 | 1000 | 34448.6767 | 204.3631 | 2.0838 | 1000000 | 29527.8684 | 2.3693 | 2.0413 | 1591.5
L4 | 1591.5 | 4Z | 396262 | 405112 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.039635 | 3.1585e-06 | 2.0441 | 40.5044 | 0.24034 | 2.0837 | 5.7039e-07 | 0.0041966 | 5.5374 | 57.271 | L41591.5 | 1000 | 40504.4121 | 240.3414 | 2.0837 | 1000000 | 39634.6824 | 3.1585 | 2.0441 | 1591.5
L5 | 1591.5 | 4Z | 494255 | 458953 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.049436 | 3.9273e-06 | 2.0454 | 45.8889 | 0.27235 | 2.0835 | 7.0493e-07 | 0.0051413 | 5.5712 | 52.9738 | L51591.5 | 1000 | 45888.8992 | 272.3525 | 2.0835 | 1000000 | 49436.1295 | 3.9273 | 2.0454 | 1591.5
L6 | 1591.5 | 4Z | 595223 | 509993 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.059535 | 4.7212e-06 | 2.0462 | 50.9933 | 0.30272 | 2.0833 | 8.4446e-07 | 0.0061266 | 5.5908 | 49.411 | L61591.5 | 1000 | 50993.2662 | 302.7189 | 2.0833 | 1000000 | 59535.1418 | 4.7212 | 2.0462 | 1591.5
L7 | 1591.5 | 4Z | 693271 | 557033 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.069342 | 5.4932e-06 | 2.0466 | 55.6976 | 0.33072 | 2.0832 | 9.8048e-07 | 0.0070903 | 5.6025 | 46.6442 | L71591.5 | 1000 | 55697.6044 | 330.7205 | 2.0832 | 1000000 | 69342.0902 | 5.4932 | 2.0466 | 1591.5
L8 | 1591.5 | 4Z | 791139 | 602627 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.079131 | 6.2643e-06 | 2.0469 | 60.2573 | 0.35787 | 2.083 | 1.1166e-06 | 0.0080567 | 5.6103 | 44.4194 | L81591.5 | 1000 | 60257.3322 | 357.8717 | 2.083 | 1000000 | 79131.0345 | 6.2643 | 2.0469 | 1591.5
L9 | 1591.5 | 4Z | 893586 | 647741 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.089378 | 7.072e-06 | 2.0472 | 64.7691 | 0.38475 | 2.0829 | 1.2593e-06 | 0.0090712 | 5.616 | 42.4146 | L91591.5 | 1000 | 64769.0566 | 384.7532 | 2.0829 | 1000000 | 89377.9793 | 7.072 | 2.0472 | 1591.5
L10 | 1591.5 | 4Z | 1013237 | 698740 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10135 | 8.0157e-06 | 2.0473 | 69.8693 | 0.41516 | 2.0827 | 1.4261e-06 | 0.010259 | 5.6207 | 40.4674 | L101591.5 | 1000 | 69869.3233 | 415.155 | 2.0827 | 1000000 | 101345.7009 | 8.0157 | 2.0473 | 1591.5
L0 | 1591.5 | 5Z | 0 | -20 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2571e-06 | 0.012571 | 0 | 0 | L01591.5 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
L1 | 1591.5 | 5Z | 100956 | 165599 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.10098 | 9.0058e-06 | 2.0132 | 165.6309 | 0.98253 | 2.0837 | 1.899e-06 | 0.01629 | 4.7424 | 60.3147 | L11591.5 | 1 | 165.6309 | 0.98253 | 2.0837 | 1000000 | 100978.1204 | 9.0058 | 2.0132 | 1591.5
L2 | 1591.5 | 5Z | 204539 | 241310 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.20458 | 1.668e-05 | 2.0355 | 241.3474 | 1.4318 | 2.0837 | 3.1372e-06 | 0.024248 | 5.3169 | 59.0453 | L21591.5 | 1 | 241.3474 | 1.4318 | 2.0837 | 1000000 | 204583.8164 | 16.6802 | 2.0355 | 1591.5
L3 | 1591.5 | 5Z | 302342 | 299040 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.30241 | 2.4241e-05 | 2.0417 | 299.0815 | 1.7747 | 2.0835 | 4.4265e-06 | 0.033016 | 5.4764 | 53.7539 | L31591.5 | 1 | 299.0815 | 1.7747 | 2.0835 | 1000000 | 302408.2459 | 24.2412 | 2.0417 | 1591.5
L4 | 1591.5 | 5Z | 395720 | 347550 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.39581 | 3.1536e-05 | 2.0442 | 347.595 | 2.0632 | 2.0833 | 5.6927e-06 | 0.041813 | 5.5396 | 49.3445 | L41591.5 | 1 | 347.595 | 2.0632 | 2.0833 | 1000000 | 395806.7059 | 31.5355 | 2.0442 | 1591.5
L5 | 1591.5 | 5Z | 492735 | 393669 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.49284 | 3.9146e-05 | 2.0455 | 393.7173 | 2.3378 | 2.0831 | 7.0244e-06 | 0.05116 | 5.5729 | 45.6952 | L51591.5 | 1 | 393.7173 | 2.3378 | 2.0831 | 1000000 | 492842.9628 | 39.1465 | 2.0455 | 1591.5
L6 | 1591.5 | 5Z | 592525 | 438249 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.59265 | 4.6993e-05 | 2.0462 | 438.3005 | 2.6033 | 2.0829 | 8.4033e-06 | 0.060894 | 5.5922 | 42.7517 | L61591.5 | 1 | 438.3005 | 2.6033 | 2.0829 | 1000000 | 592654.8277 | 46.9928 | 2.0462 | 1591.5
L7 | 1591.5 | 5Z | 691041 | 479415 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.69119 | 5.4749e-05 | 2.0467 | 479.4695 | 2.8487 | 2.0827 | 9.7698e-06 | 0.070573 | 5.6039 | 40.3659 | L71591.5 | 1 | 479.4695 | 2.8487 | 2.0827 | 1000000 | 691192.4135 | 54.749 | 2.0467 | 1591.5
L8 | 1591.5 | 5Z | 783785 | 517505 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.78396 | 6.2057e-05 | 2.047 | 517.5622 | 3.0759 | 2.0825 | 1.1059e-05 | 0.079727 | 5.6113 | 38.5805 | L81591.5 | 1 | 517.5622 | 3.0759 | 2.0825 | 1000000 | 783956.7346 | 62.0566 | 2.047 | 1591.5
L9 | 1591.5 | 5Z | 876862 | 554555 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.87705 | 6.9394e-05 | 2.0472 | 554.6149 | 3.297 | 2.0823 | 1.2355e-05 | 0.088941 | 5.6165 | 37.0692 | L91591.5 | 1 | 554.6149 | 3.297 | 2.0823 | 1000000 | 877054.1286 | 69.3942 | 2.0472 | 1591.5
L10 | 1591.5 | 5Z | 980684 | 594698 | 1591 | 20 | 0.5 | 10 | 5e-05 | 5e-06 | 10 | 0.004 | 0.0004 | 10 | 0.9809 | 7.7582e-05 | 2.0473 | 594.7608 | 3.5366 | 2.0821 | 1.3803e-05 | 0.099242 | 5.6207 | 35.6365 | L101591.5 | 1 | 594.7608 | 3.5366 | 2.0821 | 1000000 | 980898.877 | 77.582 | 2.0473 | 1591.5

## pyReport
| TW | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 |  | 10 | 101.3457 | 0.0080157 | 2.0473 | 69.8693 | 0.41516 | 2.0827
L9 |  | 9 | 89.378 | 0.007072 | 2.0472 | 64.7691 | 0.38475 | 2.0829
L8 |  | 8 | 79.131 | 0.0062643 | 2.0469 | 60.2573 | 0.35787 | 2.083
L7 |  | 7 | 69.3421 | 0.0054932 | 2.0466 | 55.6976 | 0.33072 | 2.0832
L6 |  | 6 | 59.5351 | 0.0047212 | 2.0462 | 50.9933 | 0.30272 | 2.0833
L5 |  | 5 | 49.4361 | 0.0039273 | 2.0454 | 45.8889 | 0.27235 | 2.0835
L4 |  | 4 | 39.6347 | 0.0031585 | 2.0441 | 40.5044 | 0.24034 | 2.0837
L3 |  | 3 | 29.5279 | 0.0023693 | 2.0413 | 34.4487 | 0.20436 | 2.0838
L2 |  | 2 | 20.083 | 0.00164 | 2.0349 | 28.0448 | 0.16635 | 2.084
L1 |  | 1 | 10.1684 | 0.00090609 | 2.0134 | 19.5277 | 0.11583 | 2.084
|  | Tap | Inductance Change (mH) | Expanded Uncertainty (mH) | Coverage Factor | Resistance Change (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
L10 |  | 10 | 980.8989 | 0.077582 | 2.0473 | 594.7608 | 3.5366 | 2.0821
L9 |  | 9 | 877.0541 | 0.069394 | 2.0472 | 554.6149 | 3.297 | 2.0823
L8 |  | 8 | 783.9567 | 0.062057 | 2.047 | 517.5622 | 3.0759 | 2.0825
L7 |  | 7 | 691.1924 | 0.054749 | 2.0467 | 479.4695 | 2.8487 | 2.0827
L6 |  | 6 | 592.6548 | 0.046993 | 2.0462 | 438.3005 | 2.6033 | 2.0829
L5 |  | 5 | 492.843 | 0.039146 | 2.0455 | 393.7173 | 2.3378 | 2.0831
L4 |  | 4 | 395.8067 | 0.031536 | 2.0442 | 347.595 | 2.0632 | 2.0833
L3 |  | 3 | 302.4082 | 0.024241 | 2.0417 | 299.0815 | 1.7747 | 2.0835
L2 |  | 2 | 204.5838 | 0.01668 | 2.0355 | 241.3474 | 1.4318 | 2.0837
L1 |  | 1 | 100.9781 | 0.0090058 | 2.0132 | 165.6309 | 0.98253 | 2.0837

## exampleU
Budget for L8 1591.5
ureal(60.08704790830853,0.1713140491986457,20.451673221652594)
uuttempL8 | 0.12049
temperature | 0.12043
Amp_phase_err | 0.013329
rtempcoL8 | 0.012019
Adialresureal(62.0,10.0,10.0) | 0.0011776
Adialresureal(600890.0,10.0,10.0) | 0.0011776
Adial_lin62 | 0.0010001
Adial_lin600890 | 0.0010001
Amperr | 0.00084121
R4Aacdc | 0.00015022
G1acdc | 0.00015022
G2acdc | 0.00015022
G1alpha | 6.0087e-05
G1stab | 6.0087e-05
R4Aalpha | 6.0087e-05
G2alpha | 6.0087e-05
R4Astab | 6.0087e-05
G2stab | 6.0087e-05
cal_temp | 4.2488e-05
G2err | 3.0043e-05
G1err | 3.0043e-05
R4Aerr | 3.0042e-05
uuttempL0 | 1.2432e-05
rtempcoL0 | 1.2401e-06
C1err | 0
C1alpha | 0
C1stab | 0
Bdial_lin0 | 0
Bdialresureal(0.0,1.0,10.0, label='Bdial_lin0') | 0
Bdial_lin788633 | 0
Bdialresureal(788633.0,1.0,10.0) | 0
ureal(0.0788795386314938,3.031875053076101e-06,27.783955120984952)
uuttempL8 | 1.9769e-06
temperature | 1.9211e-06
Amperr | 1.1043e-06
C1stab | 3.944e-07
R4Aacdc | 1.972e-07
G2acdc | 1.972e-07
xtempcoL8 | 1.972e-07
Bdialres0.0 | 1.1778e-07
Bdialres788633.0 | 1.1778e-07
Amp_phase_err | 1.0161e-07
Bdial_lin0 | 1.0002e-07
Bdial_lin788633 | 1.0002e-07
C1alpha | 7.888e-08
R4Aalpha | 7.888e-08
G2alpha | 7.888e-08
R4Astab | 7.888e-08
G2stab | 7.888e-08
C1err | 7.8868e-08
cal_temp | 5.5776e-08
G2err | 3.9439e-08
R4Aerr | 3.9438e-08
G1err | 0
G1acdc | 0
G1alpha | 0
G1stab | 0
Adial_lin62 | 0
Adialres62.0 | 0
uuttempL0 | 0
xtempcoL0 | 0
Adial_lin600890 | 0
Adialres600890.0 | 0

## Room Conditions
2021-04-19 | 20.3 | 52.3
2021-04-19 00:05 | 20.3 | 52.3
2021-04-19 00:10 | 20.3 | 52.3 |  | RH correction | -2.128 | from Y:\Quality\TECHPROC\E058-Lab Air Cond\loggercalibration_2017.xls
2021-04-19 00:15 | 20.3 | 52.2 | Temperature | RH
2021-04-19 00:20 | 20.3 | 52.2 | #REF! | #REF! | #REF!
2021-04-19 00:25 | 20.3 | 51.9 | #REF! | #REF!
2021-04-19 00:30 | 20.3 | 51.7
2021-04-19 00:35 | 20.4 | 51.6 | from GMH
2021-04-19 00:40 | 20.4 | 51.5
2021-04-19 00:45 | 20.4 | 51.3
2021-04-19 00:50 | 20.3 | 51.3
2021-04-19 00:55 | 20.3 | 51.2
2021-04-19 01:00 | 20.3 | 51.3
2021-04-19 01:05 | 20.3 | 51.3
2021-04-19 01:10 | 20.3 | 51.4
2021-04-19 01:15 | 20.3 | 51.6
2021-04-19 01:20 | 20.3 | 51.8
2021-04-19 01:25 | 20.3 | 52
2021-04-19 01:30 | 20.3 | 52
2021-04-19 01:35 | 20.3 | 52
2021-04-19 01:40 | 20.3 | 52
2021-04-19 01:45 | 20.3 | 52
2021-04-19 01:50 | 20.3 | 52.1
2021-04-19 01:55 | 20.3 | 52.1
2021-04-19 02:00 | 20.3 | 52.3
2021-04-19 02:05 | 20.3 | 52.5
2021-04-19 02:10 | 20.3 | 52.8
2021-04-19 02:15 | 20.3 | 52.9
2021-04-19 02:20 | 20.3 | 53.1
2021-04-19 02:25 | 20.3 | 53.3
2021-04-19 02:30 | 20.3 | 53.4
2021-04-19 02:35 | 20.3 | 53.6
2021-04-19 02:40 | 20.3 | 53.7
2021-04-19 02:45 | 20.3 | 53.8
2021-04-19 02:50 | 20.3 | 53.9
2021-04-19 02:55 | 20.3 | 54
2021-04-19 03:00 | 20.3 | 54
2021-04-19 03:05 | 20.3 | 54
2021-04-19 03:10 | 20.4 | 53.8
2021-04-19 03:15 | 20.3 | 53.9
2021-04-19 03:20 | 20.3 | 54
2021-04-19 03:25 | 20.3 | 54
2021-04-19 03:30 | 20.3 | 54
2021-04-19 03:35 | 20.3 | 53.9
2021-04-19 03:40 | 20.3 | 53.9
2021-04-19 03:45 | 20.4 | 53.8
2021-04-19 03:50 | 20.4 | 53.7
2021-04-19 03:55 | 20.3 | 53.6
2021-04-19 04:00 | 20.3 | 53.6
2021-04-19 04:05 | 20.3 | 53.6
2021-04-19 04:10 | 20.3 | 53.6
2021-04-19 04:15 | 20.3 | 53.5
2021-04-19 04:20 | 20.4 | 53.4
2021-04-19 04:25 | 20.4 | 53.4
2021-04-19 04:30 | 20.3 | 53.5
2021-04-19 04:35 | 20.4 | 53.4
2021-04-19 04:40 | 20.4 | 53.5
2021-04-19 04:45 | 20.4 | 53.5
2021-04-19 04:50 | 20.3 | 53.6
2021-04-19 04:55 | 20.3 | 53.7
2021-04-19 05:00 | 20.3 | 53.7
2021-04-19 05:05 | 20.4 | 53.7
2021-04-19 05:10 | 20.3 | 53.8
2021-04-19 05:15 | 20.3 | 54
2021-04-19 05:20 | 20.3 | 54
2021-04-19 05:25 | 20.3 | 54.1
2021-04-19 05:30 | 20.3 | 54.1
2021-04-19 05:35 | 20.4 | 53.9
2021-04-19 05:40 | 20.4 | 53.9
2021-04-19 05:45 | 20.4 | 53.9
2021-04-19 05:50 | 20.3 | 54
2021-04-19 05:55 | 20.3 | 53.8
2021-04-19 06:00 | 20.3 | 53.7
2021-04-19 06:05 | 20.3 | 53.6
2021-04-19 06:10 | 20.3 | 53.5
2021-04-19 06:15 | 20.3 | 53.4
2021-04-19 06:20 | 20.3 | 53.2
2021-04-19 06:25 | 20.4 | 53.1
2021-04-19 06:30 | 20.4 | 53.2
2021-04-19 06:35 | 20.4 | 53.2
2021-04-19 06:40 | 20.3 | 53.4
2021-04-19 06:45 | 20.3 | 53.4
2021-04-19 06:50 | 20.3 | 53.4
2021-04-19 06:55 | 20.3 | 53.5
2021-04-19 07:00 | 20.2 | 53.8
2021-04-19 07:05 | 20.3 | 53.8
2021-04-19 07:10 | 20.3 | 53.6
2021-04-19 07:15 | 20.3 | 53.6
2021-04-19 07:20 | 20.3 | 53.6
2021-04-19 07:25 | 20.4 | 53.4
2021-04-19 07:30 | 20.3 | 53.4
2021-04-19 07:35 | 20.3 | 53.4
2021-04-19 07:40 | 20.4 | 53.4
2021-04-19 07:45 | 20.3 | 53.5
2021-04-19 07:50 | 20.4 | 53.7
2021-04-19 07:55 | 20.4 | 53.9
2021-04-19 08:00 | 20.3 | 54.2
2021-04-19 08:05 | 20.3 | 54.5
2021-04-19 08:10 | 20.3 | 54.7
2021-04-19 08:15 | 20.4 | 54.6
2021-04-19 08:20 | 20.3 | 54.9
2021-04-19 08:25 | 20.3 | 55.2
2021-04-19 08:30 | 20.3 | 55.3
2021-04-19 08:35 | 20.3 | 55.3
2021-04-19 08:40 | 20.3 | 55.5
2021-04-19 08:45 | 20.3 | 55.5
2021-04-19 08:50 | 20.3 | 55.8
2021-04-19 08:55 | 20.3 | 55.7
2021-04-19 09:00 | 20.3 | 55.9
2021-04-19 09:05 | 20.3 | 56
2021-04-19 09:10 | 20.3 | 56
2021-04-19 09:15 | 20.3 | 56
2021-04-19 09:20 | 20.3 | 56.1
2021-04-19 09:25 | 20.3 | 56.2
2021-04-19 09:30 | 20.2 | 56.3
2021-04-19 09:35 | 20.2 | 56.4
2021-04-19 09:40 | 20.2 | 56.3
2021-04-19 09:45 | 20.2 | 56.3
2021-04-19 09:50 | 20.3 | 56.2
2021-04-19 09:55 | 20.3 | 56
2021-04-19 10:00 | 20.3 | 55.9
2021-04-19 10:05 | 20.3 | 55.6
2021-04-19 10:10 | 20.3 | 55.4
2021-04-19 10:15 | 20.3 | 55.4
2021-04-19 10:20 | 20.3 | 55.2
2021-04-19 10:25 | 20.3 | 55
2021-04-19 10:30 | 20.3 | 54.8
2021-04-19 10:35 | 20.3 | 54.7
2021-04-19 10:40 | 20.3 | 54.8
2021-04-19 10:45 | 20.3 | 54.9
2021-04-19 10:50 | 20.3 | 55.1
2021-04-19 10:55 | 20.2 | 55.4
2021-04-19 11:00 | 20.3 | 55.4
2021-04-19 11:05 | 20.3 | 55.6
2021-04-19 11:10 | 20.3 | 55.7
2021-04-19 11:15 | 20.3 | 56.1
2021-04-19 11:20 | 20.3 | 56.2
2021-04-19 11:25 | 20.3 | 56.2
2021-04-19 11:30 | 20.3 | 56.6
2021-04-19 11:35 | 20.3 | 56.6
2021-04-19 11:40 | 20.3 | 57.1
2021-04-19 11:45 | 20.2 | 57.4
2021-04-19 11:50 | 20.3 | 57.6
2021-04-19 11:55 | 20.2 | 57.8
2021-04-19 12:00 | 20.2 | 57.9
2021-04-19 12:05 | 20.2 | 57.8
2021-04-19 12:10 | 20.2 | 58
2021-04-19 12:15 | 20.3 | 57.8
2021-04-19 12:20 | 20.2 | 57.8
2021-04-19 12:25 | 20.2 | 57.7
2021-04-19 12:30 | 20.2 | 57.6
2021-04-19 12:35 | 20.2 | 57.4
2021-04-19 12:40 | 20.2 | 57.3
2021-04-19 12:45 | 20.3 | 57.3
2021-04-19 12:50 | 20.2 | 57.3
2021-04-19 12:55 | 20.3 | 57.1
2021-04-19 13:00 | 20.2 | 57.1
2021-04-19 13:05 | 20.2 | 57.1
2021-04-19 13:10 | 20.2 | 57.2
2021-04-19 13:15 | 20.2 | 57.3
2021-04-19 13:20 | 20.2 | 57.2
2021-04-19 13:25 | 20.2 | 57.1
2021-04-19 13:30 | 20.2 | 57.1
2021-04-19 13:35 | 20.3 | 56.9
2021-04-19 13:40 | 20.3 | 56.8
2021-04-19 13:45 | 20.3 | 56.7
2021-04-19 13:50 | 20.2 | 56.7
2021-04-19 13:55 | 20.2 | 56.6
2021-04-19 14:00 | 20.2 | 56.4
2021-04-19 14:05 | 20.3 | 56.3
2021-04-19 14:10 | 20.2 | 56.2
2021-04-19 14:15 | 20.3 | 56
2021-04-19 14:20 | 20.3 | 55.8
2021-04-19 14:25 | 20.3 | 55.8
2021-04-19 14:30 | 20.2 | 56
2021-04-19 14:35 | 20.2 | 56.1
2021-04-19 14:40 | 20.2 | 56.2
2021-04-19 14:45 | 20.2 | 56.3
2021-04-19 14:50 | 20.3 | 56.3
2021-04-19 14:55 | 20.2 | 56.6
2021-04-19 15:00 | 20.2 | 56.7
2021-04-19 15:05 | 20.2 | 56.8
2021-04-19 15:10 | 20.2 | 56.8
2021-04-19 15:15 | 20.3 | 56.9
2021-04-19 15:20 | 20.3 | 57
2021-04-19 15:25 | 20.3 | 56.9
2021-04-19 15:30 | 20.3 | 57.1
2021-04-19 15:35 | 20.2 | 57.2
2021-04-19 15:40 | 20.2 | 57.3
2021-04-19 15:45 | 20.2 | 57.1
2021-04-19 15:50 | 20.2 | 57.1
2021-04-19 15:55 | 20.3 | 57
2021-04-19 16:00 | 20.3 | 57.3
2021-04-19 16:05 | 20.3 | 57.2
2021-04-19 16:10 | 20.3 | 57.4
2021-04-19 16:15 | 20.2 | 57.6
2021-04-19 16:20 | 20.2 | 58
2021-04-19 16:25 | 20.2 | 57.9
2021-04-19 16:30 | 20.2 | 57.8
2021-04-19 16:35 | 20.3 | 57.7
2021-04-19 16:40 | 20.3 | 57.7
2021-04-19 16:45 | 20.3 | 57.9
2021-04-19 16:50 | 20.3 | 58
2021-04-19 16:55 | 20.2 | 58.2
2021-04-19 17:00 | 20.2 | 58.2
2021-04-19 17:05 | 20.3 | 58.1
2021-04-19 17:10 | 20.3 | 58.1
2021-04-19 17:15 | 20.2 | 58.1
2021-04-19 17:20 | 20.3 | 58
2021-04-19 17:25 | 20.2 | 58
2021-04-19 17:30 | 20.2 | 57.9
2021-04-19 17:35 | 20.2 | 57.9
2021-04-19 17:40 | 20.2 | 57.9
2021-04-19 17:45 | 20.3 | 57.7
2021-04-19 17:50 | 20.3 | 57.5
2021-04-19 17:55 | 20.3 | 57.5
2021-04-19 18:00 | 20.4 | 57.4
2021-04-19 18:05 | 20.3 | 57.7
2021-04-19 18:10 | 20.3 | 57.8
2021-04-19 18:15 | 20.3 | 57.7
2021-04-19 18:20 | 20.2 | 57.9
2021-04-19 18:25 | 20.2 | 57.9
2021-04-19 18:30 | 20.2 | 58
2021-04-19 18:35 | 20.2 | 58
2021-04-19 18:40 | 20.2 | 58
2021-04-19 18:45 | 20.2 | 57.9
2021-04-19 18:50 | 20.3 | 57.8
2021-04-19 18:55 | 20.2 | 57.9
2021-04-19 19:00 | 20.2 | 57.9
2021-04-19 19:05 | 20.2 | 57.9
2021-04-19 19:10 | 20.2 | 57.9
2021-04-19 19:15 | 20.3 | 57.7
2021-04-19 19:20 | 20.3 | 57.7
2021-04-19 19:25 | 20.3 | 57.8
2021-04-19 19:30 | 20.3 | 57.6
2021-04-19 19:35 | 20.3 | 57.8
2021-04-19 19:40 | 20.3 | 57.6
2021-04-19 19:45 | 20.3 | 57.5
2021-04-19 19:50 | 20.4 | 57.4
2021-04-19 19:55 | 20.4 | 57.4
2021-04-19 20:00 | 20.4 | 57.3
2021-04-19 20:05 | 20.3 | 57.5
2021-04-19 20:10 | 20.3 | 57.3
2021-04-19 20:15 | 20.3 | 57.4
2021-04-19 20:20 | 20.3 | 57.3
2021-04-19 20:25 | 20.4 | 57.2
2021-04-19 20:30 | 20.4 | 57.1
2021-04-19 20:35 | 20.4 | 57
2021-04-19 20:40 | 20.4 | 57
2021-04-19 20:45 | 20.3 | 57.1
2021-04-19 20:50 | 20.4 | 57
2021-04-19 20:55 | 20.4 | 57
2021-04-19 21:00 | 20.3 | 56.9
2021-04-19 21:05 | 20.4 | 56.9
2021-04-19 21:10 | 20.3 | 56.9
2021-04-19 21:15 | 20.4 | 56.8
2021-04-19 21:20 | 20.4 | 56.8
2021-04-19 21:25 | 20.4 | 56.8
2021-04-19 21:30 | 20.3 | 57
2021-04-19 21:35 | 20.3 | 56.9
2021-04-19 21:40 | 20.4 | 56.8
2021-04-19 21:45 | 20.4 | 56.8
2021-04-19 21:50 | 20.4 | 56.9
2021-04-19 21:55 | 20.3 | 56.9
2021-04-19 22:00 | 20.3 | 57
2021-04-19 22:05 | 20.4 | 56.7
2021-04-19 22:10 | 20.4 | 56.8
2021-04-19 22:15 | 20.3 | 56.9
2021-04-19 22:20 | 20.3 | 57
2021-04-19 22:25 | 20.3 | 57
2021-04-19 22:30 | 20.3 | 57.1
2021-04-19 22:35 | 20.3 | 57.1
2021-04-19 22:40 | 20.3 | 57.1
2021-04-19 22:45 | 20.3 | 57
2021-04-19 22:50 | 20.3 | 56.9
2021-04-19 22:55 | 20.3 | 57
2021-04-19 23:00 | 20.3 | 57.1
2021-04-19 23:05 | 20.3 | 56.9
2021-04-19 23:10 | 20.3 | 56.9
2021-04-19 23:15 | 20.3 | 57
2021-04-19 23:20 | 20.3 | 57
2021-04-19 23:25 | 20.3 | 57.1
2021-04-19 23:30 | 20.2 | 57.1
2021-04-19 23:35 | 20.3 | 57
2021-04-19 23:40 | 20.3 | 57
2021-04-19 23:45 | 20.3 | 57
2021-04-19 23:50 | 20.3 | 57
2021-04-19 23:55 | 20.3 | 57
2021-04-20 | 20.3 | 57
2021-04-20 00:05 | 20.3 | 56.9
2021-04-20 00:10 | 20.3 | 56.9
2021-04-20 00:15 | 20.3 | 57
2021-04-20 00:20 | 20.3 | 56.9
2021-04-20 00:25 | 20.3 | 57
2021-04-20 00:30 | 20.3 | 56.9
2021-04-20 00:35 | 20.3 | 56.9
2021-04-20 00:40 | 20.3 | 56.8
2021-04-20 00:45 | 20.3 | 56.7
2021-04-20 00:50 | 20.3 | 56.8
2021-04-20 00:55 | 20.3 | 56.8
2021-04-20 01:00 | 20.3 | 56.8
2021-04-20 01:05 | 20.4 | 56.7
2021-04-20 01:10 | 20.3 | 56.7
2021-04-20 01:15 | 20.3 | 56.9
2021-04-20 01:20 | 20.3 | 56.9
2021-04-20 01:25 | 20.3 | 56.8
2021-04-20 01:30 | 20.3 | 56.8
2021-04-20 01:35 | 20.3 | 56.8
2021-04-20 01:40 | 20.3 | 56.8
2021-04-20 01:45 | 20.3 | 56.8
2021-04-20 01:50 | 20.3 | 56.8
2021-04-20 01:55 | 20.3 | 56.6
2021-04-20 02:00 | 20.3 | 56.8
2021-04-20 02:05 | 20.4 | 56.8
2021-04-20 02:10 | 20.3 | 56.7
2021-04-20 02:15 | 20.4 | 56.7
2021-04-20 02:20 | 20.4 | 56.7
2021-04-20 02:25 | 20.3 | 56.9
2021-04-20 02:30 | 20.4 | 56.8
2021-04-20 02:35 | 20.3 | 56.9
2021-04-20 02:40 | 20.4 | 56.9
2021-04-20 02:45 | 20.3 | 57
2021-04-20 02:50 | 20.3 | 57
2021-04-20 02:55 | 20.3 | 57
2021-04-20 03:00 | 20.3 | 57
2021-04-20 03:05 | 20.3 | 57
2021-04-20 03:10 | 20.3 | 57.1
2021-04-20 03:15 | 20.3 | 57.1
2021-04-20 03:20 | 20.3 | 57.1
2021-04-20 03:25 | 20.3 | 57.2
2021-04-20 03:30 | 20.2 | 57.2
2021-04-20 03:35 | 20.3 | 57.2
2021-04-20 03:40 | 20.3 | 57.2
2021-04-20 03:45 | 20.3 | 57.2
2021-04-20 03:50 | 20.3 | 57.2
2021-04-20 03:55 | 20.3 | 57.3
2021-04-20 04:00 | 20.3 | 57.2
2021-04-20 04:05 | 20.3 | 57.2
2021-04-20 04:10 | 20.2 | 57.3
2021-04-20 04:15 | 20.3 | 57.2
2021-04-20 04:20 | 20.3 | 57.2
2021-04-20 04:25 | 20.3 | 57.1
2021-04-20 04:30 | 20.3 | 57.1
2021-04-20 04:35 | 20.3 | 57.3
2021-04-20 04:40 | 20.3 | 57.3
2021-04-20 04:45 | 20.3 | 57.2
2021-04-20 04:50 | 20.3 | 57.1
2021-04-20 04:55 | 20.3 | 57.1
2021-04-20 05:00 | 20.3 | 57.1
2021-04-20 05:05 | 20.3 | 57.2
2021-04-20 05:10 | 20.3 | 57.3
2021-04-20 05:15 | 20.3 | 57.2
2021-04-20 05:20 | 20.4 | 57.1
2021-04-20 05:25 | 20.3 | 57.1
2021-04-20 05:30 | 20.3 | 57.1
2021-04-20 05:35 | 20.3 | 57.2
2021-04-20 05:40 | 20.3 | 57.1
2021-04-20 05:45 | 20.3 | 57.1
2021-04-20 05:50 | 20.3 | 57.1
2021-04-20 05:55 | 20.3 | 57.1
2021-04-20 06:00 | 20.3 | 57.1
2021-04-20 06:05 | 20.3 | 57.1
2021-04-20 06:10 | 20.3 | 57
2021-04-20 06:15 | 20.3 | 56.9
2021-04-20 06:20 | 20.4 | 56.8
2021-04-20 06:25 | 20.4 | 56.7
2021-04-20 06:30 | 20.3 | 56.8
2021-04-20 06:35 | 20.3 | 56.9
2021-04-20 06:40 | 20.3 | 56.8
2021-04-20 06:45 | 20.3 | 56.7
2021-04-20 06:50 | 20.4 | 56.6
2021-04-20 06:55 | 20.3 | 56.8
2021-04-20 07:00 | 20.3 | 56.9
2021-04-20 07:05 | 20.3 | 56.8
2021-04-20 07:10 | 20.3 | 56.7
2021-04-20 07:15 | 20.4 | 56.6
2021-04-20 07:20 | 20.3 | 56.6
2021-04-20 07:25 | 20.3 | 56.6
2021-04-20 07:30 | 20.3 | 56.7
2021-04-20 07:35 | 20.3 | 56.6
2021-04-20 07:40 | 20.3 | 56.5
2021-04-20 07:45 | 20.3 | 56.7
2021-04-20 07:50 | 20.3 | 56.8
2021-04-20 07:55 | 20.3 | 56.7
2021-04-20 08:00 | 20.4 | 56.5
2021-04-20 08:05 | 20.4 | 56.6
2021-04-20 08:10 | 20.3 | 56.7
2021-04-20 08:15 | 20.4 | 56.7
2021-04-20 08:20 | 20.3 | 57
2021-04-20 08:25 | 20.3 | 57.1
2021-04-20 08:30 | 20.3 | 57.1
2021-04-20 08:35 | 20.3 | 57.2
2021-04-20 08:40 | 20.3 | 57.4
2021-04-20 08:45 | 20.3 | 57.7
2021-04-20 08:50 | 20.3 | 58
2021-04-20 08:55 | 20.2 | 58.2
2021-04-20 09:00 | 20.3 | 57.9
2021-04-20 09:05 | 20.2 | 57.9
2021-04-20 09:10 | 20.3 | 57.6
2021-04-20 09:15 | 20.2 | 57.8
2021-04-20 09:20 | 20.3 | 57.6
2021-04-20 09:25 | 20.3 | 57.5
2021-04-20 09:30 | 20.3 | 57.5
2021-04-20 09:35 | 20.3 | 57.4
2021-04-20 09:40 | 20.3 | 57.3
2021-04-20 09:45 | 20.3 | 57.4
2021-04-20 09:50 | 20.3 | 57.5
2021-04-20 09:55 | 20.3 | 57.3
2021-04-20 10:00 | 20.3 | 57.3
2021-04-20 10:05 | 20.3 | 57.3
2021-04-20 10:10 | 20.3 | 57.3
2021-04-20 10:15 | 20.3 | 57.3
2021-04-20 10:20 | 20.3 | 57.3
2021-04-20 10:25 | 20.3 | 57.3
2021-04-20 10:30 | 20.3 | 57.3
2021-04-20 10:35 | 20.3 | 57.3
2021-04-20 10:40 | 20.3 | 57.4
2021-04-20 10:45 | 20.3 | 57.3
2021-04-20 10:50 | 20.3 | 57.3
2021-04-20 10:55 | 20.3 | 57.4
2021-04-20 11:00 | 20.3 | 57.3
2021-04-20 11:05 | 20.3 | 57.3
2021-04-20 11:10 | 20.3 | 57.3
2021-04-20 11:15 | 20.4 | 57.1
2021-04-20 11:20 | 20.3 | 57.2
2021-04-20 11:25 | 20.3 | 57.1
2021-04-20 11:30 | 20.3 | 57
2021-04-20 11:35 | 20.4 | 57
2021-04-20 11:40 | 20.3 | 56.9
2021-04-20 11:45 | 20.3 | 56.9
2021-04-20 11:50 | 20.3 | 57.1
2021-04-20 11:55 | 20.3 | 57.2
2021-04-20 12:00 | 20.3 | 57
2021-04-20 12:05 | 20.3 | 56.9
2021-04-20 12:10 | 20.3 | 56.8
2021-04-20 12:15 | 20.3 | 56.7
2021-04-20 12:20 | 20.4 | 56.6
2021-04-20 12:25 | 20.3 | 56.6
2021-04-20 12:30 | 20.4 | 56.5
2021-04-20 12:35 | 20.3 | 56.5
2021-04-20 12:40 | 20.3 | 56.5
2021-04-20 12:45 | 20.4 | 56.4
2021-04-20 12:50 | 20.3 | 56.6
2021-04-20 12:55 | 20.3 | 56.6
2021-04-20 13:00 | 20.3 | 56.7
2021-04-20 13:05 | 20.3 | 56.7
2021-04-20 13:10 | 20.3 | 56.7
2021-04-20 13:15 | 20.3 | 56.6
2021-04-20 13:20 | 20.3 | 56.5
2021-04-20 13:25 | 20.3 | 56.4
2021-04-20 13:30 | 20.3 | 56.4
2021-04-20 13:35 | 20.3 | 56.3
2021-04-20 13:40 | 20.4 | 56.2
2021-04-20 13:45 | 20.3 | 56.2
2021-04-20 13:50 | 20.4 | 56.2
2021-04-20 13:55 | 20.4 | 56.1
2021-04-20 14:00 | 20.3 | 56.2
2021-04-20 14:05 | 20.4 | 55.9
2021-04-20 14:10 | 20.3 | 56
2021-04-20 14:15 | 20.3 | 55.9
2021-04-20 14:20 | 20.3 | 55.8
2021-04-20 14:25 | 20.3 | 55.7
2021-04-20 14:30 | 20.4 | 55.6
2021-04-20 14:35 | 20.4 | 55.5
2021-04-20 14:40 | 20.4 | 55.4
2021-04-20 14:45 | 20.4 | 55.3
2021-04-20 14:50 | 20.4 | 55.3
2021-04-20 14:55 | 20.4 | 55.3
2021-04-20 15:00 | 20.4 | 55.1
2021-04-20 15:05 | 20.4 | 55
2021-04-20 15:10 | 20.3 | 55.1
2021-04-20 15:15 | 20.4 | 55.1
2021-04-20 15:20 | 20.5 | 54.8
2021-04-20 15:25 | 20.5 | 54.8
2021-04-20 15:30 | 20.4 | 55
2021-04-20 15:35 | 20.3 | 55
2021-04-20 15:40 | 20.4 | 55
2021-04-20 15:45 | 20.4 | 55
2021-04-20 15:50 | 20.4 | 55.1
2021-04-20 15:55 | 20.4 | 55.1
2021-04-20 16:00 | 20.4 | 55.1
2021-04-20 16:05 | 20.4 | 55.3
2021-04-20 16:10 | 20.4 | 55.3
2021-04-20 16:15 | 20.4 | 55.3
2021-04-20 16:20 | 20.4 | 55.3
2021-04-20 16:25 | 20.4 | 55.4
2021-04-20 16:30 | 20.3 | 55.5
2021-04-20 16:35 | 20.3 | 55.6
2021-04-20 16:40 | 20.3 | 55.5
2021-04-20 16:45 | 20.3 | 55.7
2021-04-20 16:50 | 20.3 | 55.6
2021-04-20 16:55 | 20.3 | 55.5
2021-04-20 17:00 | 20.3 | 55.4
2021-04-20 17:05 | 20.4 | 55.4
2021-04-20 17:10 | 20.4 | 55.4
2021-04-20 17:15 | 20.3 | 55.4
2021-04-20 17:20 | 20.3 | 55.4
2021-04-20 17:25 | 20.4 | 55.4
2021-04-20 17:30 | 20.4 | 55.3
2021-04-20 17:35 | 20.4 | 55.1
2021-04-20 17:40 | 20.4 | 54.9
2021-04-20 17:45 | 20.4 | 54.7
2021-04-20 17:50 | 20.3 | 54.6
2021-04-20 17:55 | 20.4 | 54.5
2021-04-20 18:00 | 20.4 | 54.4
2021-04-20 18:05 | 20.4 | 54.6
2021-04-20 18:10 | 20.4 | 54.5
2021-04-20 18:15 | 20.4 | 54.6
2021-04-20 18:20 | 20.4 | 54.7
2021-04-20 18:25 | 20.4 | 54.9
2021-04-20 18:30 | 20.4 | 54.9
2021-04-20 18:35 | 20.4 | 55.1
2021-04-20 18:40 | 20.3 | 55.3
2021-04-20 18:45 | 20.4 | 55.3
2021-04-20 18:50 | 20.4 | 55.2
2021-04-20 18:55 | 20.4 | 55.2
2021-04-20 19:00 | 20.4 | 55.3
2021-04-20 19:05 | 20.4 | 55.5
2021-04-20 19:10 | 20.4 | 55.5
2021-04-20 19:15 | 20.3 | 55.6
2021-04-20 19:20 | 20.4 | 55.6
2021-04-20 19:25 | 20.4 | 55.6
2021-04-20 19:30 | 20.4 | 55.5
2021-04-20 19:35 | 20.3 | 55.7
2021-04-20 19:40 | 20.3 | 55.7
2021-04-20 19:45 | 20.4 | 55.7
2021-04-20 19:50 | 20.4 | 55.7
2021-04-20 19:55 | 20.3 | 55.9
2021-04-20 20:00 | 20.3 | 55.9
2021-04-20 20:05 | 20.4 | 55.9
2021-04-20 20:10 | 20.4 | 55.9
2021-04-20 20:15 | 20.4 | 55.9
2021-04-20 20:20 | 20.4 | 56
2021-04-20 20:25 | 20.4 | 56
2021-04-20 20:30 | 20.4 | 56.1
2021-04-20 20:35 | 20.4 | 56.1
2021-04-20 20:40 | 20.4 | 56.1
2021-04-20 20:45 | 20.4 | 56.2
2021-04-20 20:50 | 20.4 | 56.1
2021-04-20 20:55 | 20.4 | 56.3
2021-04-20 21:00 | 20.4 | 56.3
2021-04-20 21:05 | 20.4 | 56.3
2021-04-20 21:10 | 20.4 | 56.2
2021-04-20 21:15 | 20.4 | 56.2
2021-04-20 21:20 | 20.4 | 56.2
2021-04-20 21:25 | 20.4 | 56.1
2021-04-20 21:30 | 20.4 | 56.1
2021-04-20 21:35 | 20.4 | 56
2021-04-20 21:40 | 20.4 | 56
2021-04-20 21:45 | 20.4 | 56.1
2021-04-20 21:50 | 20.4 | 56.2
2021-04-20 21:55 | 20.4 | 56.2
2021-04-20 22:00 | 20.4 | 56.4
2021-04-20 22:05 | 20.4 | 56.4
2021-04-20 22:10 | 20.4 | 56.4
2021-04-20 22:15 | 20.4 | 56.5
2021-04-20 22:20 | 20.3 | 56.5
2021-04-20 22:25 | 20.3 | 56.5
2021-04-20 22:30 | 20.4 | 56.4
2021-04-20 22:35 | 20.4 | 56.3
2021-04-20 22:40 | 20.4 | 56.3
2021-04-20 22:45 | 20.4 | 56.3
2021-04-20 22:50 | 20.4 | 56.1
2021-04-20 22:55 | 20.4 | 56.1
2021-04-20 23:00 | 20.4 | 56.2
2021-04-20 23:05 | 20.4 | 56.2
2021-04-20 23:10 | 20.4 | 56.1
2021-04-20 23:15 | 20.4 | 56.2
2021-04-20 23:20 | 20.4 | 56.2
2021-04-20 23:25 | 20.4 | 56.2
2021-04-20 23:30 | 20.4 | 56.1
2021-04-20 23:35 | 20.4 | 56.1
2021-04-20 23:40 | 20.4 | 56.2
2021-04-20 23:45 | 20.4 | 56.2
2021-04-20 23:50 | 20.5 | 56.1
2021-04-20 23:55 | 20.4 | 56.2

## Conditions
| Temp |  | RH
All GMH5I readings | #REF! | #REF!
Room | #REF! | #REF! | #REF! | #REF!

## LHistory
History
40585
[12 empty] | as ratio | U in 2021
|  | S21034 | S21285 | S21470 | S21971 | S22091
| Nominal | 1995-05-09 | 1998-11-17 | 2002-09-04 | 2015-11-02 | 2021-04-30 | 1995-05-09 | 1998-11-17 | 2002-09-04 | 2015-11-02 | 2021-04-30
| 100 | 101.3351 | 101.3413 | 101.3353 | 101.3478 | 101.3457 | 1 | 1.0001 | 1 | 1.0001 | 1.0001 | 7.9092e-05 | 0.0080157
| 90 | 89.3654 | 89.37 | 89.3663 | 89.3870 | 89.378 | 1 | 1.0001 | 1 | 1.0002 | 1.0001 | 7.9124e-05 | 0.007072
| 80 | 79.1191 | 79.1228 | 79.1206 | 79.1421 | 79.131 | 1 | 1 | 1 | 1.0003 | 1.0002 | 7.9164e-05 | 0.0062643
| 70 | 69.3314 | 69.3341 | 69.3327 | 69.3526 | 69.3421 | 1 | 1 | 1 | 1.0003 | 1.0002 | 7.9218e-05 | 0.0054931
| 60 | 59.5257 | 59.5279 | 59.5269 | 59.5442 | 59.5351 | 1 | 1 | 1 | 1.0003 | 1.0002 | 7.9301e-05 | 0.0047212
| 50 | 49.429 | 49.4296 | 49.4294 | 49.4444 | 49.4361 | 1 | 1 | 1 | 1.0003 | 1.0001 | 7.9441e-05 | 0.0039272
| 40 | 39.6293 | 39.6292 | 39.6293 | 39.6415 | 39.6347 | 1 | 1 | 1 | 1.0003 | 1.0001 | 7.9689e-05 | 0.0031584
| 30 | 29.5243 | 29.5235 | 29.5239 | 29.5333 | 29.5279 | 1 | 0.99997 | 0.99998 | 1.0003 | 1.0001 | 8.024e-05 | 0.0023693
| 20 | 20.0807 | 20.08 | 20.0803 | 20.0867 | 20.083 | 1 | 0.99996 | 0.99998 | 1.0003 | 1.0001 | 8.166e-05 | 0.00164
| 10 | 10.167 | 10.1668 | 10.1671 | 10.1707 | 10.1684 | 1 | 0.99999 | 1 | 1.0004 | 1.0001 | 8.9108e-05 | 0.00090608
| 1000 | 980.805 | 981.2267 | 980.6299 | 980.752 | 980.8989 | 1 | 1.0004 | 0.99982 | 0.99995 | 1.0001 | 7.9092e-05 | 0.077582
| 900 | 876.945 | 877.2898 | 876.8131 | 876.980 | 877.0541 | 1 | 1.0004 | 0.99985 | 1 | 1.0001 | 7.9121e-05 | 0.069394
| 800 | 783.848 | 784.1268 | 783.7482 | 783.919 | 783.9567 | 1 | 1.0004 | 0.99987 | 1.0001 | 1.0001 | 7.9158e-05 | 0.062056
| 700 | 691.11 | 691.2989 | 691.0102 | 691.160 | 691.1924 | 1 | 1.0003 | 0.99986 | 1.0001 | 1.0001 | 7.9209e-05 | 0.054749
| 600 | 592.604 | 592.7309 | 592.5184 | 592.646 | 592.6548 | 1 | 1.0002 | 0.99986 | 1.0001 | 1.0001 | 7.9292e-05 | 0.046993
| 500 | 492.814 | 492.8877 | 492.7433 | 492.843 | 492.843 | 1 | 1.0001 | 0.99986 | 1.0001 | 1.0001 | 7.9429e-05 | 0.039146
| 400 | 395.793 | 395.8311 | 395.7387 | 395.818 | 395.8067 | 1 | 1.0001 | 0.99986 | 1.0001 | 1 | 7.9674e-05 | 0.031535
| 300 | 302.405 | 302.412 | 302.3617 | 302.419 | 302.4082 | 1 | 1 | 0.99986 | 1 | 1 | 8.016e-05 | 0.024241
| 200 | 204.59 | 204.5802 | 204.5619 | 204.598 | 204.5838 | 1 | 0.99995 | 0.99986 | 1 | 0.99997 | 8.1532e-05 | 0.01668
| 100 | 100.987 | 100.9724 | 100.9693 | 100.986 | 100.9781 | 1 | 0.99986 | 0.99982 | 0.99999 | 0.99991 | 8.9185e-05 | 0.0090057
Note on D16: Christian: Actually UB read 985.254