## pyUBreadings
Template for capacitors
Capacitors | GR-1409Y |  |  | Client | [5 empty] | Enter, say, 5 ppm as 5e-6
Date | 2021-12-07 |  |  | Job number | [5 empty] | 1 sigma uncertainties from specifications
|  |  |  | S Number | [5 empty] | specs say -140 ppm for capacitance, assume 100 ppm for G
|  |  | Data from lab book |  |  |  | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff C | u | df | Temp coeff G | u | df | Other | u | df
| Hz |  |  |  | Hz | °C
coax zero | 1591.5494 | 3Y | 0 | -2 | 1591.5494 | 20.61 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10
1 microF | 1591.5494 | 3Y | 999875 | 814 | 1591.5494 | 20.61 | 0.2 | 10 | 0 | 2e-05 | 10 | 0 | 0 | 10

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | td | utd | ktd | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for G | G | uG | k | multiplier for C | C | uC | k | nom Freq | td | utd | ktd
coax zero | 1591.5494 | 3Y | 0 | -2 | 1591.5494 | 20.61 | 0.2 | 10 | 0 | 0 | 10 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2566e-16 | 0 | 0 | #DIV/0! | coax zero1591.54943091895 | 1000 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5494 | 0 | 0 | 0
1 microF | 1591.5494 | 3Y | 999875 | 814 | 1591.5494 | 20.61 | 0.2 | 10 | 0 | 2e-05 | 10 | 0 | 0 | 10 | 9.9998e-07 | 4.7198e-11 | 2.0331 | 8.1597e-07 | 3.7824e-07 | 2.2184 | 0.081598 | 0.037825 | 2.2184 | 2.2e-11 | 2.7e-07 | 2.1454 | 1.4009 | 1 microF1591.54943091895 | 1000 | 0.00081597 | 0.00037824 | 2.2184 | 1000000 | 0.99998 | 4.7198e-05 | 2.0331 | 1591.5494 | 0.081598 | 0.037825 | 2.2184

## report
Table 1
|  | Nominal Value (μF) | Capacitance (μF) | Expanded Uncertainty (μF) | Coverage Factor | Loss Angle (mrad) | Expanded Uncertainty (mrad) | Coverage Factor
1 microF |  | 1 | 0.99998 | 4.7198e-05 | 2.0331 | 0.081598 | 0.037825 | 2.2184
dial correction |  | 108.4942 | ppm
cap expanded uncertainty |  | 47.199 | ppm | (largely driven by 20 ppm temperature coefficient)