## Read Me
resolve use/value of 'c' correction term (5 ppm issue)
resolved, derived on 'Resistors' page as mean of corrections from ESI boxes.
On 7 May 2009 am preparing for Fitzroy inductor calibration.
apply temperature correction to SR104 for completeness
use March 09 value for SR104, prior to BIPM correction
attempt full cal in August
completed in November

## Calibration
Calibration of Universal Bridge | [13 empty] | This block selects data from the docx Python output
A. Measurement of bridge resistors using dc. | [13 empty] | From docx
Date: | 2014-07-07 |  |  | Enter data into light blue areas | [9 empty] | Mon Jul 07 11:11:43 2014 | Sensor SR104 | 9968.8514 | 0.0016853
DMM: | 3458A S/N 2823A16452 (QHR) | [12 empty] | Mon Jul 07 11:22:27 2014 | SR104 | 9999.9736 | 0.001206
DMM settings: | NPLC 20, 4 wire ohms | [5 empty] | Data collected in 'July2014.docx' | [6 empty] | Mon Jul 07 11:25:05 2014 | SR104 | 9999.9735 | 0.001011
[6 empty] | GMH51 | using MeasureR.py to gather HP3458 readings over GPIB | [6 empty] | Mon Jul 07 13:20:44 2014 | R4A | 10000.5143 | 0.0015186
Observation No | Measurement | OCOMP ON? | DELAY | Mean of 20 | sd of 20 | temp | sd ppm | sensor R | Temp | R to add to 20 value |  |  |  | Mon Jul 07 13:46:06 2014 | Sensor SR104 | 9969.7553 | 0.0014158
1 | SR104 | on | 0 | 9999.9719 | 0.0010976 | 19.91 | 0.10976 | 9969.7553 | 19.9748 | -8.6478e-05 |  |  |  | Mon Jul 07 13:59:39 2014 | SR104 | 9999.9719 | 0.0010976
2 | 100k esi parallel | on | 0 | 10000.0391 | 0.0019682 | 20.01 | 0.19682 | [6 empty] | Mon Jul 07 14:45:30 2014 | 100k esi parallel | 10000.0391 | 0.0019682
3 | 100k esi series-para | on | 0 | 100000.1965 | 0.010205 | 20.07 | 0.10205 | [6 empty] | Mon Jul 07 14:57:05 2014 | 100k esi series-para | 100000.1965 | 0.010205
4 | 100k esi R10 | on | 0 | 99999.8548 | 0.01038 | 20.1 | 0.1038 | [6 empty] | Mon Jul 07 15:04:43 2014 | 100k esi R10 | 99999.8548 | 0.01038
5 | 100k esi series | on | 0 | 1000002.5445 | 0.23377 | 20.11 | 0.23377 | [6 empty] | Mon Jul 07 15:13:35 2014 | 100k esi series | 1000002.5445 | 0.23377
6 | G1 | on | 0 | 99998.2075 | 0.011203 | 19.65 | 0.11203 | [6 empty] | Mon Jul 07 15:53:39 2014 | R4B | 99998.5169 | 0.011883
7 | G2 | on | 0 | 99998.7528 | 0.0075329 | 19.65 | 0.07533 | [6 empty] | Mon Jul 07 16:00:55 2014 | G1 | 99998.2075 | 0.011203
8 | R4A | on | 0 | 10000.5143 | 0.0015186 | 19.55 | 0.15185 | [6 empty] | Mon Jul 07 16:04:27 2014 | R2 | 99998.7528 | 0.0075329
9 | R4B | on | 0 | 99998.5169 | 0.011883 | 19.64 | 0.11883 |  | SR104 cal val | 9999.9989 | at 20 |  |  | Mon Jul 07 16:12:22 2014 | R4C | 1000063.1377 | 0.30782
10 | R4C | on | 0 | 1000063.1377 | 0.30782 | 19.58 | 0.3078
Analysis
Resistor | Nominal Value (ohms) | ppm of nominal resistance | ppm of nominal conductance | mean DMM reading (ohms) | stdev (ohms) | DMM correction (ohms) | Value of Resistor (ohms) | Traceability - last cal date | Reference |  |  | DMM corrn ppm | stdev ppm
SR104 | 10000 |  |  | 9999.9719 | 0.0010976 | 0.026946 | 9999.9988 | 2014-01-23 | ...\Electricity\Ongoing\OHM\Buildup2014.xls, |  |  | 2.6946 | 0.10976 | sheet:ResultsFromT16_T81; cell AO38
100k esi parallel | 10000 | 6.6 |  | 10000.0391 | 0.0019682 |  | 10000.066 | [5 empty] | 0.19682
100k esi series-para | 100000 | 6.9417 |  | 100000.1965 | 0.010205 | 0.49763 | 100000.6942 | Note: see ESI model SR1010 handbook for series/parallel formula |  |  |  | 4.9763 | 0.10205
100k esi R10 | 100000 |  |  | 99999.8548 | 0.01038 | [7 empty] | 0.1038
100k esi series | 1000000 | 6.6 |  | 1000002.5445 | 0.23377 | 4.0556 | 1000006.6 |  |  |  |  | 4.0556 | 0.23377
G1 | 100000 |  | 12.9485 | 99998.2075 | 0.011203 |  | 99998.7051 | [5 empty] | 0.11203
G2 | 100000 |  | 7.4961 | 99998.7528 | 0.0075329 |  | 99999.2504 | [5 empty] | 0.07533
R4A | 10000 | 54.1261 |  | 10000.5143 | 0.0015186 |  | 10000.5413 | [5 empty] | 0.15185
R4B | 100000 | -9.8547 |  | 99998.5169 | 0.011883 |  | 99999.0145 | [5 empty] | 0.11883 | 1.08e-07
R4C | 1000000 | 67.1933 |  | 1000063.1377 | 0.30782 |  | 1000067.1933 | [5 empty] | 0.3078
B. Measurement of 1000 pF GR capacitor S/N 2248
|  |  |  | PRT | UB1000
Traceability
|  |  |  | Last cal date | Reference
value of GR1000A |  | 27.82 | -6.19 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
value of GR1000B |  | 14.18 | -6.76 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
Measurement with AH2700A
Date: | 2014-07-08
|  |  |  | PRT | UB1000
Observation No. | Measurement | C pF | G nS | data logger | temperature (deg C)
1 | GR1000 A | 1000.0296 | 0.12809 | 23.99
2 | GR 1000B | 1000.016 | 0.13532 | 23.99
3 | UB 1000 | 1000.1449 | 0.12131 |  | 19.7 | deg C
Meter Correction
ppm
GR1000 A |  |  | -1.78
GR 1000B |  |  | -1.82
UB 1000 |  | 1000.1431 | -1.8
Summary of results for pasting into AudioRegC
C_1 | R_4A | R_4B | R_4C | G_1 | G_2
143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961

## Python notation
Python variable names
| measured mean |  | temperature
| true_sr104_20 | u_true_sr104_20 |  | traceable value of SR104
sensor R | sens_sr104 | u_sens_sr104 |  | mean of HP3458 measured values with standard deviation and temperature
SR104 | sr104 | u_sr104 | t_sr104
100k esi parallel | ep100k | u_ep100k | t_ep100k
100k esi series-para | esp100k | u_esp100k | t_esp100k
100k esi R10 | er10_100k | u_er10_100k | t_er10_100k
100k esi series | es100k | u_es100k | t_es100k
G1 | g1 | u_g1 | t_g1 | calculated values of the internal standards from the above
G2 | g2 | u_g2 | t_g2
R4A | r4a | u_r4a | t_r4a
R4B | r4b | u_r4b | t_r4b
R4C | r4c | u_r4c | t_r4c

## Short, Open, Phase
Calibration of Universal Bridge
Shorts and Opens
Date: | 2014-07-24
Frequency | 1592
| Shorts |  |  |  | error, ppm | error, ppm
O bservation No. | Range | L | R |  | L | R
86 | 7Z | -1e-06 | -7e-07 | coaxial | -1 | -0.7
50 | 6Z | -1e-06 | -1e-06 | coaxial | -1 | -1
46 | 5Z | 0 | -2.2e-06 | coaxial | 0 | -2.2
70 | 4Z | -1e-06 | -1.3e-06 | coaxial | -1 | -1.3
74 | 3Z | -1e-06 | -8e-07 | coaxial | -1 | -0.8
17 | 2Z | 2e-06 | -1e-06 | flying leads | 2 | -1 | flying lead zeros used with ESI
14 | 1Z | 9e-06 | 2.3e-06 | flying leads | 9 | 2.3 | boxes
O bservation No. | Opens | C | G |  | C | G
82 | 7Y | 6e-06 | 1.7e-06 | coaxial | 6 | 1.7
54 | 6Y | 0 | 1.3e-06 | coaxial | 0 | 1.3
42 | 5Y | 1e-06 | 2.4e-06 | coaxial | 1 | 2.4
66 | 4Y | 0 | -2e-07 | coaxial | 0 | -0.2
78 | 3Y | 0 | -1e-07 | coaxial | 0 | -0.1
22 | 2Y | 0 | -1e-07 | flying leads | 0 | -0.1
10 | 1Y | 0 | -1e-07 | flying leads | 0 | -0.1
Phase angle standards
O bservation No. | Range | Frequency | Standard | value (pF) | C | G |  | Range | L | R
57/58 | 6Y | 1592 | 100k No 4 | 0.014 | 0 | 1.0003 |  | 6Z | -2.3e-05 | 0.99967
59/60 | 6Y | 160 | 100k No 4 | 0.014 | 1.9e-05 | 1.0003 |  | 6Z | -2.7e-05 | 0.99968
61/62 | 5Y | 1592 | 10k No 4 | 0.028 | -9e-06 | 0.99938 |  | 5Z | 5e-06 | 1.0006
63/64 | 5Y | 160 | 10k No 4 | 0.028 | -8e-06 | 0.99938 |  | 5Z | 4e-06 | 1.0006
65/69 | 4Y | 1592 | 1k No 4 | 0.01 | -9e-06 | 1.0024 |  | 4Z | 1e-05 | 0.99759
67/71 | 4Y | 160 | 1k No 4 | 0.01 | -1.7e-05 | 1.0024 |  | 4Z | 1.3e-05 | 0.99767
77/73 | 3Y | 1592 | 100 ohm No 2 | 0 | 0 | 0.99669 |  | 3Z | -1e-06 | 1.0033
79/75 | 3Y | 160 | 100 ohm No 2 | 0 | -6e-06 | 0.99667 |  | 3Z | 4e-06 | 1.0033
|  | Phase error, pF |  |  | Phase error ppm of fsd |  | phase error
Range | Full scale pF | Y ranges | Z ranges | Frequency | Y ranges | Z ranges | of ratio G1/G2
6 | 1000 | -0.014 | -0.009 | 1592 | -14 | -9 | -11.5
6 | 1000 | 0.005 | -0.013 | 160 | 5 | -13 | -4
5 | 10000 | -0.118 | 0.078 | 1592 | -11.8 | 7.8 | -2
5 | 10000 | -0.108 | 0.068 | 160 | -10.8 | 6.8 | -2
4 | 100000 | -0.91 | 1.01 | 1592 | -9.1 | 10.1 | 0.5
4 | 100000 | -1.71 | 1.31 | 160 | -17.1 | 13.1 | -2
3 | 1000000 | 0 | -1 | 1592 | 0 | -1 | -0.5
3 | 1000000 | -6 | 4 | 160 | -6 | 4 | -1
|  |  |  | mean | -7.975 | 2.35 | -2.8125
|  |  |  | stdev | 7.361 | 9.2853 | 3.7506
2014 comment
Comparable to 2009 result, retain 16 ppm as conservative estimate
2009 comment
Phase results similar to 2005, retain 16 ppm as conservative estimate
Comment
Previously, the phase uncertainty was taken as 16 ppm. This is now clearly
too pessimistic for the Z ranges. Assuming no corrections are made, a 1-sigma
uncertainty for the Z ranges is conservatively 5 ppm.
There is an unexplained mean error of -10 ppm in the Y ranges. If no corrections
are made, 16 ppm remains a conservative 1-sigma uncertainty for the Y ranges.

## Resistors
Calibration of Universal Bridge
Calibration of check resistors using DC
Date: | 2009-08-14 |  |  | Enter data into light blue areas
DMM: |  |  |  | DC resistance data linked to docx page
DMM settings: | NPLC 20, 4 wire ohms | [5 empty] | Correction | -0.03
GMH51 uncorrected readings
Observation No | Measurement | OCOMP ON? | DELAY | Mean | sd | sd ppm | temp
| SR104 | on |  | 9999.9676 | 0.001136 | 0.1136 | 20.07
| 1k ESI, series | on |  | 9999.645 | 0.00084579 | 0.084582 | 20.05
| 1k ESI S/P | on |  | 999.963 | 0.00015074 | 0.15075
| 1k ESI, R10 | on |  | 999.9736 | 0.0001751 | 0.17511
| 1k ESI, R1 | on |  | 999.9638 | 0.0001199 | 0.11991 | 20.17
| 100 ohm ESI, series | on |  | 999.9594 | 9.9641e-05 | 0.099645 | 19.98
| 1k ESI, parallel | on |  | 99.9964 | 0.00010128 | 1.0128
| 100 ohm ESI, R1 | on |  | 99.9971 | 9.3599e-05 | 0.93602 | 19.8
| 100 ohm ESI, parallel | on |  | 9.9996 | 1.2468e-05 | 1.2468 | 19.89
| 10 ohm ESI, R1 | on |  | 9.9999 | 1.05e-05 | 1.05 | 20.06
| 10 ohm ESI, series | on |  | 99.999 | 0.00010012 | 1.0012 | 20.14
| 100k Vishay | on |  | 100001.5383 | 0.0092966 | 0.092965
| 10 k Vishay | on |  | 10000.276 | 0.0015076 | 0.15075 | 20.17
| 100 ohm Vishay | on |  | 100.0132 | 0.00011932 | 1.1931 | 20.11 | [6 empty] | sensor R | Temp | R to add to 20 value
[14 empty] | 9971.4967 | 20.1489 | 0.00050326
Analysis
Resistor | Nominal Value (ohms) | ppm of nominal resistance | ppm of nominal conductance | mean DMM reading (ohms) | stdev (ohms) | DMM correction (ohms) | Value of Resistor (ohms) | Traceability - last cal date | Reference |  |  |  | DMM Corrn ppm
SR104 | 10000 |  |  | 9999.9676 | 0.001136 | 0.031802 | 9999.9994 | 2014-01-23 | ...\Electricity\Ongoing\OHM\Buildup2014.xls, |  |  |  | 3.1802
| 100000 |  |  |  |  | 0.49763 | DMM correction at 100k from "Calibration" worksheet | [5 empty] | 4.9763 | 2005 values
1k ESI, series | 10000 | -32.3191 |  | 9999.645 | 0.00084579 |  | 9999.6768 | [6 empty] | 9999.7443 | -6.7513
1k ESI S/P | 1000 | -33.3761 |  | 999.963 | 0.00015074 | 0.003621 | 999.9666 | Note: see ESI model SR1010 handbook for series/parallel formula |  |  |  |  | 3.6211 | 999.9745 | -7.8957
1k ESI, R10 | 1000 |  |  | 999.9736 | 0.0001751
1k ESI, R1 | 1000 | -32.572 |  | 999.9638 | 0.0001199 |  | 999.9674 | [6 empty] | 999.9761 | -8.6716
100 ohm ESI, series | 1000 | -37.0145 |  | 999.9594 | 9.9641e-05 |  | 999.963 | [6 empty] | 999.9723 | -9.3141
1k ESI, parallel | 100 |  |  | 99.9964 | 0.00010128 | 0.00041456 | 99.9968 | [5 empty] | 4.1458 | 99.9974 | -6.7513
100 ohm ESI, R1 | 100 | -25.0692 |  | 99.9971 | 9.3599e-05 |  | 99.9975 | [6 empty] | 99.9975 | -0.28124
100 ohm ESI, parallel | 10 | -37.0145 |  | 9.9996 | 1.2468e-05 | 7.1285e-05 | 9.9996 | [5 empty] | 7.1288 | 9.9997 | -9.3141
10 ohm ESI, R1 | 10 | 1.8034 |  | 9.9999 | 1.05e-05 |  | 10 | [6 empty] | 10.0001 | -5.4959
10 ohm ESI, series | 100 |  |  | 99.999 | 0.00010012 |  | 99.9995 | [6 empty] | 99.9993 | 2.0007
10 ohm ESI, para | 1 | [5 empty] | 0.99999 | [6 empty] | 0.99999 | 2.0007
100k Vishay | 100000 |  |  | 100001.5383 | 0.0092966 |  | 100002.0359 | [6 empty] | 100002.938 | -9.0204
10 k Vishay | 10000 |  |  | 10000.276 | 0.0015076 |  | 10000.3078 | [6 empty] | 10000.3196 | -1.1812
100 ohm Vishay | 100 |  |  | 100.0132 | 0.00011932 |  | 100.0136 | [6 empty] | 100.012 | 15.6151
New build up to 1M Vishay on 31 July 2014
SR104 | 10000 |  |  | 9999.9698 | 0.0012382 | 0.028869 | 9999.9986 | 2014-01-23 | ...\Electricity\Ongoing\OHM\Buildup2014.xls,
100k esi parallel | 10000 |  |  | 10000.0534 | 0.00137 |  | 10000.0823
100k esi series | 1000000 |  |  | 1000004.8675 | 0.14039 | 3.3582 | 1000008.2258
1M Vishay | 1000000 |  |  | 999997.087 | 0.11704 |  | 1000000.4453 | [6 empty] | sensor R | Temp | R to add to 20 value
[14 empty] | 9969.188 | 19.918 | -0.00028216
estimate of F
Date: | 2014-07-22 | [6 empty] | -0.035 | [11 empty] | G1/G2
Observation No | Range | Nominal value ohms | Resistor | Frequency | C | G | Range | L | R | temp | DC resistance | Corrected UB G | Corrected UB R | Gac -Gdc ppm | Rac - Rdc ppm |  | Ratio error of G1/G2 | phase error of G1/G2 |  | 5.4524
81/85 | 7Y | 1000000 | Vishay | 1592 | 0.000356 | 1.0001 | 7Z | -0.00035 | 0.99992 | 20.08 | 1 | 1 | 0.99999 | 0.50445 | -10.6996 | 0.99999 | -5.0976 | 3 | -10.5526 | -5.1002
83/87 | 7Y | 1000000 | Vishay | 160 | 0.000264 | 1 | 7Z | -0.000259 | 0.99992 | 20.45 | 1 | 0.99998 | 1 | -14.9956 | -2.7996 | 0.99999 | -8.8976 | 2.5 | -14.3518 | -8.8994
53/49 | 6Y | 100000 | Vishay | 1592 | 0.000625 | 0.99998 | 6Z | -0.000625 | 1 | 20.08 | 1 | 0.99999 | 0.99999 | 12.666 | -25.6616 | 0.99999 | -6.4976 | 0 | -11.9501 | -6.4976
55/51 | 6Y | 100000 | Vishay | 160 | 0.000641 | 0.99997 | 6Z | -0.000646 | 1 | 19.96 | 1 | 0.99999 | 1 | 7.266 | -22.9716 | 0.99999 | -7.8526 | -2.5 | -13.3051 | -7.8527
41/45 | 5Y | 10000 | Vishay | 1592 | 5.4e-05 | 1 | 5Z | -5.5e-05 | 0.99996 | 20.16 | 1 | 0.99997 | 1 | 0.20422 | -14.3003 | 0.99999 | -7.0478 | -0.5 | -12.5005 | -7.048
43/47 | 5Y | 10000 | Vishay | 160 | 5.5e-05 | 1 | 5Z | -5.6e-05 | 0.99996 | 20.04 | 1 | 0.99998 | 1 | 5.9042 | -6.8003 | 0.99999 | -0.44785 | -0.5 | -5.9004 | -0.448
37/33 | 4Y | 1000 | ESI 1000 R1 | 1592 | -0.000197 | 1.0001 | 4Z | 0.000191 | 0.99989 | 19.8 | 0.99997 | 1 | 0.99995 | 11.0533 | -13.9495 | 0.99999 | -1.4486 | -3 | -6.9049 | -1.4525
39/35 | 4Y | 1000 | ESI 1000 R1 | 160 | -0.0002 | 1.0001 | 4Z | 0.000201 | 0.9999 | 19.78 | 0.99997 | 1 | 0.99996 | 10.0533 | -4.5495 | 1 | 2.7517 | 0.5 | -2.7044 | 2.748
25/29 | 3Y | 100 | ESI 100 R1 | 1592 | -0.000532 | 1.0001 | 3Z | 0.000532 | 0.9999 | 19.92 | 0.99997 | 1 | 0.99996 | 13.1565 | -17.8522 | 0.99999 | -2.3484 | 0 | -7.8045 | -2.352
27/31 | 3Y | 100 | ESI 100 R1 | 160 | -0.00054 | 1.0001 | 3Z | 0.000538 | 0.99991 | 19.76 | 0.99997 | 1 | 0.99997 | 10.2565 | -4.6522 | 1 | 2.8019 | -1 | -2.6538 | 2.7987
21/17 | 2Y | 10 | ESI 10 R1 | 1592 | -0.000358 | 1.0001 | 2Z | 0.000356 | 0.99993 | 20.13 | 1 | 1 | 0.99999 | 8.9297 | -13.9248 | 0.99999 | -2.4976 | -1 | -7.952 | -2.4996
23/19 | 2Y | 10 | ESI 10 R1 | 160 | -0.00037 | 1.0001 | 2Z | 0.000369 | 0.99994 |  | 1 | 1 | 1 | 7.5297 | -3.0248 | 1 | 2.2524 | -0.5 | -3.2017 | 2.2508
9/13 | 1Y | 1 | ESI 10 para | 1592 | -0.000355 | 1.0001 | 1Z | 0.000359 | 0.99992 | 20.02 | 0.99999 | 1 | 0.99998 | 13.9391 | -12.7342 | 1 | 0.60227 | 2 | -4.8527 | 0.59979
11/15 | 1Y | 1 | ESI 10 para | 160 | -0.00036 | 1.0001 | 1Z | 0.00037 | 0.99993 | 20.21 | 0.99999 | 1 | 0.99999 | 20.9391 | -1.0342 | 1 | 9.9523 | 5 | 4.4975 | 9.95 | 1.6 kHz | 160 Hz
[14 empty] | 8.6362 | -15.5889 |  |  |  | -7.1526 | -1.7002 | -3.7417 | 0.078201
|  | 10.97 | [10 empty] | average: | 10.1581 | -11.7879 | average: | -0.81498 | -0.125
[13 empty] | std dev: | 5.0505 | 7.9281 | std dev: | 5.0411 | 2.0684
| test_c | 0
[14 empty] | 10.973 | average +/- error
Using test_c |  | correction to | correction to | [10 empty] | 6.7076 | -6.5475
|  | L/C dial | R/G dial
Range | [12 empty] | Comment, 2014 calibration
|  | (ppm) | (ppm) | [9 empty] | The 'c' term has increased significantly(from 2 to 11 ppm) but the standard deviation in the ac-dc differences has decreased.
1Z |  | 189.73 | 59.5785 | [9 empty] | c' was measured as 5.7 ppm in 2005, so it is not clear any repair is required, set to 10.97 ppm.
2Z |  | 189.73 | 59.5785 | [9 empty] | This suggests the calibration method is performing well with more careful use of the HP3458 (eg measuring frequency each time for NPLC)
3Z |  | 189.73 | 59.5785 | [9 empty] | The increase in 'c' is a concern
4Z |  | 189.73 | 59.5785 | [9 empty] | G1/G2 ratio scatter has increased.
5Z |  | 189.73 | 59.5785 | [9 empty] | Scatter in Gac-Gdc, and Rac-Rdc, is 8 ppm, down from 17.5 ppm.
6Z |  | 125.7492 | -4.4023 | [9 empty] | Cautiously leave the uncertainty at 17 ppm as in 2009.
7Z |  | 202.7972 | 72.6457
1Y |  | 81.4779 | -48.6736 | [9 empty] | Comment, 2009 calibration
2Y |  | 81.4779 | -48.6736 | [9 empty] | The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
3Y |  | 81.4779 | -48.6736 | [9 empty] | The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
4Y |  | 81.4779 | -48.6736 | [9 empty] | However, the scatter in Gac-Gdc, and Rac-Rdc, is 17.5 ppm with an average error of 2 ppm in G, or -2 in R. Note the 2005 version has an error in the 100 k dc value
5Y |  | 81.4779 | -48.6736 | [9 empty] | The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
6Y |  | 145.4587 | 15.3072 | [9 empty] | errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
7Y |  | 68.4107 | -61.7408 | [9 empty] | of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 2.0 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 17.5
Hence, increase "amplifier ratio" in the uncertainty budget to 17 ppm.
Comment, 2005 calibration
The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
However, the scatter in Gac-Gdc, and Rac-Rdc, is 13.8 ppm with an average error of 5.7 ppm in G, or -5.7 in R.
The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 5.7 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 13.9589
Hence, increase "amplifier ratio" in the uncertainty budget to 14 ppm.
Comment, 2004 calibration
The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
However, the scatter in Gac-Gdc, and Rac-Rdc, is 6.3 ppm with an average error of 5.2 ppm in G, or -5.2 in R.
The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 5.2 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 6.2642
Hence, increase "amplifier ratio" in the uncertainty budget to 7 ppm.
Set 'test_c' to give a zero average +/- error.
Note on A8: Keith Jones: Linked to docx page, no longer a notebook entry.
Note on Q53: Keith Jones: Range 7 is a new addition to the calibration process. Values not used for setting c.

## Results
DC calibration of resistors
Name | Value | Uncertainty |  |  |  |  | ppm from nominal | nominal
| ohm | ohm
true_sr104_t | 9999.9988 | 0.001 |  |  |  |  | -0.11665 | 10000
true_ep_100k | 10000.066 | 0.0024655 |  |  |  |  | 6.6 | 10000
true_esp_100k | 100000.6942 | 0.024697 |  |  |  |  | 6.9417 | 100000
true_es_100k | 1000006.6 | 0.24655 |  |  |  |  | 6.6 | 1000000
true_g1 | 99998.7051 | 0.028614 |  |  |  |  | -12.9485 | 100000
true_g2 | 99999.2504 | 0.027387 |  |  |  |  | -7.4961 | 100000
true_r4a | 10000.5413 | 0.0021239 |  |  |  |  | 54.1261 | 10000
true_r4b | 99999.0145 | 0.028887 |  |  |  |  | -9.8547 | 100000
true_r4c | 1000067.1933 | 0.45846 |  |  |  |  | 67.1933 | 1000000
true_sr104b_t | 9999.9994 | 0.001 |  |  |  |  | -0.057674 | 10000
true_es_1k | 9999.6768 | 0.0017338 |  |  |  |  | -32.3191 | 10000
true_esp_1k | 999.9666 | 0.00017491 |  |  |  |  | -33.3761 | 1000
true_er1_1k | 999.9674 | 0.00025129 |  |  |  |  | -32.572 | 1000
true_es_100 | 999.963 | 0.00024228 |  |  |  |  | -37.0145 | 1000
true_ep_1k | 99.9968 | 1.7338e-05 |  |  |  |  | -32.3191 | 100
true_ep_100 | 9.9996 | 2.4228e-06 |  |  |  |  | -37.0145 | 10
true_er1_100 | 99.9975 | 0.00013899 |  |  |  |  | -25.0692 | 100
true_er1_10 | 10 | 1.648e-05 |  |  |  |  | 1.8034 | 10
true_es_10 | 99.9995 | 0.00014346 |  |  |  |  | -5.3873 | 100
true_v_100k | 100002.0359 | 0.027923 |  |  |  |  | 20.3593 | 100000
true_v_10k | 10000.3078 | 0.0021362 |  |  |  |  | 30.7788 | 10000
true_v_100 | 100.0136 | 0.00015747 |  |  |  |  | 135.829 | 100
true_sr104c_t | 9999.9986 | 0.001 |  |  |  |  | -0.13622 | 10000
true_epc_100k | 10000.0823 | 0.0021 |  |  |  |  | 8.2258 | 10000
true_esc_100k | 1000008.2258 | 0.21 |  |  |  |  | 8.2258 | 1000000
true_v_1M | 1000000.4453 | 0.2784 |  |  |  |  | 0.44528 | 1000000
Phase angle corrections from Thompson set
Label | Range | Freq/Hz | F | H | ppm C | ppm L
thomp_6a | 6 | 1592 | 1.4e-14 | 8.0184e-05 | 14 | 8.0184
thomp_6b | 6 | 160 | -5e-15 | 0.00013018 | -5 | 13.0181
thomp_5a | 5 | 1592 | 1.28e-13 | -7.807e-06 | 12.8 | -7.807
thomp_5b | 5 | 160 | 1.18e-13 | -6.8069e-06 | 11.8 | -6.8069
thomp_4a | 4 | 1592 | 9.1e-13 | -1.1099e-06 | 9.1 | -11.099
thomp_4b | 4 | 160 | 1.71e-12 | -1.0099e-06 | 17.1 | -10.099
thomp_3a | 3 | 1592 | 0 | 0 | 0 | 0
thomp_3b | 3 | 160 | 6e-12 | -4e-08 | 6 | -4
Gain Factor
Range | prod / ppm | factor / ppm | Freq / Hz
7 | -10.5526 | -5.1001 | 1592
7 | -14.3518 | -8.8994 | 160
6 | -11.9501 | -6.4976 | 1592
6 | -13.3051 | -7.8526 | 160
5 | -12.5005 | -7.048 | 1592
5 | -5.9004 | -0.44796 | 160
4 | -6.9049 | -1.4524 | 1592
4 | -2.7044 | 2.7481 | 160
3 | -7.8045 | -2.352 | 1592
3 | -2.6538 | 2.7987 | 160
2 | -7.952 | -2.4995 | 1592
2 | -3.2017 | 2.2508 | 160
1 | -4.8527 | 0.59983 | 1592
1 | 4.4975 | 9.95 | 160
UB measured values
Label | G / siemen | R / ohm |  |  |  |  | 1/G | R | GxR | DC R | rel to 1/G | rel to R
v_1M_1592 | 1e-06 | 999989.7343 |  |  |  |  | 999999.9397 | 999989.7343 | 0.99999 | 1000000.4453 | -0.50557 | -10.711
v_1M_160 | 9.9998e-07 | 999997.6349 |  |  |  |  | 1000015.439 | 999997.6349 | 0.99998 | 1000000.4453 | 14.9937 | -2.8104
v_100k_1592 | 9.9999e-06 | 99999.47 |  |  |  |  | 100000.77 | 99999.47 | 0.99999 | 100002.0359 | -12.6587 | -25.6589
v_100k_160 | 9.9999e-06 | 99999.739 |  |  |  |  | 100001.31 | 99999.739 | 0.99998 | 100002.0359 | -7.2587 | -22.9689
v_10k_1592 | 9.9997e-05 | 10000.1648 |  |  |  |  | 10000.3058 | 10000.1648 | 0.99999 | 10000.3078 | -0.19957 | -14.3007
v_10k_160 | 9.9998e-05 | 10000.2398 |  |  |  |  | 10000.2488 | 10000.2398 | 1 | 10000.3078 | -5.8994 | -6.8005
er1_1k_1592 | 0.001 | 999.9535 |  |  |  |  | 999.9564 | 999.9535 | 1 | 999.9674 | -11.0446 | -13.9545
er1_1k_160 | 0.001 | 999.9629 |  |  |  |  | 999.9574 | 999.9629 | 1 | 999.9674 | -10.0447 | -4.5536
er1_100_1592 | 0.01 | 99.9957 |  |  |  |  | 99.9962 | 99.9957 | 1 | 99.9975 | -13.148 | -17.8571
er1_100_160 | 0.01 | 99.997 |  |  |  |  | 99.9965 | 99.997 | 1 | 99.9975 | -10.2483 | -4.6559
er1_10_1592 | 0.1 | 9.9999 |  |  |  |  | 9.9999 | 9.9999 | 0.99999 | 10 | -8.9232 | -13.9273
er1_10_160 | 0.1 | 10 |  |  |  |  | 9.9999 | 10 | 1 | 10 | -7.5233 | -3.0267
ep_10_1592 | 1 | 0.99998 |  |  |  |  | 0.99998 | 0.99998 | 1 | 0.99999 | -13.9317 | -12.7372
ep_10_160 | 1 | 0.99999 |  |  |  |  | 0.99997 | 0.99999 | 1 | 0.99999 | -20.9311 | -1.0364
Inductor set
Label | L / henry | R / ohm |  |  |  |  | L / microhenry
gr100uh_1k | 9.9901e-05 | 0.081306 |  |  |  |  | 99.9013
gr1000uh_1k | 0.00099894 | 0.86068 |  |  |  |  | 998.9426
gr10000uh_1k | 0.0099994 | 8.6456 |  |  |  |  | 9999.3776
gr100000uh_1k | 0.099979 | 86.4954 |  |  |  |  | 99979.3731
gr100000uh_100 | 1.0022 | 864.9885 |  |  |  |  | 1002239.1952
Capacitor set
Label | C / farad | G / siemen |  |  |  |  | C / pF
su100000_1k | 1.0002e-07 | 4.3298e-08 |  |  |  |  | 100016.9542
su1000000_1k | 1.0002e-06 | 5.7497e-07 |  |  |  |  | 1000176.5423
gr1000a_1592 | 1e-09 | -1.9e-10 |  |  |  |  | 1000.0435
gr1000b_1592 | 1e-09 | -2.1e-10 |  |  |  |  | 1000.0295
gr100_1592 | 1e-10 | 1.4399e-10 |  |  |  |  | 100.0015
gr_10x100_1k | 9.9944e-07 | 5.9997e-07 |  |  |  |  | 999444.4826
gr_10x10n_1k | 1.0021e-07 | 3.7598e-08 |  |  |  |  | 100210.57
gr_10x1n_1k | 1.0043e-08 | 3.8898e-09 |  |  |  |  | 10043.2588
gr_0_1k | 4.7954e-11 | 5.6997e-10 |  |  |  |  | 47.9539

## csv_data
Resistor cal set a | Value | Stdev | dof | Comment
true_sr104_20 | 9999.9989 | 0.001 | 50 | cal_cert interpretation
sens_sr104 | 9969.7553 | 0.0014158 | 19
sr104 | 9999.9719 | 0.0010976 | 19
ep_100k | 10000.0391 | 0.0019682 | 19
esp_100k | 100000.1965 | 0.010205 | 19
er10_100k | 99999.8548 | 0.01038 | 19
es_100k | 1000002.5445 | 0.23377 | 19
g1 | 99998.2075 | 0.011203 | 19
g2 | 99998.7528 | 0.0075329 | 19
r4a | 10000.5143 | 0.0015186 | 19
r4b | 99998.5169 | 0.011883 | 19
r4c | 1000063.1377 | 0.30782 | 19
Zeros | dial a | dial b
zero7z_1592 | -1e-06 | -7e-07
zero6z_1592 | -1e-06 | -1e-06
zero5z_1592 | 0 | -2.2e-06
zero4z_1592 | -1e-06 | -1.3e-06
zero3z_1592 | -1e-06 | -8e-07
zero2z_1592 | 2e-06 | -1e-06
zero1z_1592 | 9e-06 | 2.3e-06
zero7y_1592 | 6e-06 | 1.7e-06
zero6y_1592 | 0 | 1.3e-06
zero5y_1592 | 1e-06 | 2.4e-06
zero4y_1592 | 0 | -2e-07
zero3y_1592 | 0 | -1e-07
zero2y_1592 | 0 | -1e-07 |  | flying leads
zero1y_1592 | 0 | -1e-07 |  | flying leads
zero7z_160 | 0 | 0 |  | not measured at 160 Hz!
zero6z_160 | 0 | 0
zero5z_160 | 0 | 0
zero4z_160 | 0 | 0
zero3z_160 | 0 | 0
zero2z_160 | 0 | 0
zero1z_160 | 0 | 0
zero7y_160 | 0 | 0
zero6y_160 | 0 | 0
zero5y_160 | 0 | 0
zero4y_160 | 0 | 0
zero3y_160 | 0 | 0
zero2y_160 | 0 | 0
zero1y_160 | 0 | 0
Thompson ac Cal | dial a | dial b | cap pF
thomp_100k_4_z_1592 | -2.3e-05 | 0.99967 | 0.014
thomp_100k_4_y_1592 | 0 | 1.0003 | 0.014
thomp_10k_4_z_1592 | 5e-06 | 1.0006 | 0.028
thomp_10k_4_y_1592 | -9e-06 | 0.99938 | 0.028
thomp_1k_4_z_1592 | 1e-05 | 0.99759 | 0.01
thomp_1k_4_y_1592 | -9e-06 | 1.0024 | 0.01
thomp_100_2_z_1592 | -1e-06 | 1.0033 | 0
thomp_100_2_y_1592 | 0 | 0.99669 | 0
thomp_100k_4_z_160 | -2.7e-05 | 0.99968 | 0.014
thomp_100k_4_y_160 | 1.9e-05 | 1.0003 | 0.014
thomp_10k_4_z_160 | 4e-06 | 1.0006 | 0.028
thomp_10k_4_y_160 | -9e-06 | 0.99938 | 0.028
thomp_1k_4_z_160 | 1e-05 | 0.99759 | 0.01
thomp_1k_4_y_160 | -1.7e-05 | 1.0024 | 0.01
thomp_100_2_z_160 | 4e-06 | 1.0033 | 0
thomp_100_2_y_160 | -6e-06 | 0.99667 | 0
Resistor cal set b | Value | Stdev | dof | Comment
sens_sr104b | 9971.4967 | 0.0017802 | 19 | assume true value of sr104 as in set1, just different temperature
sr104b | 9999.9676 | 0.001136 | 19
es_1k | 9999.645 | 0.00084579 | 19
esp_1k | 999.963 | 0.00015074 | 19
er10_1k | 999.9736 | 0.0001751 | 19
er1_1k | 999.9638 | 0.0001199 | 19
es_100 | 999.9594 | 9.9641e-05 | 19
ep_1k | 99.9964 | 0.00010128 | 19
er1_100 | 99.9971 | 9.3599e-05 | 19
ep_100 | 9.9996 | 1.2468e-05 | 19
er1_10 | 9.9999 | 1.05e-05 | 19
es_10 | 99.999 | 0.00010012 | 19
v_100k | 100001.5383 | 0.0092966 | 19
v_10k | 10000.276 | 0.0015076 | 19
v_100 | 100.0132 | 0.00011932 | 19
Resistor cal set c | Value | Stdev | dof | Comment
sr104c | 9999.9698 | 0.0012382 | 19
sens_sr104c | 9969.188 | 0.0017682 | 19
ep_100k | 10000.0534 | 0.00137 | 19
es_100k | 1000004.8675 | 0.14039 | 19
v_1M | 999997.087 | 0.11704 | 19
Resistor ac Cal | Y dial a | Y dial b | Z dial a | Z dial b | Range | Frequency
v_1M_1592 | 0.000356 | 1.0001 | -0.00035 | 0.99992 | 7 | 1592
v_1M_160 | 0.000264 | 1 | -0.000259 | 0.99992 | 7 | 160
v_100k_1592 | 0.000625 | 0.99998 | -0.000625 | 1 | 6 | 1592
v_100k_160 | 0.000641 | 0.99997 | -0.000646 | 1 | 6 | 160
v_10k_1592 | 5.4e-05 | 1 | -5.5e-05 | 0.99996 | 5 | 1592
v_10k_160 | 5.5e-05 | 1 | -5.6e-05 | 0.99996 | 5 | 160
er1_1k_1592 | -0.000197 | 1.0001 | 0.000191 | 0.99989 | 4 | 1592
er1_1k_160 | -0.0002 | 1.0001 | 0.000201 | 0.9999 | 4 | 160
er1_100_1592 | -0.000532 | 1.0001 | 0.000532 | 0.9999 | 3 | 1592
er1_100_160 | -0.00054 | 1.0001 | 0.000538 | 0.99991 | 3 | 160
er1_10_1592 | -0.000358 | 1.0001 | 0.000356 | 0.99993 | 2 | 1592
er1_10_160 | -0.00037 | 1.0001 | 0.000369 | 0.99994 | 2 | 160
ep_10_1592 | -0.000355 | 1.0001 | 0.000359 | 0.99992 | 1 | 1592
ep_10_160 | -0.00036 | 1.0001 | 0.00037 | 0.99993 | 1 | 160
Resistor ac Cal zeros | Y dial a | Y dial b | Z dial a | Z dial b | Range | Frequency | need to be sourced for this spreadsheet
v_1M_1592 | 0 | 0 | 0 | 0 | 7 | 1592 | zeros not used in this spreadsheet
v_1M_160 | 0 | 0 | 0 | 0 | 7 | 160
v_100k_1592 | 0 | 0 | 0 | 0 | 6 | 1592
v_100k_160 | 0 | 0 | 0 | 0 | 6 | 160
v_10k_1592 | 0 | 0 | 0 | 0 | 5 | 1592
v_10k_160 | 0 | 0 | 0 | 0 | 5 | 160
er1_1k_1592 | 0 | 0 | 0 | 0 | 4 | 1592
er1_1k_160 | 0 | 0 | 0 | 0 | 4 | 160
er1_100_1592 | 0 | 0 | 0 | 0 | 3 | 1592
er1_100_160 | 0 | 0 | 0 | 0 | 3 | 160
er1_10_1592 | 0 | 0 | 0 | 0 | 2 | 1592
er1_10_160 | 0 | 0 | 0 | 0 | 2 | 160
ep_10_1592 | 0 | 0 | 0 | 0 | 1 | 1592
ep_10_160 | 0 | 0 | 0 | 0 | 1 | 160
Inductors | dial a | dial b | dial a zero | dial b zero | Range | Frequency
gr100uh_1k | 1.0003 | 0.086494 | 0.001438 | 0.0051924 | 1Z | 1000
gr1000uh_1k | 0.99878 | 0.086494 | 2.3e-05 | 0.0004305 | 2Z | 1000
gr10000uh_1k | 0.99975 | 0.086494 | 2e-06 | 4.27e-05 | 3Z | 1000
gr100000uh_1k | 0.9996 | 0.086494 | 0 | 3.5e-06 | 4Z | 1000
gr100000uh_100 | 1.002 | 0.086494 | 0 | 0 | 5Z | 1000
Capacitors | dial a | dial b | dial a zero | dial b zero | Range | Frequency
su100000_1k | 1.0001 | 4.33e-05 | 0 | 0 | 4Y | 1000
su1000000_1k | 1.0001 | 5.75e-05 | 0 | 0 | 3Y | 1000
gr1000a_1592 | 0.9999 | -1.9e-05 | 0 | 0 | 6Y | 1592
gr1000b_1592 | 0.99988 | -2.1e-05 | 0 | 0 | 6Y | 1592
gr100_1592 | 0.99995 | 0.000144 | 0 | 0 | 7Y | 1592
gr_10x100_1k | 0.99936 | 6e-05 | 0 | 0 | 3Y | 1000
gr_10x10n_1k | 1.002 | 3.76e-05 | 0 | 0 | 4Y | 1000
gr_10x1n_1k | 1.0042 | 3.89e-05 | 0 | 0 | 5Y | 1000
gr_0_1k | 0.004795 | 5.7e-06 | 0 | 0 | 5Y | 1000

## L and C
Calibration of Universal Bridge
Measurements on MSL inductors and capacitors
Date: | 2014-08-01
| Inductors | [15 empty] | 2004-02-18
Observation No. | Range | Nominal value microhenries | GR inductor S/N | Frequency | L | R | Lo | Ro | temp | Full scale inductance (uH) | Full scale resistance (ohms) | Corrected L dial | Corrected R dial | L, millihenires | R, ohms |  | L, millihenires | R, ohms | ppm shift in L
101/102 | 1Z | 100 | 18290 | 1000 | 1.0003 | 0.086494 | 0.001438 | 0.0051924 |  | 100 | 1 | 0.99902 | 0.081307 | 0.099902 | 0.081307 |  | 0.099907 | 0.081795 | 44.3825
103/108 | 2Z | 1000 | 18142 | 1000 | 0.99878 | 0.086004 | 2.3e-05 | 0.0004305 |  | 1000 | 10 | 0.99895 | 0.08558 | 0.99895 | 0.8558 |  | 0.99901 | 0.85334 | 59.4055
104/109 | 3Z | 10000 | 18413 | 1000 | 0.99975 | 0.084789 | 2e-06 | 4.27e-05 |  | 10000 | 100 | 0.99995 | 0.084753 | 9.9995 | 8.4753 |  | 9.9997 | 8.4806 | 24.3167
105/110 | 4Z | 100000 | 1879 | 1000 | 0.9996 | 0.083082 | 0 | 3.5e-06 |  | 100000 | 1000 | 0.9998 | 0.083084 | 99.9805 | 83.0842 |  | 99.9815 | 83.1956 | 10.3181
106 | 5Z | 1000000 | 18430 | 1000 | 1.002 | 0.064003 |  |  |  | 1000000 | 10000 | 1.0023 | 0.064008 | 1002.2501 | 640.0752 |  | 1002.2554 | 641.3039 | 5.2503
107 | 5Z | 1000000 | 18430 | 107 | 0.99972 | 0.063419 |  |  |  | 1000000 | 10000 | 0.99992 | 0.063423 | 999.9246 | 634.2347 |  | 999.935 | 635.4936 | 10.3143
Date: | 2014-08-01
| Capacitors | [15 empty] | Value | Last cal date | Reference
Observation No. | Range | Nominal value pF | Capacitor | Frequency | C | G |  |  |  | Full scale capacitance (pF) | Full scale conductance (uS) | Corrected C dial | G dial | C, (pF) | tan delta (urad) | Capacitance error ppm |  | Traceability
95 | 4Y | 100000 | Sullivan 57709/1960 | 1000 | 1.0001 | 4.33e-05 |  |  |  | 100000 | 1000 | 1.0002 | 4.3297e-05 | 100015.8514 | 68.91 | #DIV/0!
93 | 3Y | 1000000 | Sullivan 57708/1960 | 1000 | 1.0001 | 5.75e-05 |  |  |  | 1000000 | 10000 | 1.0002 | 5.7497e-05 | 1000165.5146 | 91.5086 | #DIV/0!
90 | 6Y | 1000 | GR 1000A | 1592 | 0.9999 | -1.9e-05 |  |  |  | 1000 | 10 | 1 | -1.9e-05 | 1000.0325 | -18.9947 | 4.6548 | 1000.0278 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
91 | 6Y | 1000 | GR1000B | 1592 | 0.99988 | -2.1e-05 |  |  |  | 1000 | 10 | 1 | -2.1e-05 | 1000.0185 | -20.9941 | 4.293 | 1000.0142 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
92 | 7Y | 100 | GR100 | 1592 | 0.99995 | 0.000144 |  |  |  | 100 | 1 | 1 | 0.00014399 | 100.0004 | 143.9488 | -32.1212 | 100.0037 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
97 | 3Y | 1000000 | GR box 3024 10x100nF | 1000 | 0.99936 | 6e-05 |  |  |  | 1000000 | 10000 | 0.99943 | 5.9996e-05 | 999385.5096 | 95.4873 | #DIV/0!
98 | 4Y | 100000 | GR box 3024 10x10nF | 1000 | 1.002 | 3.76e-05 |  |  |  | 100000 | 1000 | 1.0021 | 3.7598e-05 | 100161.5117 | 59.8387 | #DIV/0!
99 | 5Y | 10000 | GR box 3024 10x1nF | 1000 | 1.0042 | 3.89e-05 |  |  |  | 10000 | 100 | 1.0043 | 3.8898e-05 | 9995.1947 | 61.9076 | #DIV/0!
100 | 5Y | 0 | GR box 3024 dials zero | 1000 | 0.004795 | 5.7e-06 |  |  |  | 10000 | 100 | 0.0047953 | 5.6997e-06 | 47.9534
box zero subtracted
note that the AH2700A is two terminal-pair and expect lead corrections of the order of 50 ppm for 1 microfarad load.
With correction term "c"
|  |  | correction to L/C dial | correction to R/G dial
Range
|  |  | (ppm) | (ppm)
| 1Z |  | 200.7 | 70.5485
| 2Z |  | 200.7 | 70.5485
| 3Z |  | 200.7 | 70.5485
| 4Z |  | 200.7 | 70.5485
| 5Z |  | 200.7 | 70.5485
| 6Z |  | 136.7192 | 6.5677
| 7Z |  | 213.7672 | 83.6157
| 1Y |  | 70.5079 | -59.6436
| 2Y |  | 70.5079 | -59.6436
| 3Y |  | 70.5079 | -59.6436
| 4Y |  | 70.5079 | -59.6436
| 5Y |  | 70.5079 | -59.6436
| 6Y |  | 134.4887 | 4.3372
| 7Y |  | 57.4407 | -72.7108
Note on T10: Keith Jones: corrected transcription error in old spreadsheet

## Corrections
This table is the usual form for calculating spreadsheet corrections
Component | Value
(ppm)
C_1 | 143.1
R_4A | 54.1261
R_4B | -9.8547
R_4C | 67.1933
G_1 | 12.9485
G_2 | 7.4961 |  | Correction term d is
assumed negligible.
Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | Correction term c is
where Z3=R+jwC |  | correction_c | 10.97 | see 'Resistors' sheet, Comment block N74
Y4=(G3/G2)(aG1+ßjwC1)(1-c-jd) |  |  | (G3=1/R4*)
where Y4=G+jwC
R4A is used for ranges 5 to 1 inclusive
R4B is used on range 6 only
R4C is used on range 7 only
Without correction term "c" |  |  |  |  | With correction term "c"
|  | correction to L/C dial | correction to R/G dial |  |  | correction to L/C dial | correction to R/G dial
Range |  | b correction | a correction |  | Range | b correction | a correction
|  | (ppm) | (ppm) |  |  | (ppm) | (ppm)
1Z |  | 189.73 | 59.5785 |  | 1Z | 200.7 | 70.5485
2Z |  | 189.73 | 59.5785 |  | 2Z | 200.7 | 70.5485
3Z |  | 189.73 | 59.5785 |  | 3Z | 200.7 | 70.5485
4Z |  | 189.73 | 59.5785 |  | 4Z | 200.7 | 70.5485
5Z |  | 189.73 | 59.5785 |  | 5Z | 200.7 | 70.5485
6Z |  | 125.7492 | -4.4023 |  | 6Z | 136.7192 | 6.5677
7Z |  | 202.7972 | 72.6457 |  | 7Z | 213.7672 | 83.6157
1Y |  | 81.4779 | -48.6736 |  | 1Y | 70.5079 | -59.6436
2Y |  | 81.4779 | -48.6736 |  | 2Y | 70.5079 | -59.6436
3Y |  | 81.4779 | -48.6736 |  | 3Y | 70.5079 | -59.6436
4Y |  | 81.4779 | -48.6736 |  | 4Y | 70.5079 | -59.6436
5Y |  | 81.4779 | -48.6736 |  | 5Y | 70.5079 | -59.6436
6Y |  | 145.4587 | 15.3072 |  | 6Y | 134.4887 | 4.3372
7Y |  | 68.4107 | -61.7408 |  | 7Y | 57.4407 | -72.7108

## Uncertainties
Universal Bridge Uncertainties compared with uncertainties prior to Aug-Nov 2009 calibration
Note: This sheet is not protected.
Assumptions |  | Changes in uncertainties
1) The uncertainties apply to ranges 1 to 6 |  |  |  |  | Revert phase interaction to 16, otherwise
2) The uncertainties apply for frequencies from 160 Hz to 1592 Hz |  |  |  |  | July 2014 uncertainties
are same as for 2009
| Previous uncertainties | Feb 2004 uncertainties | Feb 2005 uncertainties | Aug 2009 uncertainties | July 2014 uncertainties
| Standard | Standard | Standard | Standard | Standard
Source of uncertainty | Uncertainties | Uncertainties | Uncertainties | Uncertainties | Uncertainties | Degrees of
| u (ppm of full dial scale) | u (ppm of full dial scale) | u (ppm of full dial scale) |  |  | Freedom, v | Previous Comment | Feb 2004 Comment | Feb 2005 Comment | Aug 2009 comment
INDUCTANCE
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here |  | problem is similar, increase is not clearly significant
component stability | 2 | 5 | 5 | 5 | 5 | 2 | some doubt on stability of capacitor in particular | the 1000 pF GR capacitor can change 5 to 10 ppm short term
dial linearity | 1 | 1 | 1 | 1 | 1 | 5.5556 | 1 ppm of full scale
inductance error due to phase error in resistance | 16 | 5 | 5 | 5 | 16 | 13 | 16 ppm of resistance dial setting, see comment for resistance | If no correction applied. See comment E53, Short, Open, Phase.
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be range and frequency dependent
temperature coefficient of capacitor | 2 | 2 | 2 | 2 | 2 | 5.5556 | 2 ppm per degree for capacitor
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
RESISTANCE of an Inductor
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 2 | 2 | 2 | 2 | 5.5556 | stability of resistors only
dial linearity | 2 | 2 | 2 | 1 | 1 | 5.6 |  |  |  | match to inductance linearity
resistance error due to phase error in inductance | 16 | 5 | 5 | 5 | 5 | 13 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | If no correction applied. See comment E53, Short, Open, Phase.
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be dependent on frequency and range, included below.
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
CAPACITANCE
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 5 | 5 | 5 | 5 | 2 | some doubt on stability of capacitor in particular | the 1000 pF GR capacitor can change 5 to 10 ppm short term
dial linearity | 1 | 1 | 1 | 1 | 1 | 5.5556 | 1 ppm of full scale
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be range and frequency dependent
temperature coefficient of capacitor | 2 | 2 | 2 | 2 | 2 | 5.5556 | 2 ppm per degree for capacitor
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
CONDUCTANCE of a Capacitor
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 2 | 2 | 2 | 2 | 5.5556 | stability of resistors only
conductance error due to phase error in capacitance | 16 | 16 | 16 | 16 | 16 | 13 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | If no correction applied. See comment E53, Short, Open, Phase. |  | Phase error same as for inductor
dial linearity | 2 | 2 | 2 | 1 | 1 | 5.6 |  |  |  | match to capacitance linearity
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be dependent on frequency and range, included below.
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components

## docx
Calibration Log
This is a first time to track the calibration of dc resistors for the universal bridge by copying in the output of Python scripts for operating the HP3458A.
ACAL ALL
Do this at the beginning of every day.
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 49.9868902E+00 Hz
Temperature = 32.8
ACAL ALL starting
ACAL finished | .
Time taken = 16.0816833337 minutes
Temperature = 33.0
Mon Jul 07 10:02:35 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
Measure SR104 10k
Using DIN lead to connect
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104 temperature sensor
Trial 9.968847642E+03
0 9.968848812E+03
1 9.968849556E+03
2 9.968848774E+03
3 9.968851888E+03
4 9.968849861E+03
5 9.968850095E+03
6 9.968851475E+03
7 9.968851193E+03
8 9.968850259E+03
9 9.968848787E+03
10 9.968852180E+03
11 9.968852252E+03
12 9.968850656E+03
13 9.968852848E+03
14 9.968853640E+03
15 9.968853569E+03
16 9.968852335E+03
17 9.968853911E+03
18 9.968853262E+03
19 9.968851806E+03
Result
Mon Jul 07 11:11:43 2014
9968.8514
0.0016853
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104
Trial 9.999973446E+03
0 9.999974730E+03
1 9.999975014E+03
2 9.999974547E+03
3 9.999974136E+03
4 9.999972150E+03
5 9.999972246E+03
6 9.999974452E+03
7 9.999971986E+03
8 9.999972802E+03
9 9.999974782E+03
10 9.999975478E+03
11 9.999974745E+03
12 9.999972656E+03
13 9.999971060E+03
14 9.999973176E+03
15 9.999973711E+03
16 9.999973923E+03
17 9.999974118E+03
18 9.999973191E+03
19 9.999972615E+03
Result
Mon Jul 07 11:22:27 2014
9999.9736
0.001206
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104 repeat
Trial 9.999973826E+03
0 9.999973572E+03
1 9.999974375E+03
2 9.999972206E+03
3 9.999975232E+03
4 9.999973291E+03
5 9.999972529E+03
6 9.999973886E+03
7 9.999972752E+03
8 9.999973631E+03
9 9.999974736E+03
10 9.999971425E+03
11 9.999975183E+03
12 9.999974039E+03
13 9.999973422E+03
14 9.999972873E+03
15 9.999972462E+03
16 9.999974698E+03
17 9.999973321E+03
18 9.999974010E+03
19 9.999973134E+03
Result
Mon Jul 07 11:25:05 2014
9999.9735
0.001011
>Exit code: 0
GMH51 reads 19.81 °C in well of SR104
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? UB R4A guard open
Trial 1.000051105E+04
0 1.000051677E+04
1 1.000051329E+04
2 1.000051552E+04
3 1.000051277E+04
4 1.000051393E+04
5 1.000051192E+04
6 1.000051119E+04
7 1.000051371E+04
8 1.000051637E+04
9 1.000051443E+04
10 1.000051400E+04
11 1.000051349E+04
12 1.000051429E+04
13 1.000051503E+04
14 1.000051428E+04
15 1.000051693E+04
16 1.000051126E+04
17 1.000051550E+04
18 1.000050864E+04
19 1.000051388E+04
Result
Mon Jul 07 11:43:03 2014
10000.5139
0.0020299
>Exit code: 0
GMH51 reads 19.65°C at R4A
This next run had error messag…wrong range?
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :!.0E4
Item? UB R4A, guard to low, guard to case in both instances
Trial 1.000051184E+04
0 1.000050288E+04
1 1.000054114E+04
2 1.000051742E+04
3 1.000050780E+04
4 1.000027609E+04
5 1.000047247E+04
6 1.000068019E+04
7 1.000052316E+04
8 1.000054238E+04
9 1.000060744E+04
10 1.000046489E+04
11 1.000052452E+04
12 1.000051911E+04
13 1.000051009E+04
14 1.000052685E+04
15 1.000052598E+04
16 1.000051702E+04
17 1.000051372E+04
18 1.000051516E+04
19 1.000051513E+04
Result
Mon Jul 07 11:46:34 2014
10000.5152
0.072305
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, guard to low and case, correct range?
Trial 1.000052034E+04
0 1.000047879E+04
1 1.000051572E+04
2 1.000051427E+04
3 1.000051448E+04
4 1.000051173E+04
5 1.000052494E+04
6 1.000050944E+04
7 1.000051340E+04
8 1.000051265E+04
9 1.000051301E+04
10 1.000051536E+04
11 1.000051105E+04
12 1.000051434E+04
13 1.000051467E+04
14 1.000051164E+04
15 1.000051268E+04
16 1.000050846E+04
17 1.000051377E+04
18 1.000051271E+04
19 1.000051169E+04
Result
Mon Jul 07 11:49:40 2014
10000.5117
0.0084162
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E4
Item? UB R4A, guard open and to screen
Trial 1.000044113E+04
0 1.000051299E+04
1 1.000051810E+04
2 1.000050990E+04
3 1.000051338E+04
4 1.000051482E+04
5 1.000050946E+04
6 1.000051668E+04
7 1.000050979E+04
8 1.000051601E+04
9 1.000051061E+04
10 1.000051345E+04
11 1.000051100E+04
12 1.000051316E+04
13 1.000051236E+04
14 1.000051058E+04
15 1.000051413E+04
16 1.000051436E+04
17 1.000051031E+04
18 1.000051388E+04
19 1.000051327E+04
Result
Mon Jul 07 11:52:54 2014
10000.5129
0.0024255
>Exit code: 0
GMH51 19.60 °C at R4A.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, after lunch break, now with DELAY set to 2 seconds
Trial 1.000052589E+04
0 1.000052074E+04
1 1.000051917E+04
2 1.000051844E+04
3 1.000051013E+04
4 1.000051697E+04
5 1.000051094E+04
6 1.000051846E+04
7 1.000051544E+04
8 1.000049867E+04
9 1.000051633E+04
10 1.000053928E+04
11 1.000051701E+04
12 1.000052019E+04
13 1.000051840E+04
14 1.000051891E+04
15 1.000051976E+04
16 1.000051501E+04
17 1.000052022E+04
18 1.000052817E+04
19 1.000050408E+04
Result
Mon Jul 07 13:02:18 2014
10000.5173
0.0081742
>Exit code: 0
GMH51 19.52 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2e4
Item? UB R4A, no special delay after trigger, guard still open
Trial 1.000051283E+04
0 1.000052131E+04
1 1.000051804E+04
2 1.000051886E+04
3 1.000051461E+04
4 1.000055984E+04
5 1.000046729E+04
6 1.000055989E+04
7 1.000047512E+04
8 1.000054184E+04
9 1.000049248E+04
10 1.000052779E+04
11 1.000048655E+04
12 1.000053447E+04
13 1.000048190E+04
14 1.000054823E+04
15 1.000054815E+04
16 1.000047037E+04
17 1.000052479E+04
18 1.000050109E+04
19 1.000051294E+04
Result
Mon Jul 07 13:08:38 2014
10000.5153
0.029127
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, no special delay, Guard still open
Trial 1.000051478E+04
0 1.000054698E+04
1 1.000051195E+04
2 1.000051439E+04
3 1.000051689E+04
4 1.000050801E+04
5 1.000052710E+04
6 1.000050310E+04
7 1.000052782E+04
8 1.000050876E+04
9 1.000051984E+04
10 1.000051878E+04
11 1.000051292E+04
12 1.000052282E+04
13 1.000051277E+04
14 1.000054627E+04
15 1.000050765E+04
16 1.000051395E+04
17 1.000052917E+04
18 1.000050868E+04
19 1.000051352E+04
Result
Mon Jul 07 13:12:05 2014
10000.5186
0.011933
>Exit code: 0
The above 2 runs after lunch seem to start with one high value that shows as a higher standard deviation.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? UB R4A, try again to see if standard deviation comes down
Trial 1.000052412E+04
0 1.000049257E+04
1 1.000051917E+04
2 1.000051294E+04
3 1.000052710E+04
4 1.000050427E+04
5 1.000051553E+04
6 1.000051450E+04
7 1.000051694E+04
8 1.000051604E+04
9 1.000051280E+04
10 1.000051749E+04
11 1.000051597E+04
12 1.000051586E+04
13 1.000051458E+04
14 1.000051526E+04
15 1.000051668E+04
16 1.000051302E+04
17 1.000052021E+04
18 1.000051263E+04
19 1.000051755E+04
Result
Mon Jul 07 13:17:35 2014
10000.5146
0.0066755
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? UB R4A, one more preliminary trial added
Trial 1.000050991E+04
Trial 2 1.000051573E+04
0 1.000051664E+04
1 1.000051542E+04
2 1.000051310E+04
3 1.000051525E+04
4 1.000051363E+04
5 1.000051658E+04
6 1.000051659E+04
7 1.000051617E+04
8 1.000051439E+04
9 1.000051410E+04
10 1.000051457E+04
11 1.000051342E+04
12 1.000051518E+04
13 1.000051140E+04
14 1.000051286E+04
15 1.000051382E+04
16 1.000051251E+04
17 1.000051281E+04
18 1.000051286E+04
19 1.000051499E+04
Result
Mon Jul 07 13:20:44 2014
10000.5143
0.0015186
>Exit code: 0
GMH51 19.55 °C
Note better behaviour above with the extra preliminary reading. Data could be reprocessed to eliminate odd initial readings…then keeping guard to low would make sense.
Now proceed with an ESI buildup to give us 100 k and 1 M values.
Start with SR104 again
Guard to lo and connected to screen is now the default
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? SR104 thermometer
Trial 9.969754446E+03
Trial 2 9.969752259E+03
0 9.969753171E+03
1 9.969755265E+03
2 9.969752744E+03
3 9.969753893E+03
4 9.969755510E+03
5 9.969754812E+03
6 9.969756225E+03
7 9.969756247E+03
8 9.969755079E+03
9 9.969754888E+03
10 9.969755070E+03
11 9.969753998E+03
12 9.969755202E+03
13 9.969754467E+03
14 9.969755483E+03
15 9.969755763E+03
16 9.969755918E+03
17 9.969758751E+03
18 9.969757845E+03
19 9.969756461E+03
Result
Mon Jul 07 13:46:06 2014
9969.7553
0.0014158
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
1Range? :1.2e4
Item? SR104 thermometer repeat
Trial 9.969757629E+03
Trial 2 9.969768426E+03
0 9.969783854E+03
1 9.969764005E+03
2 9.969757494E+03
3 9.969774507E+03
4 9.969777351E+03
5 9.969757395E+03
6 9.969778043E+03
7 9.969772016E+03
8 9.969764523E+03
9 9.969780099E+03
10 9.969762923E+03
11 9.969765847E+03
12 9.969763095E+03
13 9.969780969E+03
14 9.969768253E+03
15 9.969782385E+03
16 9.969779386E+03
17 9.969754849E+03
18 9.969784565E+03
19 9.969766940E+03
Result
Mon Jul 07 13:48:45 2014
9969.7709
0.0095478
>Exit code: 0
GMH51 19.88 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2e4
Item? SR104
Trial 9.999972622E+03
Trial 2 9.999970557E+03
0 9.999972805E+03
1 9.999973710E+03
2 9.999973138E+03
3 9.999975826E+03
4 9.999973500E+03
5 9.999971629E+03
6 9.999971824E+03
7 9.999971934E+03
8 9.999972801E+03
9 9.999971968E+03
10 9.999972916E+03
11 9.999971712E+03
12 9.999973280E+03
13 9.999974768E+03
14 9.999972485E+03
15 9.999974247E+03
16 9.999973615E+03
17 9.999973429E+03
18 9.999972278E+03
19 9.999972582E+03
Result
Mon Jul 07 13:56:58 2014
9999.973
0.0010819
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
1Range? :1.2e4
Item? SR104 repeat
Trial 9.999972281E+03
Trial 2 9.999969798E+03
0 9.999970333E+03
1 9.999973018E+03
2 9.999971955E+03
3 9.999972642E+03
4 9.999974570E+03
5 9.999970514E+03
6 9.999971585E+03
7 9.999970321E+03
8 9.999972314E+03
9 9.999972744E+03
10 9.999971399E+03
11 9.999971101E+03
12 9.999971577E+03
13 9.999971836E+03
14 9.999972996E+03
15 9.999971721E+03
16 9.999973228E+03
17 9.999971913E+03
18 9.999971466E+03
19 9.999970512E+03
Result
Mon Jul 07 13:59:39 2014
9999.9719
0.0010976
>Exit code: 0
GMH51 19.91 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2e4
Item? ESI 100 k/step in parallel
Trial 1.000004489E+04
Trial 2 1.000003969E+04
0 1.000004267E+04
1 1.000004232E+04
2 1.000004242E+04
3 1.000004483E+04
4 1.000004193E+04
5 1.000004357E+04
6 1.000004127E+04
7 1.000004281E+04
8 1.000004205E+04
9 1.000004411E+04
10 1.000004482E+04
11 1.000004401E+04
12 1.000004213E+04
13 1.000004370E+04
14 1.000004166E+04
15 1.000004122E+04
16 1.000004310E+04
17 1.000004374E+04
18 1.000004264E+04
19 1.000004363E+04
Result
Mon Jul 07 14:06:30 2014
10000.0429
0.0010864
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? ESI 100 k/step, with parallel bars and network as before
Trial 1.000004600E+04
Trial 2 1.000004386E+04
0 1.000004403E+04
1 1.000004318E+04
2 1.000004348E+04
3 1.000004129E+04
4 1.000004017E+04
5 1.000004470E+04
6 1.000004281E+04
7 1.000004242E+04
8 1.000004286E+04
9 1.000004135E+04
10 1.000004320E+04
11 1.000004179E+04
12 1.000004017E+04
13 1.000004132E+04
14 1.000004045E+04
15 1.000004219E+04
16 1.000004128E+04
17 1.000004109E+04
18 1.000004435E+04
19 1.000004015E+04
Result
Mon Jul 07 14:11:36 2014
10000.0421
0.0014255
>Exit code: 0
GMH51 20.07 °C next to ESI box
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? ESI 100k/step in parallel, repeat (distracted by email to Laly)
Trial 1.000003918E+04
Trial 2 1.000003865E+04
0 1.000004056E+04
1 1.000004283E+04
2 1.000003867E+04
3 1.000004095E+04
4 1.000004042E+04
5 1.000004087E+04
6 1.000004012E+04
7 1.000003737E+04
8 1.000003809E+04
9 1.000003920E+04
10 1.000003750E+04
11 1.000004008E+04
12 1.000003788E+04
13 1.000003379E+04
14 1.000003932E+04
15 1.000003986E+04
16 1.000003920E+04
17 1.000003612E+04
18 1.000003831E+04
19 1.000003994E+04
Result
Mon Jul 07 14:45:30 2014
10000.0391
0.0019682
>Exit code: 0
GMH51 20.01 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step in series parallel on 1-9
Trial 1.000002318E+05
Trial 2 1.000002333E+05
0 1.000001983E+05
1 1.000002073E+05
2 1.000001877E+05
3 1.000001935E+05
4 1.000002095E+05
5 1.000002126E+05
6 1.000001944E+05
7 1.000002033E+05
8 1.000002193E+05
9 1.000001998E+05
10 1.000002189E+05
11 1.000002049E+05
12 1.000002117E+05
13 1.000002004E+05
14 1.000001984E+05
15 1.000001999E+05
16 1.000001974E+05
17 1.000002061E+05
18 1.000001741E+05
19 1.000002034E+05
Result
Mon Jul 07 14:53:59 2014
100000.202
0.010456
>Exit code: 0
GMH51 20.07 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step in s-p 1-9, repeat
Trial 1.000002125E+05
Trial 2 1.000002080E+05
0 1.000002132E+05
1 1.000002136E+05
2 1.000002040E+05
3 1.000002001E+05
4 1.000001778E+05
5 1.000002069E+05
6 1.000001998E+05
7 1.000001934E+05
8 1.000001969E+05
9 1.000002015E+05
10 1.000001930E+05
11 1.000001828E+05
12 1.000001848E+05
13 1.000001809E+05
14 1.000001838E+05
15 1.000001979E+05
16 1.000002046E+05
17 1.000002011E+05
18 1.000001961E+05
19 1.000001986E+05
Result
Mon Jul 07 14:57:05 2014
100000.1965
0.010205
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
120000
Item? ESI 100 k/step coil 10 only
Trial 9.999990254E+04
Trial 2 9.999988844E+04
0 9.999989239E+04
1 9.999988490E+04
2 9.999987999E+04
3 9.999988814E+04
4 9.999987455E+04
5 9.999987447E+04
6 9.999987710E+04
7 9.999986630E+04
8 9.999986782E+04
9 9.999988442E+04
10 9.999987215E+04
11 9.999985706E+04
12 9.999987768E+04
13 9.999987282E+04
14 9.999985879E+04
15 9.999986546E+04
16 9.999985970E+04
17 9.999986167E+04
18 9.999986196E+04
19 9.999986205E+04
Result
Mon Jul 07 15:02:38 2014
99999.872
0.010464
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step coil 10 repeat
Trial 9.999987439E+04
Trial 2 9.999987560E+04
0 9.999986246E+04
1 9.999984082E+04
2 9.999985943E+04
3 9.999984823E+04
4 9.999986350E+04
5 9.999986447E+04
6 9.999984148E+04
7 9.999984114E+04
8 9.999986801E+04
9 9.999984083E+04
10 9.999983898E+04
11 9.999984853E+04
12 9.999984752E+04
13 9.999986121E+04
14 9.999986661E+04
15 9.999985507E+04
16 9.999986677E+04
17 9.999985653E+04
18 9.999985792E+04
19 9.999986690E+04
Result
Mon Jul 07 15:04:43 2014
99999.8548
0.01038
>Exit code: 0
GMH51 20.10 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E6
Item? ESI 100k/step 1-10 in series
Trial 9.999999411E+05
Trial 2 1.000000261E+06
0 1.000000511E+06
1 1.000000485E+06
2 1.000000661E+06
3 1.000000075E+06
4 1.000000824E+06
5 1.000000551E+06
6 1.000000941E+06
7 1.000000650E+06
8 1.000000773E+06
9 1.000000617E+06
10 1.000000732E+06
11 1.000001057E+06
12 1.000001001E+06
13 1.000001072E+06
14 1.000001005E+06
15 1.000001410E+06
16 1.000001335E+06
17 1.000001110E+06
18 1.000001290E+06
19 1.000001315E+06
Result
Mon Jul 07 15:11:06 2014
1000000.8708
0.3444
>Exit code: 0
GMH51 20.11 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? ESI 100k/step 1-10 in series, repeat
Trial 1.000002540E+06
Trial 2 1.000002393E+06
0 1.000002853E+06
1 1.000002656E+06
2 1.000002477E+06
3 1.000002456E+06
4 1.000002490E+06
5 1.000002296E+06
6 1.000002549E+06
7 1.000002195E+06
8 1.000002461E+06
9 1.000002326E+06
10 1.000002483E+06
11 1.000002379E+06
12 1.000002663E+06
13 1.000002427E+06
14 1.000002362E+06
15 1.000002331E+06
16 1.000002983E+06
17 1.000002585E+06
18 1.000002991E+06
19 1.000002926E+06
Result
Mon Jul 07 15:13:35 2014
1000002.5445
0.23377
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R4B,screen connected to guard
Trial 9.999851622E+04
Trial 2 9.999852071E+04
0 9.999851887E+04
1 9.999852840E+04
2 9.999852880E+04
3 9.999851685E+04
4 9.999852670E+04
5 9.999849643E+04
6 9.999850678E+04
7 9.999854300E+04
8 9.999848452E+04
9 9.999850813E+04
10 9.999853209E+04
11 9.999851728E+04
12 9.999855155E+04
13 9.999846786E+04
14 9.999850982E+04
15 9.999852877E+04
16 9.999850358E+04
17 9.999852062E+04
18 9.999852889E+04
19 9.999851465E+04
Result
Mon Jul 07 15:50:00 2014
99998.5167
0.019288
>Exit code: 0
GMH51 19.64 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R4B, repeat
Trial 9.999851985E+04
Trial 2 9.999854826E+04
0 9.999854139E+04
1 9.999851160E+04
2 9.999851911E+04
3 9.999852385E+04
4 9.999850656E+04
5 9.999852170E+04
6 9.999849829E+04
7 9.999850536E+04
8 9.999851757E+04
9 9.999851992E+04
10 9.999851615E+04
11 9.999851062E+04
12 9.999852176E+04
13 9.999853130E+04
14 9.999853575E+04
15 9.999851119E+04
16 9.999850832E+04
17 9.999849515E+04
18 9.999851274E+04
19 9.999852954E+04
Result
Mon Jul 07 15:53:39 2014
99998.5169
0.011883
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB G1 through BPO tees
Trial 9.999821890E+04
Trial 2 9.999825721E+04
0 9.999822478E+04
1 9.999824312E+04
2 9.999822837E+04
3 9.999822644E+04
4 9.999831956E+04
5 9.999816724E+04
6 9.999827530E+04
7 9.999823707E+04
8 9.999820014E+04
9 9.999825013E+04
10 9.999820758E+04
11 9.999823278E+04
12 9.999824534E+04
13 9.999819385E+04
14 9.999823569E+04
15 9.999824552E+04
16 9.999823241E+04
17 9.999823336E+04
18 9.999821847E+04
19 9.999820986E+04
Result
Mon Jul 07 15:58:11 2014
99998.2314
0.031135
>Exit code: 0
GMH51 19.65 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E5
Item? UB G1, repeat
Trial 9.999820850E+04
Trial 2 9.999824112E+04
0 9.999823716E+04
1 9.999820791E+04
2 9.999821928E+04
3 9.999820587E+04
4 9.999822005E+04
5 9.999821540E+04
6 9.999820422E+04
7 9.999820865E+04
8 9.999819905E+04
9 9.999820661E+04
10 9.999819155E+04
11 9.999820612E+04
12 9.999819771E+04
13 9.999819191E+04
14 9.999820451E+04
15 9.999821002E+04
16 9.999821671E+04
17 9.999819260E+04
18 9.999819907E+04
19 9.999821590E+04
Result
Mon Jul 07 16:00:55 2014
99998.2075
0.011203
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E5
Item? UB R2, through tees
Trial 9.999872717E+04
Trial 2 9.999875119E+04
0 9.999875371E+04
1 9.999876633E+04
2 9.999874843E+04
3 9.999873617E+04
4 9.999874592E+04
5 9.999874689E+04
6 9.999875441E+04
7 9.999875031E+04
8 9.999875558E+04
9 9.999874941E+04
10 9.999875476E+04
11 9.999875096E+04
12 9.999876137E+04
13 9.999875080E+04
14 9.999875900E+04
15 9.999874084E+04
16 9.999874885E+04
17 9.999876419E+04
18 9.999876074E+04
19 9.999875652E+04
Result
Mon Jul 07 16:04:27 2014
99998.7528
0.0075329
>Exit code: 0
GMH51 19.65 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R2, repeat
Trial 9.999873303E+04
Trial 2 9.999876652E+04
0 9.999875159E+04
1 9.999874877E+04
2 9.999874038E+04
3 9.999875179E+04
4 9.999875089E+04
5 9.999876352E+04
6 9.999873818E+04
7 9.999874177E+04
8 9.999875559E+04
9 9.999873250E+04
10 9.999874771E+04
11 9.999875003E+04
12 9.999873891E+04
13 9.999876193E+04
14 9.999873841E+04
15 9.999873134E+04
16 9.999873348E+04
17 9.999875415E+04
18 9.999875171E+04
19 9.999873934E+04
Result
Mon Jul 07 16:07:34 2014
99998.7461
0.0094139
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E6
Item? UB R4C
Trial 1.000064944E+06
Trial 2 1.000063309E+06
0 1.000063088E+06
1 1.000063522E+06
2 1.000062799E+06
3 1.000063077E+06
4 1.000062732E+06
5 1.000063157E+06
6 1.000063340E+06
7 1.000063490E+06
8 1.000063180E+06
9 1.000063129E+06
10 1.000062883E+06
11 1.000062625E+06
12 1.000062786E+06
13 1.000063017E+06
14 1.000063443E+06
15 1.000063546E+06
16 1.000063724E+06
17 1.000063227E+06
18 1.000063196E+06
19 1.000062793E+06
Result
Mon Jul 07 16:12:22 2014
1000063.1377
0.30782
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E6
Item? UB R4C, repeat
Trial 1.000065819E+06
Trial 2 1.000067191E+06
0 1.000066664E+06
1 1.000064486E+06
2 1.000063662E+06
3 1.000063648E+06
4 1.000065137E+06
5 1.000067281E+06
6 1.000068879E+06
7 1.000069332E+06
8 1.000068557E+06
9 1.000066801E+06
10 1.000065168E+06
11 1.000064379E+06
12 1.000064598E+06
13 1.000065889E+06
14 1.000067184E+06
15 1.000068094E+06
16 1.000068104E+06
17 1.000067190E+06
18 1.000067680E+06
19 1.000067224E+06
Result
Mon Jul 07 16:40:29 2014
1000066.4979
1.7619
>Exit code: 0
GMH51 19.58 °C
Next calibrate ESI set of resistors
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 50.0514113E+00 Hz
Temperature = 32.9
ACAL ALL starting
ACAL finished
Time taken = 16.2855000019 minutes
Temperature = 32.9
Thu Jul 17 13:34:55 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E4
Item? SR104 thermistor
Trial 9.971495330E+03
Trial 2 9.971496304E+03
0 9.971495573E+03
1 9.971496996E+03
2 9.971496629E+03
3 9.971495857E+03
4 9.971494343E+03
5 9.971495549E+03
6 9.971495693E+03
7 9.971495905E+03
8 9.971494767E+03
9 9.971497310E+03
10 9.971497057E+03
11 9.971501974E+03
12 9.971499281E+03
13 9.971497535E+03
14 9.971498418E+03
15 9.971495111E+03
16 9.971497004E+03
17 9.971495674E+03
18 9.971498310E+03
19 9.971495692E+03
Result
Thu Jul 17 13:45:11 2014
9971.4967
0.0017802
>Exit code: 0
GMH51 reads 20.06 degrees
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? SR104
Trial 9.999969676E+03
Trial 2 9.999969541E+03
0 9.999967457E+03
1 9.999968973E+03
2 9.999967447E+03
3 9.999970984E+03
4 9.999967750E+03
5 9.999970241E+03
6 9.999968971E+03
7 9.999970669E+03
8 9.999970184E+03
9 9.999969089E+03
10 9.999968120E+03
11 9.999970090E+03
12 9.999970392E+03
13 9.999967860E+03
14 9.999967198E+03
15 9.999968372E+03
16 9.999968818E+03
17 9.999966314E+03
18 9.999967115E+03
19 9.999968399E+03
Result
Thu Jul 17 13:52:00 2014
9999.9687
0.0013473
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? SR104 repeat
Trial 9.999969196E+03
Trial 2 9.999968014E+03
0 9.999968116E+03
1 9.999969293E+03
2 9.999968547E+03
3 9.999965997E+03
4 9.999966624E+03
5 9.999967819E+03
6 9.999966747E+03
7 9.999966218E+03
8 9.999967544E+03
9 9.999967935E+03
10 9.999970729E+03
11 9.999967323E+03
12 9.999966437E+03
13 9.999967657E+03
14 9.999966065E+03
15 9.999967762E+03
16 9.999967547E+03
17 9.999967953E+03
18 9.999967937E+03
19 9.999968175E+03
Result
Thu Jul 17 13:55:30 2014
9999.9676
0.001136
>Exit code: 0
GMH51 in SR104 well 20.07 degrees
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E4
Item? ES! 1k 1-10 series
Trial 9.999645202E+03
Trial 2 9.999645817E+03
0 9.999644582E+03
1 9.999645575E+03
2 9.999644686E+03
3 9.999645225E+03
4 9.999645158E+03
5 9.999644638E+03
6 9.999644933E+03
7 9.999645746E+03
8 9.999646619E+03
9 9.999644517E+03
10 9.999646996E+03
11 9.999644402E+03
12 9.999644341E+03
13 9.999643298E+03
14 9.999644679E+03
15 9.999644022E+03
16 9.999645659E+03
17 9.999645362E+03
18 9.999644955E+03
19 9.999644741E+03
Result
Thu Jul 17 14:05:00 2014
9999.645
0.00084579
>Exit code: 0
GMH51 20.05 degrees by ESI box
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E3
Item? ESI 1k 1-9 S/P
Trial 9.999632104E+02
Trial 2 9.999630321E+02
0 9.999631495E+02
1 9.999631178E+02
2 9.999631765E+02
3 9.999628805E+02
4 9.999632166E+02
5 9.999631175E+02
6 9.999629689E+02
7 9.999629700E+02
8 9.999629707E+02
9 9.999628537E+02
10 9.999630094E+02
11 9.999630536E+02
12 9.999628324E+02
13 9.999630765E+02
14 9.999632411E+02
15 9.999629403E+02
16 9.999630600E+02
17 9.999627381E+02
18 9.999626925E+02
19 9.999629922E+02
Result
Thu Jul 17 14:14:26 2014
999.963
0.00015074
>Exit code: 0
GMH51 20.09
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E3
Item? ESI 1k R10
Trial 9.999737655E+02
Trial 2 9.999737940E+02
0 9.999736473E+02
1 9.999737959E+02
2 9.999738792E+02
3 9.999737596E+02
4 9.999736667E+02
5 9.999735997E+02
6 9.999735344E+02
7 9.999738024E+02
8 9.999737068E+02
9 9.999735789E+02
10 9.999734382E+02
11 9.999736382E+02
12 9.999733559E+02
13 9.999733443E+02
14 9.999736660E+02
15 9.999735552E+02
16 9.999735056E+02
17 9.999732994E+02
18 9.999733617E+02
19 9.999733233E+02
Result
Thu Jul 17 14:19:33 2014
999.9736
0.0001751
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E3
Item? ESI 1k R1
Trial 9.999639804E+02
Trial 2 9.999638130E+02
0 9.999636698E+02
1 9.999636799E+02
2 9.999638243E+02
3 9.999636862E+02
4 9.999638516E+02
5 9.999640541E+02
6 9.999636862E+02
7 9.999637291E+02
8 9.999639401E+02
9 9.999638077E+02
10 9.999640010E+02
11 9.999637841E+02
12 9.999638811E+02
13 9.999638383E+02
14 9.999637598E+02
15 9.999638422E+02
16 9.999636058E+02
17 9.999637151E+02
18 9.999639641E+02
19 9.999638198E+02
Result
Thu Jul 17 14:24:20 2014
999.9638
0.0001199
>Exit code: 0
GMH51 20.17 degrees.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 1k 1-10 p
Trial 9.999619231E+01
Trial 2 9.999641379E+01
0 9.999607254E+01
1 9.999640437E+01
2 9.999643328E+01
3 9.999634847E+01
4 9.999621716E+01
5 9.999633592E+01
6 9.999662691E+01
7 9.999640136E+01
8 9.999658124E+01
9 9.999635024E+01
10 9.999617473E+01
11 9.999634126E+01
12 9.999633012E+01
13 9.999650638E+01
14 9.999662631E+01
15 9.999648935E+01
16 9.999625468E+01
17 9.999632772E+01
18 9.999617991E+01
19 9.999641203E+01
Result
Thu Jul 17 14:33:16 2014
99.9964
0.00014916
>Exit code: 0
GMH51 20.20
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 1k 1-10 p, repeat
Trial 9.999633590E+01
Trial 2 9.999622223E+01
0 9.999640251E+01
1 9.999638511E+01
2 9.999638605E+01
3 9.999638538E+01
4 9.999637363E+01
5 9.999637178E+01
6 9.999624219E+01
7 9.999628705E+01
8 9.999636234E+01
9 9.999613280E+01
10 9.999620267E+01
11 9.999646790E+01
12 9.999656748E+01
13 9.999633301E+01
14 9.999622343E+01
15 9.999637672E+01
16 9.999640555E+01
17 9.999649081E+01
18 9.999630989E+01
19 9.999636415E+01
Result
Thu Jul 17 14:36:11 2014
99.9964
0.00010128
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E3
Item? ESI 100 1-10 S
Trial 9.999593322E+02
Trial 2 9.999593652E+02
0 9.999594681E+02
1 9.999594293E+02
2 9.999593828E+02
3 9.999594268E+02
4 9.999592960E+02
5 9.999592070E+02
6 9.999593703E+02
7 9.999594132E+02
8 9.999593939E+02
9 9.999593433E+02
10 9.999595731E+02
11 9.999594137E+02
12 9.999593631E+02
13 9.999594964E+02
14 9.999594233E+02
15 9.999592255E+02
16 9.999593256E+02
17 9.999593129E+02
18 9.999592013E+02
19 9.999592254E+02
Result
Thu Jul 17 14:46:46 2014
999.9594
9.9641e-05
>Exit code: 0
GMH51 19.98
pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E3
Item?E SI 100 1-10 S repeat after tea
Trial 9.999591445E+02
Trial 2 9.999589284E+02
0 9.999588058E+02
1 9.999588937E+02
2 9.999588314E+02
3 9.999589860E+02
4 9.999590854E+02
5 9.999590110E+02
6 9.999590416E+02
7 9.999589177E+02
8 9.999589846E+02
9 9.999587510E+02
10 9.999586031E+02
11 9.999587845E+02
12 9.999590555E+02
13 9.999586938E+02
14 9.999590344E+02
15 9.999588307E+02
16 9.999588225E+02
17 9.999586650E+02
18 9.999587548E+02
19 9.999587591E+02
Result
Thu Jul 17 15:49:57 2014
999.9589
0.00014271
>Exit code: 0
GMH51 19.72
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E2
Item? ESI 100 R1
Trial 9.999711930E+01
Trial 2 9.999690324E+01
0 9.999722929E+01
1 9.999712979E+01
2 9.999714256E+01
3 9.999717471E+01
4 9.999689610E+01
5 9.999707279E+01
6 9.999708441E+01
7 9.999701128E+01
8 9.999725316E+01
9 9.999709699E+01
10 9.999697153E+01
11 9.999708017E+01
12 9.999706381E+01
13 9.999702265E+01
14 9.999718259E+01
15 9.999709527E+01
16 9.999691501E+01
17 9.999700524E+01
18 9.999704988E+01
19 9.999709301E+01
Result
Thu Jul 17 15:54:01 2014
99.9971
9.3599e-05
>Exit code: 0
GMH51 19.80
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E1
Item? ESI 100 1-10 P
Trial 9.999553033E+00
Trial 2 9.999557216E+00
0 9.999549003E+00
1 9.999584869E+00
2 9.999556193E+00
3 9.999557771E+00
4 9.999562567E+00
5 9.999546486E+00
6 9.999551302E+00
7 9.999570111E+00
8 9.999564404E+00
9 9.999560671E+00
10 9.999576982E+00
11 9.999560644E+00
12 9.999561386E+00
13 9.999566108E+00
14 9.999562171E+00
15 9.999557495E+00
16 9.999558348E+00
17 9.999548467E+00
18 9.999552776E+00
19 9.999523645E+00
Result
Thu Jul 17 16:00:49 2014
9.9996
1.2468e-05
>Exit code: 0
GMH51 19.89
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E1
Item? ESI 10 R1
Trial 9.999964368E+00
Trial 2 9.999940452E+00
0 9.999942579E+00
1 9.999961826E+00
2 9.999943242E+00
3 9.999926210E+00
4 9.999950548E+00
5 9.999969567E+00
6 9.999945126E+00
7 9.999932827E+00
8 9.999937230E+00
9 9.999955760E+00
10 9.999961918E+00
11 9.999941277E+00
12 9.999954219E+00
13 9.999935945E+00
14 9.999949951E+00
15 9.999942958E+00
16 9.999942906E+00
17 9.999948244E+00
18 9.999950680E+00
19 9.999941952E+00
Result
Thu Jul 17 16:08:50 2014
9.9999
1.05e-05
>Exit code: 0
GMH51 20.06
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 10 R1-10 series
Trial 9.999895003E+01
Trial 2 9.999891885E+01
0 9.999907150E+01
1 9.999911487E+01
2 9.999908492E+01
3 9.999914930E+01
4 9.999904286E+01
5 9.999886680E+01
6 9.999908335E+01
7 9.999895627E+01
8 9.999908324E+01
9 9.999898378E+01
10 9.999896120E+01
11 9.999899297E+01
12 9.999900346E+01
13 9.999883605E+01
14 9.999909381E+01
15 9.999921214E+01
16 9.999924453E+01
17 9.999905933E+01
18 9.999902850E+01
19 9.999906527E+01
Result
Thu Jul 17 16:14:14 2014
99.999
0.00010012
>Exit code: 0
GMH51 20.14
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E5
Item? 100k ishay
Trial 1.000017099E+05
Trial 2 1.000017357E+05
0 1.000017177E+05
1 1.000016704E+05
2 1.000017294E+05
3 1.000017696E+05
4 1.000017417E+05
5 1.000017171E+05
6 1.000017417E+05
7 1.000017366E+05
8 1.000017292E+05
9 1.000017461E+05
10 1.000017244E+05
11 1.000017205E+05
12 1.000017326E+05
13 1.000017152E+05
14 1.000017374E+05
15 1.000017273E+05
16 1.000017183E+05
17 1.000017323E+05
18 1.000017069E+05
19 1.000017188E+05
Result
Thu Jul 17 16:20:40 2014
100001.7267
0.019199
>Exit code: 0
GMH51 20.19
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E5
Item? Vishay 100k repeat
Trial 1.000015411E+05
Trial 2 1.000015510E+05
0 1.000015365E+05
1 1.000015268E+05
2 1.000015517E+05
3 1.000015355E+05
4 1.000015254E+05
5 1.000015401E+05
6 1.000015240E+05
7 1.000015304E+05
8 1.000015454E+05
9 1.000015580E+05
10 1.000015524E+05
11 1.000015321E+05
12 1.000015344E+05
13 1.000015440E+05
14 1.000015456E+05
15 1.000015365E+05
16 1.000015337E+05
17 1.000015335E+05
18 1.000015451E+05
19 1.000015348E+05
Result
Thu Jul 17 16:28:04 2014
100001.5383
0.0092966
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E4
Item? 10k Vishay
Trial 1.000027757E+04
Trial 2 1.000027674E+04
0 1.000027601E+04
1 1.000027719E+04
2 1.000027778E+04
3 1.000027750E+04
4 1.000027408E+04
5 1.000027672E+04
6 1.000027695E+04
7 1.000027494E+04
8 1.000027555E+04
9 1.000027679E+04
10 1.000027615E+04
11 1.000028002E+04
12 1.000027514E+04
13 1.000027444E+04
14 1.000027457E+04
15 1.000027570E+04
16 1.000027637E+04
17 1.000027354E+04
18 1.000027488E+04
19 1.000027540E+04
Result
Thu Jul 17 16:31:49 2014
10000.276
0.0015076
>Exit code: 0
GMH51 20.17
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E2
Item? 100 Vishay
Trial 1.000131660E+02
Trial 2 1.000130719E+02
0 1.000131073E+02
1 1.000134502E+02
2 1.000132129E+02
3 1.000131490E+02
4 1.000130690E+02
5 1.000132104E+02
6 1.000130966E+02
7 1.000131675E+02
8 1.000132009E+02
9 1.000131551E+02
10 1.000133579E+02
11 1.000132767E+02
12 1.000130706E+02
13 1.000131544E+02
14 1.000131324E+02
15 1.000130392E+02
16 1.000131840E+02
17 1.000133201E+02
18 1.000129175E+02
19 1.000130950E+02
Result
Thu Jul 17 16:35:29 2014
100.0132
0.00011932
>Exit code: 0
GMH 51 20.11
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 50.0400462E+00 Hz
Temperature = 33.4
ACAL ALL starting
ACAL finished
Time taken = 16.0772499998 minutes
Temperature = 33.4
Thu Jul 31 13:57:24 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? Vishay S106D 1M
Trial 9.999963193E+05
Trial 2 9.999967197E+05
0 9.999965833E+05
1 9.999967235E+05
2 9.999966186E+05
3 9.999964135E+05
4 9.999964585E+05
5 9.999964929E+05
6 9.999966365E+05
7 9.999964942E+05
8 9.999965459E+05
9 9.999962713E+05
10 9.999964026E+05
11 9.999962692E+05
12 9.999963153E+05
13 9.999962428E+05
14 9.999961601E+05
15 9.999962719E+05
16 9.999963667E+05
17 9.999964213E+05
18 9.999962916E+05
19 9.999963838E+05
Result
Thu Jul 31 14:32:49 2014
999996.4182
0.15059
>Exit code: 0
GMH51 20.01
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? Vishay S106D 1M
Trial 9.999965304E+05
Trial 2 9.999966892E+05
0 9.999970253E+05
1 9.999971704E+05
2 9.999969313E+05
3 9.999971360E+05
4 9.999973011E+05
5 9.999971888E+05
6 9.999971570E+05
7 9.999971115E+05
8 9.999971711E+05
9 9.999968658E+05
10 9.999970695E+05
11 9.999969482E+05
12 9.999971898E+05
13 9.999969724E+05
14 9.999971814E+05
15 9.999968986E+05
16 9.999970115E+05
17 9.999970907E+05
18 9.999971375E+05
19 9.999971830E+05
Result
Thu Jul 31 14:37:33 2014
999997.087
0.11704
>Exit code: 0
GMH51 19.95
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E4
Item? SR104 temperature sensor
Trial 9.969186052E+03
Trial 2 9.969185708E+03
0 9.969184643E+03
1 9.969185981E+03
2 9.969186238E+03
3 9.969186144E+03
4 9.969188423E+03
5 9.969189547E+03
6 9.969187924E+03
7 9.969188497E+03
8 9.969186384E+03
9 9.969187898E+03
10 9.969188013E+03
11 9.969187596E+03
12 9.969185831E+03
13 9.969189290E+03
14 9.969186741E+03
15 9.969188574E+03
16 9.969190410E+03
17 9.969191168E+03
18 9.969189748E+03
19 9.969190179E+03
Result
Thu Jul 31 15:40:36 2014
9969.188
0.0017682
>Exit code: 0
GMH51 19.81
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E4
Item? SR104 10k
Trial 9.999970312E+03
Trial 2 9.999971031E+03
0 9.999970112E+03
1 9.999972805E+03
2 9.999968950E+03
3 9.999968510E+03
4 9.999968676E+03
5 9.999969939E+03
6 9.999970711E+03
7 9.999971186E+03
8 9.999970540E+03
9 9.999970725E+03
10 9.999968956E+03
11 9.999969770E+03
12 9.999970224E+03
13 9.999968849E+03
14 9.999968916E+03
15 9.999969152E+03
16 9.999971083E+03
17 9.999967072E+03
18 9.999969479E+03
19 9.999969715E+03
Result
Thu Jul 31 15:49:22 2014
9999.9698
0.0012382
>Exit code: 0
GMH51 19.82
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E4
Item? ESI 100k 1-10 p
Trial 1.000005854E+04
Trial 2 1.000005519E+04
0 1.000005444E+04
1 1.000005470E+04
2 1.000005663E+04
3 1.000005504E+04
4 1.000005708E+04
5 1.000005620E+04
6 1.000005344E+04
7 1.000005340E+04
8 1.000005378E+04
9 1.000005322E+04
10 1.000005472E+04
11 1.000005645E+04
12 1.000005765E+04
13 1.000005537E+04
14 1.000005472E+04
15 1.000005644E+04
16 1.000005663E+04
17 1.000005621E+04
18 1.000005901E+04
19 1.000005702E+04
Result
Thu Jul 31 15:55:34 2014
10000.0556
0.0015706
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E4
Item? ESI 100k/step, 1-10 p
Trial 1.000005337E+04
Trial 2 1.000005397E+04
0 1.000005361E+04
1 1.000005177E+04
2 1.000005296E+04
3 1.000005285E+04
4 1.000005567E+04
5 1.000005313E+04
6 1.000005635E+04
7 1.000005177E+04
8 1.000005225E+04
9 1.000005360E+04
10 1.000005273E+04
11 1.000005338E+04
12 1.000005544E+04
13 1.000005413E+04
14 1.000005072E+04
15 1.000005368E+04
16 1.000005260E+04
17 1.000005450E+04
18 1.000005296E+04
19 1.000005367E+04
Result
Thu Jul 31 15:58:09 2014
10000.0534
0.00137
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E6
Item? ESI 100k, 1-10 s
Trial 1.000003360E+06
Trial 2 1.000004248E+06
0 1.000003722E+06
1 1.000003822E+06
2 1.000003732E+06
3 1.000003887E+06
4 1.000004086E+06
5 1.000004042E+06
6 1.000004148E+06
7 1.000004002E+06
8 1.000004200E+06
9 1.000004054E+06
10 1.000004334E+06
11 1.000003929E+06
12 1.000004077E+06
13 1.000004348E+06
14 1.000004306E+06
15 1.000004498E+06
16 1.000004333E+06
17 1.000004624E+06
18 1.000004432E+06
19 1.000004480E+06
Result
Thu Jul 31 16:07:25 2014
1000004.1528
0.26189
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E6
Item? ESI 100k/step, 1-10 series
Trial 1.000004733E+06
Trial 2 1.000004482E+06
0 1.000004688E+06
1 1.000005104E+06
2 1.000004909E+06
3 1.000004584E+06
4 1.000004707E+06
5 1.000004897E+06
6 1.000005083E+06
7 1.000004882E+06
8 1.000004765E+06
9 1.000004879E+06
10 1.000004920E+06
11 1.000005022E+06
12 1.000004656E+06
13 1.000004950E+06
14 1.000004856E+06
15 1.000004739E+06
16 1.000004974E+06
17 1.000004938E+06
18 1.000004839E+06
19 1.000004959E+06
Result
Thu Jul 31 16:09:22 2014
1000004.8675
0.14039
>Exit code: 0
GMH51 20.31

## ESI_check
ESI box summary for comparison with MI bridge values
see Y:\Ongoing\OHM\ESI_Boxes\BuildupESIboxes2014.xls, ~10 Feb 2014
Box/step | MI(ppm) | HP3458
1 | -73.6095
10 | -6.4394 | -5.3873 | 1.0521
100 | -32.076 | -37.0145 | -4.9384
1000 | -31.7342 | -32.3191 | -0.58489
10000 | 17.9648
100000 | 8.7165 | 6.6 | -2.1165
Note that humidity/temperature/time will account for the differences.
View as a 'common sense' check only.