## Commentary
Start with spreadsheet S21740 from the July 2009 calibration.
Remove all repeats from Results sheet
Temperature correction updated for GMH51
Measured frequency included, but not really necessary
Results transcribed from AI VII
Room data logger information for Primary Lab insered in 'Room data logger' sheet
Replace 'Universal Bridge' page with the same from S21942b, change zeros block and reference.
Insert uncertainty sheet from S21942b, with subsequent modifications but not to the UB uncertainty
The copper temperature coefficient is identical for all, but I have left the parameter table that would allow variation by component.
Add worksheet 'Histroical' to give a cleaner graphical view of previous calibrations, but some of the data is still picked up off the 'Previous History' worksheet.
Resolution term for L in uncertainty, cell B40, has been made consistent with the CMC formula.
On results page there is an analysis of HP4284 measurements designed to cross-check that the 600 microH result is reasonable (oscillation problem).
Restart using S21952, 2014, but update to interact with Python for analysis.
Enter lab book readings in 'pyUBreadings' worksheet and check the temperautre coefficients entered for each line.
Run sullivan.py to create a set of results in sullivan_Results.xlsx, being sure to use the latest ubdict_date.csv for the universal bridge calibration.
Paste these results into the uncoloured part of the pyUBresults worksheet.
Final report appears in the pyReport worksheet.

## pyUBreadings
Template for impedances
1 | 2 |  | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Impedance | sullivan |  | 4: Client; 10: Enter, say, 5 ppm as 5e-6
Date from |  |  | 4: Job number; 10: 1 sigma uncertainties from specifications
to |  |  | 4: S Number
|  |  | Data from lab book |  | check! | check! | Data from specifications and other sources
Item | Nominal Frequency | Range | xdial | rdial | Frequency | Temperature | u | df | Temp coeff L | u | df | Temp coeff R | u | df | Other | u | df
Hz | bar: 5 to 6
coax zero | 1591.5 | 3Z | -1 | -6 | 1593.07 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A1 | 1591.5 | 3Z | 99380 | 569169 | 1593.33 | 19.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A2 | 1591.5 | 3Z | 198010 | 838356 | 1593.46 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A3 | 1591.5 | 3Z | 296879 | 1066481 | 1593.67 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A4 | 1591.5 | 3Z | 394832 | 1265232 | 1593.74 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A5 | 1591.5 | 3Z | 493210 | 1442337 | 1592.85 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A6 | 1591.5 | 3Z | 591464 | 1611164 | 1592.97 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A7 | 1591.5 | 3Z | 689154 | 1775693 | 1593.08 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A8 | 1591.5 | 3Z | 786400 | 1927498 | 1593.25 | 19.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A9 | 1591.5 | 3Z | 883362 | 2073610 | 1593.16 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
A10 | 1591.5 | 3Z | 980451 | 2222585 | 1592.88 | 19.86 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B1 | 1591.5 | 2Z | 99118 | 1166562 | 1594.81 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B2 | 1591.5 | 2Z | 198956 | 1739697 | 1594.95 | 19.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B3 | 1591.5 | 2Z | 298184 | 2207103 | 1595.08 | 19.75 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B4 | 1591.5 | 2Z | 396438 | 2792679 | 1595.22 | 19.74 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B5 | 1591.5 | 2Z | 494367 | 3176412 | 1595.32 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B6 | 1591.5 | 2Z | 593385 | 3538125 | 1594.66 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B7 | 1591.5 | 2Z | 691195 | 3896072 | 1594.9 | 19.78 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B8 | 1591.5 | 2Z | 788963 | 4199748 | 1595.1 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B9 | 1591.5 | 2Z | 887174 | 4517998 | 1595.29 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
B10 | 1591.5 | 2Z | 984585 | 4826330 | 1595.54 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
coax zero | 1591.5 | 2Z | -3 | -25 | 1594.71 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10
19.7427
0.043663

## pyUBresults
Item | Label | Range | xDial | rDial | frequency | temperature | utemp | dftemp | alphax | ualphax | dfalphax | alphar | ualphar | dfalphr | x | ux | kx | r | ur | kr | cmcx | cmcr | ux/cmcx | ur/cmcr | label | multiplier for R | R | uR | k | multiplier for L | L | uL | k | nom Freq
coax zero | 1591.5 | 3Z | -1 | -6 | 1593.07 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2554e-08 | 0.00012554 | 0 | 0 | coax zero1591.5 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5
A1 | 1591.5 | 3Z | 99380 | 569169 | 1593.33 | 19.73 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00099402 | 6.3213e-08 | 1.989 | 5.686 | 0.033719 | 2.0834 | 1.9586e-08 | 0.00017885 | 3.2274 | 188.5328 | A11591.5 | 1 | 5.686 | 0.033719 | 2.0834 | 1000000 | 994.0168 | 0.063213 | 1.989 | 1591.5
A2 | 1591.5 | 3Z | 198010 | 838356 | 1593.46 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0019805 | 9.8535e-08 | 2.0003 | 8.3745 | 0.049591 | 2.0845 | 3.157e-08 | 0.00026221 | 3.1212 | 189.1228 | A21591.5 | 1 | 8.3745 | 0.049591 | 2.0845 | 1000000 | 1980.5213 | 0.098535 | 2.0003 | 1591.5
A3 | 1591.5 | 3Z | 296879 | 1066481 | 1593.67 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0029694 | 1.3829e-07 | 2.0102 | 10.6533 | 0.063054 | 2.0848 | 4.4715e-08 | 0.00035527 | 3.0927 | 177.4793 | A31591.5 | 1 | 10.6533 | 0.063054 | 2.0848 | 1000000 | 2969.4166 | 0.13829 | 2.0102 | 1591.5
A4 | 1591.5 | 3Z | 394832 | 1265232 | 1593.74 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0039491 | 1.7901e-07 | 2.0157 | 12.6381 | 0.074789 | 2.085 | 5.8089e-08 | 0.00045065 | 3.0817 | 165.9583 | A41591.5 | 1 | 12.6381 | 0.074789 | 2.085 | 1000000 | 3949.1493 | 0.17901 | 2.0157 | 1591.5
A5 | 1591.5 | 3Z | 493210 | 1442337 | 1592.85 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0049331 | 2.2044e-07 | 2.0191 | 14.4077 | 0.085248 | 2.0851 | 7.1663e-08 | 0.00054764 | 3.0761 | 155.6651 | A51591.5 | 1 | 14.4077 | 0.085248 | 2.0851 | 1000000 | 4933.1344 | 0.22044 | 2.0191 | 1591.5
A6 | 1591.5 | 3Z | 591464 | 1611164 | 1592.97 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0059159 | 2.6214e-07 | 2.0212 | 16.0935 | 0.095221 | 2.0851 | 8.5304e-08 | 0.00064547 | 3.073 | 147.5225 | A61591.5 | 1 | 16.0935 | 0.095221 | 2.0851 | 1000000 | 5915.8773 | 0.26214 | 2.0212 | 1591.5
A7 | 1591.5 | 3Z | 689154 | 1775693 | 1593.08 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.006893 | 3.038e-07 | 2.0226 | 17.737 | 0.10494 | 2.0851 | 9.8922e-08 | 0.00074344 | 3.0711 | 141.1545 | A71591.5 | 1 | 17.737 | 0.10494 | 2.0851 | 1000000 | 6892.98 | 0.3038 | 2.0226 | 1591.5
A8 | 1591.5 | 3Z | 786400 | 1927498 | 1593.25 | 19.69 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0078656 | 3.4534e-07 | 2.0237 | 19.2525 | 0.11391 | 2.0852 | 1.125e-07 | 0.000841 | 3.0697 | 135.4465 | A81591.5 | 1 | 19.2525 | 0.11391 | 2.0852 | 1000000 | 7865.6403 | 0.34534 | 2.0237 | 1591.5
A9 | 1591.5 | 3Z | 883362 | 2073610 | 1593.16 | 19.7 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0088355 | 3.8682e-07 | 2.0245 | 20.7128 | 0.12254 | 2.0852 | 1.2605e-07 | 0.00093846 | 3.0688 | 130.5772 | A91591.5 | 1 | 20.7128 | 0.12254 | 2.0852 | 1000000 | 8835.463 | 0.38682 | 2.0245 | 1591.5
A10 | 1591.5 | 3Z | 980451 | 2222585 | 1592.88 | 19.86 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.0098066 | 4.2841e-07 | 2.0251 | 22.2151 | 0.13133 | 2.0853 | 1.3965e-07 | 0.0010365 | 3.0679 | 126.7045 | A101591.5 | 1 | 22.2151 | 0.13133 | 2.0853 | 1000000 | 9806.5859 | 0.42841 | 2.0251 | 1591.5
B1 | 1591.5 | 2Z | 99118 | 1166562 | 1594.81 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 9.9142e-05 | 7.2135e-09 | 1.9977 | 1.1656 | 0.0068951 | 2.085 | 2.2046e-09 | 2.2854e-05 | 3.2721 | 301.6977 | B11591.5 | 1 | 1.1656 | 0.0068951 | 2.085 | 1000000 | 99.1417 | 0.0072135 | 1.9977 | 1591.5
B2 | 1591.5 | 2Z | 198956 | 1739697 | 1594.95 | 19.76 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.000199 | 1.1164e-08 | 2.0028 | 1.7382 | 0.010279 | 2.0853 | 3.5158e-09 | 3.386e-05 | 3.1755 | 303.5668 | B21591.5 | 1 | 1.7382 | 0.010279 | 2.0853 | 1000000 | 199.0005 | 0.011164 | 2.0028 | 1591.5
B3 | 1591.5 | 2Z | 298184 | 2207103 | 1595.08 | 19.75 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00029825 | 1.5328e-08 | 2.0066 | 2.2051 | 0.013039 | 2.0854 | 4.8866e-09 | 4.474e-05 | 3.1367 | 291.4462 | B31591.5 | 1 | 2.2051 | 0.013039 | 2.0854 | 1000000 | 298.2492 | 0.015328 | 2.0066 | 1591.5
B4 | 1591.5 | 2Z | 396438 | 2792679 | 1595.22 | 19.74 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00039652 | 1.9807e-08 | 2.0094 | 2.79 | 0.016498 | 2.0854 | 6.3395e-09 | 5.7082e-05 | 3.1244 | 289.0192 | B41591.5 | 1 | 2.79 | 0.016498 | 2.0854 | 1000000 | 396.5236 | 0.019807 | 2.0094 | 1591.5
B5 | 1591.5 | 2Z | 494367 | 3176412 | 1595.32 | 19.71 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00049447 | 2.4016e-08 | 2.011 | 3.173 | 0.018765 | 2.0854 | 7.7192e-09 | 6.7676e-05 | 3.1112 | 277.2736 | B51591.5 | 1 | 3.173 | 0.018765 | 2.0854 | 1000000 | 494.4727 | 0.024016 | 2.011 | 1591.5
B6 | 1591.5 | 2Z | 593385 | 3538125 | 1594.66 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00059351 | 2.8287e-08 | 2.0124 | 3.5351 | 0.0209 | 2.0855 | 9.1179e-09 | 7.8319e-05 | 3.1024 | 266.8602 | B61591.5 | 1 | 3.5351 | 0.0209 | 2.0855 | 1000000 | 593.512 | 0.028287 | 2.0124 | 1591.5
B7 | 1591.5 | 2Z | 691195 | 3896072 | 1594.9 | 19.78 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00069134 | 3.2543e-08 | 2.0135 | 3.8929 | 0.023014 | 2.0855 | 1.0509e-08 | 8.8952e-05 | 3.0968 | 258.7269 | B71591.5 | 1 | 3.8929 | 0.023014 | 2.0855 | 1000000 | 691.3425 | 0.032543 | 2.0135 | 1591.5
B8 | 1591.5 | 2Z | 788963 | 4199748 | 1595.1 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00078913 | 3.675e-08 | 2.0145 | 4.1965 | 0.024808 | 2.0855 | 1.1886e-08 | 9.9207e-05 | 3.092 | 250.06 | B81591.5 | 1 | 4.1965 | 0.024808 | 2.0855 | 1000000 | 789.1311 | 0.03675 | 2.0145 | 1591.5
B9 | 1591.5 | 2Z | 887174 | 4517998 | 1595.29 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00088736 | 4.1013e-08 | 2.0154 | 4.5146 | 0.026688 | 2.0855 | 1.3279e-08 | 0.00010969 | 3.0887 | 243.2912 | B91591.5 | 1 | 4.5146 | 0.026688 | 2.0855 | 1000000 | 887.3627 | 0.041013 | 2.0154 | 1591.5
B10 | 1591.5 | 2Z | 984585 | 4826330 | 1595.54 | 19.79 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0.00098479 | 4.5247e-08 | 2.0161 | 4.8226 | 0.028509 | 2.0855 | 1.4661e-08 | 0.00012009 | 3.0862 | 237.3956 | B101591.5 | 1 | 4.8226 | 0.028509 | 2.0855 | 1000000 | 984.7941 | 0.045247 | 2.0161 | 1591.5
coax zero | 1591.5 | 2Z | -3 | -25 | 1594.71 | 19.77 | 0.5 | 10 | 2e-05 | 1e-06 | 10 | 0.004 | 0.0002 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 1.2541e-09 | 1.2541e-05 | 0 | 0 | coax zero1591.5 | 1 | 0 | 0 | 0 | 1000000 | 0 | 0 | 0 | 1591.5

## environment
From room logging system
| 20.6244 | 48.517
| 0.04585 | 0.83038
2020-11-09 | 20.7 | 50.5
2020-11-09 00:05 | 20.6 | 50.6
2020-11-09 00:10 | 20.6 | 50.6
2020-11-09 00:15 | 20.6 | 50.5
2020-11-09 00:20 | 20.6 | 50.4
2020-11-09 00:25 | 20.7 | 50.3
2020-11-09 00:30 | 20.7 | 50.3
2020-11-09 00:35 | 20.7 | 50.3
2020-11-09 00:40 | 20.7 | 50.3
2020-11-09 00:45 | 20.7 | 50.2
2020-11-09 00:50 | 20.7 | 50.3
2020-11-09 00:55 | 20.7 | 50.2
2020-11-09 01:00 | 20.7 | 50.3
2020-11-09 01:05 | 20.7 | 50.2
2020-11-09 01:10 | 20.7 | 50.1
2020-11-09 01:15 | 20.7 | 50.1
2020-11-09 01:20 | 20.6 | 50.1
2020-11-09 01:25 | 20.7 | 50
2020-11-09 01:30 | 20.6 | 50.1
2020-11-09 01:35 | 20.7 | 50
2020-11-09 01:40 | 20.7 | 49.9
2020-11-09 01:45 | 20.7 | 50
2020-11-09 01:50 | 20.7 | 49.9
2020-11-09 01:55 | 20.7 | 49.9
2020-11-09 02:00 | 20.6 | 50
2020-11-09 02:05 | 20.6 | 50
2020-11-09 02:10 | 20.7 | 49.9
2020-11-09 02:15 | 20.7 | 49.9
2020-11-09 02:20 | 20.6 | 49.9
2020-11-09 02:25 | 20.6 | 49.9
2020-11-09 02:30 | 20.7 | 49.8
2020-11-09 02:35 | 20.7 | 49.9
2020-11-09 02:40 | 20.7 | 49.9
2020-11-09 02:45 | 20.6 | 49.9
2020-11-09 02:50 | 20.7 | 49.8
2020-11-09 02:55 | 20.6 | 49.9
2020-11-09 03:00 | 20.7 | 49.8
2020-11-09 03:05 | 20.7 | 49.8
2020-11-09 03:10 | 20.7 | 49.7
2020-11-09 03:15 | 20.7 | 49.7
2020-11-09 03:20 | 20.7 | 49.7
2020-11-09 03:25 | 20.7 | 49.6
2020-11-09 03:30 | 20.7 | 49.5
2020-11-09 03:35 | 20.6 | 49.7
2020-11-09 03:40 | 20.6 | 49.6
2020-11-09 03:45 | 20.6 | 49.5
2020-11-09 03:50 | 20.6 | 49.6
2020-11-09 03:55 | 20.6 | 49.5
2020-11-09 04:00 | 20.6 | 49.5
2020-11-09 04:05 | 20.6 | 49.4
2020-11-09 04:10 | 20.7 | 49.2
2020-11-09 04:15 | 20.7 | 49.2
2020-11-09 04:20 | 20.6 | 49.2
2020-11-09 04:25 | 20.6 | 49.2
2020-11-09 04:30 | 20.6 | 49.2
2020-11-09 04:35 | 20.7 | 49.2
2020-11-09 04:40 | 20.7 | 49.1
2020-11-09 04:45 | 20.7 | 49
2020-11-09 04:50 | 20.6 | 49
2020-11-09 04:55 | 20.7 | 48.9
2020-11-09 05:00 | 20.6 | 48.8
2020-11-09 05:05 | 20.7 | 48.7
2020-11-09 05:10 | 20.6 | 48.7
2020-11-09 05:15 | 20.6 | 48.7
2020-11-09 05:20 | 20.6 | 48.6
2020-11-09 05:25 | 20.6 | 48.6
2020-11-09 05:30 | 20.7 | 48.6
2020-11-09 05:35 | 20.7 | 48.6
2020-11-09 05:40 | 20.7 | 48.6
2020-11-09 05:45 | 20.7 | 48.6
2020-11-09 05:50 | 20.6 | 48.6
2020-11-09 05:55 | 20.7 | 48.6
2020-11-09 06:00 | 20.6 | 48.6
2020-11-09 06:05 | 20.7 | 48.5
2020-11-09 06:10 | 20.7 | 48.4
2020-11-09 06:15 | 20.7 | 48.4
2020-11-09 06:20 | 20.6 | 48.5
2020-11-09 06:25 | 20.6 | 48.5
2020-11-09 06:30 | 20.6 | 48.6
2020-11-09 06:35 | 20.6 | 48.5
2020-11-09 06:40 | 20.6 | 48.4
2020-11-09 06:45 | 20.6 | 48.4
2020-11-09 06:50 | 20.6 | 48.4
2020-11-09 06:55 | 20.6 | 48.2
2020-11-09 07:00 | 20.6 | 48.2
2020-11-09 07:05 | 20.6 | 48.2
2020-11-09 07:10 | 20.6 | 48.1
2020-11-09 07:15 | 20.6 | 48.1
2020-11-09 07:20 | 20.6 | 48.1
2020-11-09 07:25 | 20.6 | 48.1
2020-11-09 07:30 | 20.6 | 48.2
2020-11-09 07:35 | 20.6 | 48
2020-11-09 07:40 | 20.7 | 47.9
2020-11-09 07:45 | 20.7 | 48
2020-11-09 07:50 | 20.7 | 47.9
2020-11-09 07:55 | 20.7 | 47.9
2020-11-09 08:00 | 20.7 | 48
2020-11-09 08:05 | 20.7 | 48
2020-11-09 08:10 | 20.7 | 48
2020-11-09 08:15 | 20.7 | 47.9
2020-11-09 08:20 | 20.7 | 48
2020-11-09 08:25 | 20.7 | 47.9
2020-11-09 08:30 | 20.8 | 47.8
2020-11-09 08:35 | 20.7 | 48
2020-11-09 08:40 | 20.7 | 48.2
2020-11-09 08:45 | 20.6 | 48.3
2020-11-09 08:50 | 20.6 | 48.4
2020-11-09 08:55 | 20.7 | 48.4
2020-11-09 09:00 | 20.7 | 48.4
2020-11-09 09:05 | 20.7 | 48.5
2020-11-09 09:10 | 20.7 | 48.5
2020-11-09 09:15 | 20.6 | 48.6
2020-11-09 09:20 | 20.7 | 48.5
2020-11-09 09:25 | 20.7 | 48.5
2020-11-09 09:30 | 20.7 | 48.6
2020-11-09 09:35 | 20.6 | 48.9
2020-11-09 09:40 | 20.7 | 48.8
2020-11-09 09:45 | 20.6 | 48.8
2020-11-09 09:50 | 20.6 | 48.8
2020-11-09 09:55 | 20.6 | 48.7
2020-11-09 10:00 | 20.6 | 48.7
2020-11-09 10:05 | 20.6 | 48.6
2020-11-09 10:10 | 20.6 | 48.5
2020-11-09 10:15 | 20.6 | 48.4
2020-11-09 10:20 | 20.6 | 48.4
2020-11-09 10:25 | 20.6 | 48.2
2020-11-09 10:30 | 20.6 | 48.2
2020-11-09 10:35 | 20.6 | 48.2
2020-11-09 10:40 | 20.6 | 48.4
2020-11-09 10:45 | 20.6 | 48.4
2020-11-09 10:50 | 20.6 | 48.3
2020-11-09 10:55 | 20.6 | 48.4
2020-11-09 11:00 | 20.6 | 48.3
2020-11-09 11:05 | 20.7 | 48.2
2020-11-09 11:10 | 20.6 | 48.2
2020-11-09 11:15 | 20.6 | 48.2
2020-11-09 11:20 | 20.7 | 48
2020-11-09 11:25 | 20.6 | 48.2
2020-11-09 11:30 | 20.6 | 48.3
2020-11-09 11:35 | 20.6 | 48.4
2020-11-09 11:40 | 20.6 | 48.4
2020-11-09 11:45 | 20.5 | 48.5
2020-11-09 11:50 | 20.6 | 48.4
2020-11-09 11:55 | 20.6 | 48.1
2020-11-09 12:00 | 20.6 | 48
2020-11-09 12:05 | 20.6 | 48
2020-11-09 12:10 | 20.6 | 47.9
2020-11-09 12:15 | 20.7 | 47.7
2020-11-09 12:20 | 20.7 | 47.6
2020-11-09 12:25 | 20.6 | 47.7
2020-11-09 12:30 | 20.6 | 47.7
2020-11-09 12:35 | 20.6 | 47.8
2020-11-09 12:40 | 20.6 | 47.9
2020-11-09 12:45 | 20.6 | 47.7
2020-11-09 12:50 | 20.6 | 47.7
2020-11-09 12:55 | 20.7 | 47.6
2020-11-09 13:00 | 20.6 | 47.7
2020-11-09 13:05 | 20.6 | 47.6
2020-11-09 13:10 | 20.6 | 47.6
2020-11-09 13:15 | 20.6 | 47.7
2020-11-09 13:20 | 20.6 | 47.7
2020-11-09 13:25 | 20.6 | 47.7
2020-11-09 13:30 | 20.6 | 47.5
2020-11-09 13:35 | 20.6 | 47.3
2020-11-09 13:40 | 20.6 | 47.1
2020-11-09 13:45 | 20.6 | 46.9
2020-11-09 13:50 | 20.6 | 46.5
2020-11-09 13:55 | 20.6 | 46.3
2020-11-09 14:00 | 20.6 | 46.2
2020-11-09 14:05 | 20.6 | 45.9
2020-11-09 14:10 | 20.6 | 45.9
2020-11-09 14:15 | 20.6 | 45.6
2020-11-09 14:20 | 20.6 | 45.7
2020-11-09 14:25 | 20.7 | 45.7
2020-11-09 14:30 | 20.6 | 45.9
2020-11-09 14:35 | 20.6 | 46.2
2020-11-09 14:40 | 20.6 | 46.4
2020-11-09 14:45 | 20.6 | 46.6
2020-11-09 14:50 | 20.6 | 46.8
2020-11-09 14:55 | 20.5 | 46.7
2020-11-09 15:00 | 20.6 | 46.6
2020-11-09 15:05 | 20.6 | 46.6
2020-11-09 15:10 | 20.7 | 46.5
2020-11-09 15:15 | 20.6 | 46.6
2020-11-09 15:20 | 20.6 | 46.7
2020-11-09 15:25 | 20.6 | 46.7
2020-11-09 15:30 | 20.5 | 46.9
2020-11-09 15:35 | 20.6 | 46.8
2020-11-09 15:40 | 20.6 | 46.8
2020-11-09 15:45 | 20.6 | 46.8
2020-11-09 15:50 | 20.6 | 46.8
2020-11-09 15:55 | 20.6 | 46.9
2020-11-09 16:00 | 20.6 | 47.2
2020-11-09 16:05 | 20.6 | 47.2
2020-11-09 16:10 | 20.7 | 47.3
2020-11-09 16:15 | 20.6 | 47.4
2020-11-09 16:20 | 20.6 | 47.5
2020-11-09 16:25 | 20.6 | 47.5
2020-11-09 16:30 | 20.6 | 47.7
2020-11-09 16:35 | 20.5 | 47.8
2020-11-09 16:40 | 20.6 | 47.7
2020-11-09 16:45 | 20.6 | 47.7
2020-11-09 16:50 | 20.6 | 47.8
2020-11-09 16:55 | 20.6 | 47.9
2020-11-09 17:00 | 20.6 | 47.9
2020-11-09 17:05 | 20.6 | 48
2020-11-09 17:10 | 20.6 | 48
2020-11-09 17:15 | 20.6 | 47.9
2020-11-09 17:20 | 20.6 | 47.9
2020-11-09 17:25 | 20.6 | 47.9
2020-11-09 17:30 | 20.7 | 48
2020-11-09 17:35 | 20.6 | 48
2020-11-09 17:40 | 20.6 | 48
2020-11-09 17:45 | 20.6 | 47.8
2020-11-09 17:50 | 20.6 | 47.8
2020-11-09 17:55 | 20.6 | 47.7
2020-11-09 18:00 | 20.6 | 47.7
2020-11-09 18:05 | 20.6 | 47.5
2020-11-09 18:10 | 20.6 | 47.7
2020-11-09 18:15 | 20.6 | 47.6
2020-11-09 18:20 | 20.6 | 47.6
2020-11-09 18:25 | 20.6 | 47.6
2020-11-09 18:30 | 20.6 | 47.6
2020-11-09 18:35 | 20.6 | 47.5
2020-11-09 18:40 | 20.6 | 47.5
2020-11-09 18:45 | 20.6 | 47.6
2020-11-09 18:50 | 20.6 | 47.6
2020-11-09 18:55 | 20.6 | 47.6
2020-11-09 19:00 | 20.6 | 47.7
2020-11-09 19:05 | 20.6 | 47.7
2020-11-09 19:10 | 20.6 | 47.6
2020-11-09 19:15 | 20.6 | 47.7
2020-11-09 19:20 | 20.6 | 47.8
2020-11-09 19:25 | 20.6 | 47.7
2020-11-09 19:30 | 20.6 | 47.7
2020-11-09 19:35 | 20.6 | 47.7
2020-11-09 19:40 | 20.6 | 47.7
2020-11-09 19:45 | 20.7 | 47.6
2020-11-09 19:50 | 20.6 | 47.7
2020-11-09 19:55 | 20.6 | 47.7
2020-11-09 20:00 | 20.6 | 47.7
2020-11-09 20:05 | 20.6 | 47.5
2020-11-09 20:10 | 20.6 | 47.6
2020-11-09 20:15 | 20.6 | 47.7
2020-11-09 20:20 | 20.6 | 47.6
2020-11-09 20:25 | 20.6 | 47.6
2020-11-09 20:30 | 20.6 | 47.6
2020-11-09 20:35 | 20.6 | 47.5
2020-11-09 20:40 | 20.6 | 47.4
2020-11-09 20:45 | 20.6 | 47.4
2020-11-09 20:50 | 20.6 | 47.4
2020-11-09 20:55 | 20.6 | 47.4
2020-11-09 21:00 | 20.6 | 47.4
2020-11-09 21:05 | 20.6 | 47.5
2020-11-09 21:10 | 20.6 | 47.5
2020-11-09 21:15 | 20.6 | 47.6
2020-11-09 21:20 | 20.6 | 47.5
2020-11-09 21:25 | 20.7 | 47.5
2020-11-09 21:30 | 20.6 | 47.7
2020-11-09 21:35 | 20.6 | 47.6
2020-11-09 21:40 | 20.6 | 47.7
2020-11-09 21:45 | 20.6 | 47.8
2020-11-09 21:50 | 20.6 | 47.7
2020-11-09 21:55 | 20.6 | 47.7
2020-11-09 22:00 | 20.7 | 47.7
2020-11-09 22:05 | 20.6 | 47.8
2020-11-09 22:10 | 20.6 | 47.8
2020-11-09 22:15 | 20.6 | 47.8
2020-11-09 22:20 | 20.6 | 47.9
2020-11-09 22:25 | 20.6 | 47.9
2020-11-09 22:30 | 20.6 | 47.9
2020-11-09 22:35 | 20.6 | 47.9
2020-11-09 22:40 | 20.6 | 47.9
2020-11-09 22:45 | 20.6 | 47.9
2020-11-09 22:50 | 20.7 | 47.9
2020-11-09 22:55 | 20.6 | 48
2020-11-09 23:00 | 20.6 | 48
2020-11-09 23:05 | 20.7 | 47.9
2020-11-09 23:10 | 20.6 | 47.9
2020-11-09 23:15 | 20.6 | 47.9
2020-11-09 23:20 | 20.6 | 48
2020-11-09 23:25 | 20.7 | 48
2020-11-09 23:30 | 20.6 | 48
2020-11-09 23:35 | 20.7 | 48
2020-11-09 23:40 | 20.7 | 48
2020-11-09 23:45 | 20.6 | 48.1
2020-11-09 23:50 | 20.7 | 48.1
2020-11-09 23:55 | 20.6 | 48.2
2020-11-10 | 20.6 | 48.1
2020-11-10 00:05 | 20.7 | 48.2
2020-11-10 00:10 | 20.7 | 48.2
2020-11-10 00:15 | 20.6 | 48.3
2020-11-10 00:20 | 20.7 | 48.2
2020-11-10 00:25 | 20.7 | 48.2
2020-11-10 00:30 | 20.7 | 48.2
2020-11-10 00:35 | 20.7 | 48.3
2020-11-10 00:40 | 20.6 | 48.3
2020-11-10 00:45 | 20.6 | 48.4
2020-11-10 00:50 | 20.7 | 48.3
2020-11-10 00:55 | 20.6 | 48.4
2020-11-10 01:00 | 20.6 | 48.4
2020-11-10 01:05 | 20.6 | 48.4
2020-11-10 01:10 | 20.6 | 48.4
2020-11-10 01:15 | 20.7 | 48.4
2020-11-10 01:20 | 20.6 | 48.5
2020-11-10 01:25 | 20.6 | 48.5
2020-11-10 01:30 | 20.7 | 48.3
2020-11-10 01:35 | 20.6 | 48.4
2020-11-10 01:40 | 20.6 | 48.4
2020-11-10 01:45 | 20.6 | 48.4
2020-11-10 01:50 | 20.7 | 48.4
2020-11-10 01:55 | 20.6 | 48.4
2020-11-10 02:00 | 20.6 | 48.4
2020-11-10 02:05 | 20.6 | 48.4
2020-11-10 02:10 | 20.6 | 48.4
2020-11-10 02:15 | 20.6 | 48.4
2020-11-10 02:20 | 20.6 | 48.5
2020-11-10 02:25 | 20.6 | 48.5
2020-11-10 02:30 | 20.6 | 48.5
2020-11-10 02:35 | 20.7 | 48.4
2020-11-10 02:40 | 20.7 | 48.4
2020-11-10 02:45 | 20.6 | 48.5
2020-11-10 02:50 | 20.6 | 48.4
2020-11-10 02:55 | 20.6 | 48.5
2020-11-10 03:00 | 20.6 | 48.5
2020-11-10 03:05 | 20.6 | 48.5
2020-11-10 03:10 | 20.6 | 48.6
2020-11-10 03:15 | 20.6 | 48.5
2020-11-10 03:20 | 20.6 | 48.6
2020-11-10 03:25 | 20.6 | 48.6
2020-11-10 03:30 | 20.6 | 48.6
2020-11-10 03:35 | 20.6 | 48.7
2020-11-10 03:40 | 20.6 | 48.6
2020-11-10 03:45 | 20.6 | 48.7
2020-11-10 03:50 | 20.6 | 48.8
2020-11-10 03:55 | 20.6 | 48.8
2020-11-10 04:00 | 20.6 | 48.8
2020-11-10 04:05 | 20.6 | 48.8
2020-11-10 04:10 | 20.6 | 48.9
2020-11-10 04:15 | 20.6 | 48.9
2020-11-10 04:20 | 20.7 | 48.8
2020-11-10 04:25 | 20.6 | 48.8
2020-11-10 04:30 | 20.6 | 48.8
2020-11-10 04:35 | 20.6 | 48.9
2020-11-10 04:40 | 20.6 | 48.8
2020-11-10 04:45 | 20.6 | 48.8
2020-11-10 04:50 | 20.6 | 48.9
2020-11-10 04:55 | 20.6 | 48.9
2020-11-10 05:00 | 20.6 | 48.9
2020-11-10 05:05 | 20.6 | 48.9
2020-11-10 05:10 | 20.6 | 48.9
2020-11-10 05:15 | 20.6 | 48.8
2020-11-10 05:20 | 20.6 | 48.8
2020-11-10 05:25 | 20.6 | 48.8
2020-11-10 05:30 | 20.6 | 48.8
2020-11-10 05:35 | 20.6 | 48.8
2020-11-10 05:40 | 20.6 | 48.8
2020-11-10 05:45 | 20.6 | 48.9
2020-11-10 05:50 | 20.6 | 48.8
2020-11-10 05:55 | 20.6 | 48.9
2020-11-10 06:00 | 20.6 | 48.8
2020-11-10 06:05 | 20.7 | 48.8
2020-11-10 06:10 | 20.7 | 48.7
2020-11-10 06:15 | 20.6 | 48.8
2020-11-10 06:20 | 20.6 | 48.9
2020-11-10 06:25 | 20.6 | 48.8
2020-11-10 06:30 | 20.6 | 48.8
2020-11-10 06:35 | 20.6 | 48.8
2020-11-10 06:40 | 20.6 | 48.9
2020-11-10 06:45 | 20.6 | 48.9
2020-11-10 06:50 | 20.6 | 48.8
2020-11-10 06:55 | 20.6 | 48.8
2020-11-10 07:00 | 20.7 | 48.7
2020-11-10 07:05 | 20.6 | 48.7
2020-11-10 07:10 | 20.6 | 48.8
2020-11-10 07:15 | 20.6 | 48.8
2020-11-10 07:20 | 20.6 | 48.7
2020-11-10 07:25 | 20.6 | 48.8
2020-11-10 07:30 | 20.6 | 48.8
2020-11-10 07:35 | 20.6 | 48.8
2020-11-10 07:40 | 20.6 | 48.8
2020-11-10 07:45 | 20.6 | 48.7
2020-11-10 07:50 | 20.6 | 48.7
2020-11-10 07:55 | 20.6 | 48.7
2020-11-10 08:00 | 20.7 | 48.6
2020-11-10 08:05 | 20.6 | 48.6
2020-11-10 08:10 | 20.6 | 48.7
2020-11-10 08:15 | 20.6 | 48.7
2020-11-10 08:20 | 20.6 | 48.6
2020-11-10 08:25 | 20.6 | 48.6
2020-11-10 08:30 | 20.6 | 48.7
2020-11-10 08:35 | 20.7 | 48.6
2020-11-10 08:40 | 20.7 | 48.6
2020-11-10 08:45 | 20.6 | 48.7
2020-11-10 08:50 | 20.6 | 48.7
2020-11-10 08:55 | 20.6 | 48.7
2020-11-10 09:00 | 20.6 | 48.6
2020-11-10 09:05 | 20.6 | 48.7
2020-11-10 09:10 | 20.6 | 48.7
2020-11-10 09:15 | 20.6 | 48.7
2020-11-10 09:20 | 20.6 | 48.7
2020-11-10 09:25 | 20.6 | 48.7
2020-11-10 09:30 | 20.6 | 48.8
2020-11-10 09:35 | 20.6 | 48.7
2020-11-10 09:40 | 20.6 | 48.8
2020-11-10 09:45 | 20.6 | 48.9
2020-11-10 09:50 | 20.7 | 48.8
2020-11-10 09:55 | 20.6 | 48.9
2020-11-10 10:00 | 20.6 | 48.9
2020-11-10 10:05 | 20.6 | 49.1
2020-11-10 10:10 | 20.6 | 49.1
2020-11-10 10:15 | 20.6 | 49
2020-11-10 10:20 | 20.6 | 49.1
2020-11-10 10:25 | 20.6 | 49.1
2020-11-10 10:30 | 20.6 | 49.1
2020-11-10 10:35 | 20.6 | 49.1
2020-11-10 10:40 | 20.7 | 49.1
2020-11-10 10:45 | 20.6 | 49.2
2020-11-10 10:50 | 20.6 | 49.2
2020-11-10 10:55 | 20.6 | 49.2
2020-11-10 11:00 | 20.6 | 49.2
2020-11-10 11:05 | 20.6 | 49.1
2020-11-10 11:10 | 20.7 | 49.1
2020-11-10 11:15 | 20.6 | 49.1
2020-11-10 11:20 | 20.7 | 49.1
2020-11-10 11:25 | 20.6 | 49.1
2020-11-10 11:30 | 20.6 | 49.1
2020-11-10 11:35 | 20.6 | 49
2020-11-10 11:40 | 20.6 | 49
2020-11-10 11:45 | 20.6 | 49
2020-11-10 11:50 | 20.6 | 49
2020-11-10 11:55 | 20.6 | 48.9
2020-11-10 12:00 | 20.7 | 48.9
2020-11-10 12:05 | 20.6 | 48.9
2020-11-10 12:10 | 20.7 | 48.8
2020-11-10 12:15 | 20.6 | 48.9
2020-11-10 12:20 | 20.6 | 48.8
2020-11-10 12:25 | 20.6 | 48.8
2020-11-10 12:30 | 20.6 | 48.8
2020-11-10 12:35 | 20.7 | 48.7
2020-11-10 12:40 | 20.7 | 48.7
2020-11-10 12:45 | 20.7 | 49
2020-11-10 12:50 | 20.7 | 49.1
2020-11-10 12:55 | 20.7 | 49.3
2020-11-10 13:00 | 20.7 | 49.3
2020-11-10 13:05 | 20.8 | 49.2
2020-11-10 13:10 | 20.7 | 49.5
2020-11-10 13:15 | 20.6 | 49.7
2020-11-10 13:20 | 20.6 | 49.6
2020-11-10 13:25 | 20.6 | 49.6
2020-11-10 13:30 | 20.6 | 49.3
2020-11-10 13:35 | 20.7 | 49.1
2020-11-10 13:40 | 20.7 | 49
2020-11-10 13:45 | 20.6 | 49.1
2020-11-10 13:50 | 20.7 | 49.1
2020-11-10 13:55 | 20.7 | 49
2020-11-10 14:00 | 20.7 | 48.9
2020-11-10 14:05 | 20.6 | 49.1
2020-11-10 14:10 | 20.7 | 49.1
2020-11-10 14:15 | 20.6 | 49.2
2020-11-10 14:20 | 20.7 | 49.1
2020-11-10 14:25 | 20.6 | 49.3
2020-11-10 14:30 | 20.6 | 49.2
2020-11-10 14:35 | 20.6 | 49.2
2020-11-10 14:40 | 20.6 | 49.2
2020-11-10 14:45 | 20.7 | 49.1
2020-11-10 14:50 | 20.7 | 49
2020-11-10 14:55 | 20.7 | 49.2
2020-11-10 15:00 | 20.7 | 49.2
2020-11-10 15:05 | 20.6 | 49.2
2020-11-10 15:10 | 20.6 | 49.3
2020-11-10 15:15 | 20.6 | 49.2
2020-11-10 15:20 | 20.6 | 49.2
2020-11-10 15:25 | 20.6 | 49.2
2020-11-10 15:30 | 20.6 | 49.1
2020-11-10 15:35 | 20.6 | 49.1
2020-11-10 15:40 | 20.7 | 49.2
2020-11-10 15:45 | 20.6 | 49.2
2020-11-10 15:50 | 20.6 | 49.3
2020-11-10 15:55 | 20.6 | 49.3
2020-11-10 16:00 | 20.7 | 49.3
2020-11-10 16:05 | 20.6 | 49.6
2020-11-10 16:10 | 20.6 | 49.4
2020-11-10 16:15 | 20.7 | 49.4
2020-11-10 16:20 | 20.6 | 49.5
2020-11-10 16:25 | 20.7 | 49.3
2020-11-10 16:30 | 20.7 | 49.4
2020-11-10 16:35 | 20.7 | 49.3
2020-11-10 16:40 | 20.7 | 49.2
2020-11-10 16:45 | 20.6 | 49.4
2020-11-10 16:50 | 20.7 | 49.3
2020-11-10 16:55 | 20.7 | 49.5
2020-11-10 17:00 | 20.6 | 49.6
2020-11-10 17:05 | 20.6 | 49.6
2020-11-10 17:10 | 20.6 | 49.5
2020-11-10 17:15 | 20.6 | 49.4
2020-11-10 17:20 | 20.6 | 49.5
2020-11-10 17:25 | 20.6 | 49.6
2020-11-10 17:30 | 20.6 | 49.5
2020-11-10 17:35 | 20.6 | 49.5
2020-11-10 17:40 | 20.6 | 49.4
2020-11-10 17:45 | 20.6 | 49.5
2020-11-10 17:50 | 20.6 | 49.4
2020-11-10 17:55 | 20.6 | 49.2
2020-11-10 18:00 | 20.6 | 49.2
2020-11-10 18:05 | 20.6 | 49.1
2020-11-10 18:10 | 20.6 | 49
2020-11-10 18:15 | 20.6 | 49
2020-11-10 18:20 | 20.6 | 48.9
2020-11-10 18:25 | 20.6 | 49
2020-11-10 18:30 | 20.6 | 48.9
2020-11-10 18:35 | 20.6 | 48.9
2020-11-10 18:40 | 20.6 | 48.9
2020-11-10 18:45 | 20.6 | 48.7
2020-11-10 18:50 | 20.6 | 48.7
2020-11-10 18:55 | 20.6 | 48.7
2020-11-10 19:00 | 20.6 | 48.6
2020-11-10 19:05 | 20.6 | 48.7
2020-11-10 19:10 | 20.6 | 48.7
2020-11-10 19:15 | 20.5 | 48.8
2020-11-10 19:20 | 20.6 | 48.8
2020-11-10 19:25 | 20.6 | 48.6
2020-11-10 19:30 | 20.6 | 48.6
2020-11-10 19:35 | 20.6 | 48.6
2020-11-10 19:40 | 20.6 | 48.5
2020-11-10 19:45 | 20.6 | 48.6
2020-11-10 19:50 | 20.6 | 48.7
2020-11-10 19:55 | 20.6 | 48.6
2020-11-10 20:00 | 20.6 | 48.5
2020-11-10 20:05 | 20.6 | 48.5
2020-11-10 20:10 | 20.6 | 48.5
2020-11-10 20:15 | 20.6 | 48.5
2020-11-10 20:20 | 20.6 | 48.5
2020-11-10 20:25 | 20.6 | 48.5
2020-11-10 20:30 | 20.6 | 48.6
2020-11-10 20:35 | 20.6 | 48.6
2020-11-10 20:40 | 20.6 | 48.5
2020-11-10 20:45 | 20.6 | 48.4
2020-11-10 20:50 | 20.6 | 48.5
2020-11-10 20:55 | 20.6 | 48.4
2020-11-10 21:00 | 20.6 | 48.4
2020-11-10 21:05 | 20.6 | 48.5
2020-11-10 21:10 | 20.6 | 48.5
2020-11-10 21:15 | 20.6 | 48.5
2020-11-10 21:20 | 20.6 | 48.5
2020-11-10 21:25 | 20.6 | 48.5
2020-11-10 21:30 | 20.6 | 48.4
2020-11-10 21:35 | 20.6 | 48.4
2020-11-10 21:40 | 20.6 | 48.4
2020-11-10 21:45 | 20.6 | 48.5
2020-11-10 21:50 | 20.6 | 48.4
2020-11-10 21:55 | 20.6 | 48.4
2020-11-10 22:00 | 20.6 | 48.4

## pyReport
| Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
A1 | 1 | 994.0168 | 0.063213 | 1.989 | 5.686 | 0.033719 | 2.0834
A2 | 2 | 1980.5213 | 0.098535 | 2.0003 | 8.3745 | 0.049591 | 2.0845
A3 | 3 | 2969.4166 | 0.13829 | 2.0102 | 10.6533 | 0.063054 | 2.0848
A4 | 4 | 3949.1493 | 0.17901 | 2.0157 | 12.6381 | 0.074789 | 2.085
A5 | 5 | 4933.1344 | 0.22044 | 2.0191 | 14.4077 | 0.085248 | 2.0851
A6 | 6 | 5915.8773 | 0.26214 | 2.0212 | 16.0935 | 0.095221 | 2.0851
A7 | 7 | 6892.98 | 0.3038 | 2.0226 | 17.737 | 0.10494 | 2.0851
A8 | 8 | 7865.6403 | 0.34534 | 2.0237 | 19.2525 | 0.11391 | 2.0852
A9 | 9 | 8835.463 | 0.38682 | 2.0245 | 20.7128 | 0.12254 | 2.0852
A10 | 10 | 9806.5859 | 0.42841 | 2.0251 | 22.2151 | 0.13133 | 2.0853
| Tap | Inductance (µH) | Expanded Uncertainty (µH) | Coverage Factor | Resistance (Ω) | Expanded Uncertainty (Ω) | Coverage Factor
B1 | 1 | 99.1417 | 0.0072135 | 1.9977 | 1.1656 | 0.0068951 | 2.085
B2 | 2 | 199.0005 | 0.011164 | 2.0028 | 1.7382 | 0.010279 | 2.0853
B3 | 3 | 298.2492 | 0.015328 | 2.0066 | 2.2051 | 0.013039 | 2.0854
B4 | 4 | 396.5236 | 0.019807 | 2.0094 | 2.79 | 0.016498 | 2.0854
B5 | 5 | 494.4727 | 0.024016 | 2.011 | 3.173 | 0.018765 | 2.0854
B6 | 6 | 593.512 | 0.028287 | 2.0124 | 3.5351 | 0.0209 | 2.0855
B7 | 7 | 691.3425 | 0.032543 | 2.0135 | 3.8929 | 0.023014 | 2.0855
B8 | 8 | 789.1311 | 0.03675 | 2.0145 | 4.1965 | 0.024808 | 2.0855
B9 | 9 | 887.3627 | 0.041013 | 2.0154 | 4.5146 | 0.026688 | 2.0855
B10 | 10 | 984.7941 | 0.045247 | 2.0161 | 4.8226 | 0.028509 | 2.0855

## Results
Date | 2014-10-09
Client | Babcock (NZ) Limited
Job number | 96040221
Report Number | S21952
For | 1592Hz
Results Using Universal Bridge
Dial | Label | Range | a | b | Inductance |  | Resistance |  | Frequency (kHz) | Temperature | Corrected T
[11 empty] | -0.03 | corrrection for GMH51
A | A0 | 3Z | -1 | -10 | 0 | milliH | 0 | ohm | 1.5921 | 19.43
A | A1 | 3Z | 99399 | 569099 | 0.9942 | milliH | 5.6915 | ohm | 1.5922 | 19.44 | 19.41
A | A2 | 3Z | 198037 | 838177 | 1.9808 | milliH | 8.3825 | ohm | 1.5923 | 19.42 | 19.39
A | A3 | 3Z | 296910 | 1066102 | 2.9697 | milliH | 10.6619 | ohm | 1.5924 | 19.46 | 19.43
A | A4 | 3Z | 394862 | 1264597 | 3.9494 | milliH | 12.647 | ohm | 1.5925 | 19.51 | 19.48
A | A5 | 3Z | 493235 | 1441386 | 4.9333 | milliH | 14.415 | ohm | 1.5926 | 19.52 | 19.49
A | A6 | 3Z | 591477 | 1609922 | 5.916 | milliH | 16.1005 | ohm | 1.5927 | 19.49 | 19.46
A | A7 | 3Z | 689150 | 1774172 | 6.8929 | milliH | 17.7431 | ohm | 1.5928 | 19.46 | 19.43
A | A8 | 3Z | 786374 | 1925612 | 7.8653 | milliH | 19.2576 | ohm | 1.5929 | 19.42 | 19.39
A | A9 | 3Z | 883275 | 2071149 | 8.8345 | milliH | 20.7131 | ohm | 1.5929 | 19.41 | 19.38
A | A10 | 3Z | 980309 | 2218833 | 9.8051 | milliH | 22.19 | ohm | 1.5913 | 19.44 | 19.41
B | B1 | 2Z | 99137 | 1166002 | 99.1589 | microH | 1.1661 | ohm | 1.5917 | 19.67 | 19.64
B | B2 | 2Z | 198991 | 1738805 | 199.0329 | microH | 1.7389 | ohm | 1.5918 | 19.65 | 19.62
B | B3 | 2Z | 298239 | 2205906 | 298.3009 | microH | 2.2061 | ohm | 1.5918 | 19.6 | 19.57
B | B4 | 2Z | 396505 | 2790983 | 396.5866 | microH | 2.7912 | ohm | 1.5919 | 19.56 | 19.53
B | B5 | 2Z | 494440 | 3174124 | 494.5412 | microH | 3.1744 | ohm | 1.592 | 19.5 | 19.47
B | B6 | 2Z | 593475 | 3537225 | 593.5961 | microH | 3.5375 | ohm | 1.5916 | 19.66 | 19.63
B | B7 | 2Z | 691294 | 3889800 | 691.4347 | microH | 3.8901 | ohm | 1.5922 | 19.59 | 19.56
B | B8 | 2Z | 789069 | 4192210 | 789.2294 | microH | 4.1925 | ohm | 1.5923 | 19.58 | 19.55
B | B9 | 2Z | 887314 | 4509200 | 887.4941 | microH | 4.5095 | ohm | 1.5923 | 19.5 | 19.47
B | B10 | 2Z | 984751 | 4815910 | 984.9506 | microH | 4.8163 | ohm | 1.5924 | 19.42 | 19.39
B | B0 | 2Z | -2 | -15 | 0 | microH | 0 | ohm | 1.5917 | 19.69 | 19.66
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
C |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
D |  |  |  |  | #N/A | #N/A | #N/A | #N/A
[5 empty] | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
E |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
F |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
G |  |  |  |  | #N/A | #N/A | #N/A | #N/A
HP4284 check | [5 empty] | zero corrected | L | R | UB/HP (scale factor)
700 | 695.657 | 3.8876 | 695.6523 | 3.8877 |  | 700 | 693.7174 | 3.869 | 0.99671 | 1.0054
| 695.657 | 3.8876 | 0.0044987 | 6.694e-05 |  | 600 | 595.6647 | 3.5172 | 0.99653 | 1.0058
| 695.656 | 3.8877 |  |  |  | 500 | 496.2928 | 3.162 | 0.99647 | 1.0039
| 695.648 | 3.8877
| 695.65 | 3.8877 |  |  |  |  | average of 500 and 700 |  | 0.99659 | 1.0047
| 695.652 | 3.8877 |  |  |  |  | % difference at 600 |  | -0.006318 | 0.10703
| 695.646 | 3.8878
Conclude 600 result is good to at least 0.01% for L
600 | 597.6 | 3.5359 | 597.5997 | 3.5359
| 597.596 | 3.5359 | 0.0029439 | 2.2509e-05 |  |  | Reported U | 0.0061784 | %
| 597.599 | 3.5359
| 597.599 | 3.5359 |  |  |  |  | Conclude that HP4284 is not inconsistent with UB
| 597.605 | 3.5359
| 597.599 | 3.5359
500 | 498.219 | 3.1809 | 498.2277 | 3.1806
| 498.227 | 3.1808 | 0.0043861 | 0.00018154
| 498.226 | 3.1806
| 498.231 | 3.1806
| 498.23 | 3.1805
| 498.229 | 3.1804
| 498.232 | 3.1805
short | 1.9538 | 0.01795 | 1.9349 | 0.01864
| 1.9528 | 0.018 | 0.015132 | 0.00054327
| 1.9476 | 0.01818
| 1.9384 | 0.01852
| 1.9213 | 0.01919
| 1.9178 | 0.01917
| 1.9204 | 0.01917
| 1.9271 | 0.01894

## AH2700A results
from 'FinalDecadeCapacitor.xlsx' for S20885
Start with internal cal.
AH2700 will be calibrated 'on the day' against known reference capacitors.
Use these results to have a correction factor adequate for this calibration.
[7 empty] | subtract zero |  |  |  | Value from capacitance scale, out.csv |  |  | measured | NMIA
| Reference | C | sdC | G | sdG |  | C | sdC | G | sdG | C | mean | stdev | tan delta | G | tan delta
|  | pF | pF | nS | nS |  | pF | pF | nS | nS | pF | -2.483 | 0.13805 |  |  |  | correction
| zero | -1.32e-07 | 1.36e-07 | 2.72e-05 | 1.93e-06 |  | 0 | 1.9233e-07 | 0 | 2.7294e-06 |  | ppm correction |  | mrad | nS | mrad | mrad
| AH111A | 10 | 1.8e-07 | -0.00012269 | 2.05e-06 |  | 10 | 2.256e-07 | -0.00014989 | 2.8156e-06 | 9.9999 | -2.4644 |  | -0.0014989 | 0.00018 | 0.0018 | 0.0032989
| AH111B | 10 | 2.65e-07 | -6.61e-05 | 2.22e-06 |  | 10 | 2.9786e-07 | -9.33e-05 | 2.9416e-06 | 10 | -2.4315 |  | -0.000933 | 0.00024 | 0.0024 | 0.003333
| AH111C | 99.9998 | 1.69e-07 | 0.0016369 | 1.58e-06 |  | 99.9998 | 2.1693e-07 | 0.0016097 | 2.4943e-06 | 99.9996 | -2.3617 |  | 0.0016097 | 0.0019 | 0.0019 | 0.0002903
| AH111D | 99.9998 | 2.38e-07 | 0.0020809 | 2.11e-06 |  | 99.9998 | 2.7412e-07 | 0.0020538 | 2.8595e-06 | 99.9996 | -2.3869 |  | 0.0020538 | 0.0023 | 0.0023 | 0.00024625
| GR1000A | 1000.0317 | 6.4e-06 | 0.17936 | 5.45e-05 |  | 1000.0317 | 6.4014e-06 | 0.17933 | 5.4534e-05 | 1000.0289 | -2.7431 |  | 0.017932
| GR1000B | 1000.0164 | 5.92e-06 | 0.18445 | 5.79e-05 |  | 1000.0164 | 5.9216e-06 | 0.18442 | 5.7932e-05 | 1000.0139 | -2.5101 |  | 0.018442
2020-10-12 | UB1000 | 1000.1471 | 5.89e-06 | 0.142 | 9.76e-05 |  | 1000.1471 | 5.8916e-06 | 0.14197 | 9.7619e-05 | 1000.1444 | -2.6266 |  | 0.014195
2020-10-16 | UB1000 | 1000.1471 | 1.4353e-05 | 0.10825 | 0.0001556 |  | 1000.1471 | 1.4353e-05 | 0.10822 | 0.00015561 | 1000.1444 | -2.6266 |  | 0.01082
2020-10-19 | UB1000 | 1000.1393 | 7.4833e-06 | 0.10425 | 4.8629e-05 |  | 1000.1393 | 7.4846e-06 | 0.10422 | 4.8667e-05 | 1000.1367 | -2.6266 |  | 0.010421
2020-10-20 | UB1000 | 1000.1439 | 7.4162e-06 | 0.13882 | 5.3703e-05 |  | 1000.1439 | 7.4174e-06 | 0.13879 | 5.3738e-05 | 1000.1413 | -2.6266 |  | 0.013877
2020-10-23 | UB1000 | 1000.1511 | 3e-06 | 0.10824 | 5.0326e-05 |  | 1000.1511 | 3.0031e-06 | 0.10821 | 5.0363e-05 | 1000.1485 | -2.6266 |  | 0.010819
2020-10-23 | UB1000 | 1000.1506 | 5.3852e-06 | 0.1095 | 4.8936e-05 |  | 1000.1506 | 5.3869e-06 | 0.10947 | 4.8974e-05 | 1000.148 | -2.6266 |  | 0.010945
stdev
ppm
Average | UB1000 | 1000.1465 | 0.0044079 | 0.11851 | 0.017085 |  | 1000.1465 | 0.0044079 | 0.11848 | 0.017085 | 1000.1439 | -2.6266 | 4.4072 | 0.011846
Readings from p.51 of UBIII
from cap scale
2020-11-11 |  | C | G |  | C | Correction
|  | pF | nS |  | pF
| UB1000 | 1000.1593 | 0.13362 |  | 1000.1567 | -0.0026205
| GR1000A | 1000.0317 | 0.17503 |  | 1000.0289 | -0.002776
| GR1000B | 1000.0164 | 0.17621 |  | 1000.0139 | -0.002465

## Room data logger
Primary Lab
2014-10-09 14:00 | 20.41 | 42.67
2014-10-09 14:05 | 20.41 | 42.38
2014-10-09 14:10 | 20.45 | 42.55
2014-10-09 14:15 | 20.51 | 42.72
2014-10-09 14:20 | 20.57 | 42.77
2014-10-09 14:25 | 20.52 | 42.87
2014-10-09 14:30 | 20.53 | 42.67
2014-10-09 14:35 | 20.46 | 42.42
2014-10-09 14:40 | 20.41 | 42.55
2014-10-09 14:45 | 20.41 | 42.28
2014-10-09 14:50 | 20.4 | 42.47
2014-10-09 14:55 | 20.44 | 42.3
2014-10-09 15:00 | 20.4 | 42.65
2014-10-09 15:05 | 20.41 | 42.89
2014-10-09 15:10 | 20.41 | 42.74
2014-10-09 15:15 | 20.42 | 42.77
2014-10-09 15:20 | 20.4 | 42.72
2014-10-09 15:25 | 20.33 | 42.79
2014-10-09 15:30 | 20.35 | 43.19
2014-10-09 15:35 | 20.41 | 43.04
2014-10-09 15:40 | 20.45 | 42.96
2014-10-09 15:45 | 20.41 | 43.19
2014-10-09 15:50 | 20.44 | 42.92
2014-10-09 15:55 | 20.4 | 42.89
2014-10-09 16:00 | 20.47 | 43.23
2014-10-09 16:05 | 20.45 | 42.84
2014-10-09 16:10 | 20.41 | 42.92
2014-10-09 16:15 | 20.39 | 42.79
2014-10-09 16:20 | 20.51 | 43.06
2014-10-09 16:25 | 20.51 | 43.23
2014-10-09 16:30 | 20.5 | 43.41
2014-10-09 16:35 | 20.5 | 43.26
2014-10-09 16:40 | 20.55 | 43.48
2014-10-09 16:45 | 20.51 | 43.43
2014-10-09 16:50 | 20.5 | 43.43
2014-10-09 16:55 | 20.56 | 43.65
2014-10-09 17:00 | 20.56 | 43.58
2014-10-09 17:05 | 20.55 | 43.68
2014-10-09 17:10 | 20.47 | 43.63
2014-10-09 17:15 | 20.49 | 43.75
2014-10-09 17:20 | 20.42 | 43.77
2014-10-09 17:25 | 20.49 | 43.63
2014-10-09 17:30 | 20.52 | 43.73
2014-10-09 17:35 | 20.53 | 43.48
2014-10-09 17:40 | 20.5 | 43.38
2014-10-09 17:45 | 20.46 | 43.5
2014-10-09 17:50 | 20.46 | 43.38
2014-10-09 17:55 | 20.45 | 43.53
2014-10-09 18:00 | 20.49 | 43.43
average | 20.4612 | 43.0735
sd | 0.057686 | 0.43176

## Universal Bridge
New UB calibrations need to be entered in Data block and accessed through the Date |  |  | This table is the usual form for calculating spreadsheet corrections | [6 empty] | Date | C1 | R4A | R4B | R4C | G1 | G2 | c | d
|  |  | Component | Value | Date |  |  |  | 1.6 kHz | 1985-02-14 | 32 | 30.1 | -12 | 53 | 13 | 8.7 | <10 | <10 | except for range 4, up to 44? |  |  |  | -91.32 | 0 | 0 | 0 | 0 | 0
|  |  |  | (ppm) | 2014-07-07 |  |  |  | DC for resistors | 1988-11-30 | 123.32 | 35 | -12 | #N/A | 9 | 4 | 0 | 0 |  |  |  |  | 0 | 4.9 | 0 | #N/A | -4 | -4.7 | 0
Date | Record | Comments | C_1 | 143.1 | UBIII |  |  |  | DC for resistors | 1991-03-12 | 133 | 37 | -14 | 61 | 10 | 6 | 0 | 0 |  |  |  |  | 9.68 | 6.9 | -2 | 8 | -3 | -2.7 | 0
1981-05-22 | PEL Meas V | Check Y&Z ranges 6 & 4 @ 400 Hz | R_4A | 54.1261 | UBIII |  |  |  | DC for resistors | 1995-08-22 | 127.406 | 40.34 | -17.21 | 52 | 11.21 | 6.21 | 0 | 0 |  |  |  |  | 4.086 | 10.24 | -5.21 | -1 | -1.79 | -2.49 | 0
1983-12-15 | Universal Bridge II | Full calibration of standards and of all dial ratios and performance checks. | R_4B | -9.8547 | UBIII |  |  |  | DC for resistors | 1998-11-03 | 135.27 | 43.3741 | -16.3523 | 67.0039 | 11.1844 | 6.6891 | 0 | 0 |  |  |  |  | 11.95 | 13.2741 | -4.3523 | 14.0039 | -1.8156 | -2.0109 | 0
1984-02-10 |  | completion of above | R_4C | 67.1933 | UBIII |  |  |  | DC for resistors | 2001-09-08 | 124.1565 | 46.006 | -17.3198 | 64.091 | 12.0228 | 7.4763 | 0 | 0 |  |  |  |  | 0.83653 | 15.906 | -5.3198 | 11.091 | -0.97725 | -1.2237 | 0
1988-11-30 | UB II | component values measured | G_1 | 12.9485 | UBIII |  |  |  | DC for resistors | 2004-01-22 | 131.7148 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 0 | 0 |  |  |  |  | 8.3948 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 0
1991-03-12 | UB II | component values measured | G_2 | 7.4961 | UBIII |  |  |  | DC for resistors | 2004-02-18 | 123.9 | 47.52 | -16.236 | 63.74 | 12.616 | 7.596 | 5.2 | 0 |  |  |  |  | 0.58 | 17.42 | -4.236 | 10.74 | -0.384 | -1.104 | 5.2
1991-05-16 | UB II | problem with '6' dial, repaired! | [6 empty] | DC for resistors | 2005-02-18 | 129.4493 | 46.54 | -16.96 | 61.5 | 13.84 | 9 | 5.7 | 0 |  |  |  |  | 6.1293 | 16.44 | -4.96 | 8.5 | 0.84 | 0.3 | 5.7
1991-06-18 | AI III | problem with 0.01 mF on range 5Y, lack of DC path for current pump, AVOID | Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | assume negligible c,d |  |  | DC for resistors | 2009-08-13 | 130.2987 | 50.8747 | -14.1043 | 64.5247 | 12.5073 | 7.4753 | 2 | 0 |  |  |  |  | 6.9787 | 20.7747 | -2.1043 | 11.5247 | -0.4927 | -1.2247 | 2
1995-08-22 | UB II | bridge calibration | where Z3=R+jwC | correction_c | 10.97 |  |  |  | DC for resistors | 2014-07-07 | 143.1 | 54.1261 | -9.8547 | 67.1933 | 12.9485 | 7.4961 | 10.97 | 0 |  |  |  |  | 19.78 | 24.0261 | 2.1453 | 14.1933 | -0.051497 | -1.2039 | 10.97
1998-11-03 | UB II | bridge calibration, using HP3458 as transfer instrument | Y4=(G3/G2)(aG1+ßjwC1)(1+c+jd) |  |  | (G3=1/R4*)
2001-09-08 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem fixed | where Y4=G+jwC
2004-01-22 | UBII | bridge calibration, using HP3458 as transfer instrument: Range 5Y problem REALLY fixed
2004-02-18 | UBII | remeasure capacitor, same dc resistance values as for 22/01/04 | R4A is used for ranges 5 to 1 inclusive
2005-02-18 | UBII | bridge calibration, following MSLT.E.023.004. Uncertainties have increased. | R4B is used on range 6 only
2009-08-13 | UBIII | bridge calibration, following MSLT.E.023.004/5. Uncertainties have increased slightly from 2005. | R4C is used on range 7 only | [16 empty] | AH2700A | ppm | mrad
2014-07-07 | UBIII | bridge calibration, following MSLT.E.023.006. Uncertainties unchanged from 2009. | [9 empty] | Zeros | Range | a | b |  |  |  |  | R_4C | 15.0002 | 0.096
[13 empty] | 2Z | -2 | -15
[13 empty] | 3Z | -1 | -10
|  |  |  | Range | a correction | b correction
[5 empty] | (ppm) | (ppm)
|  |  |  | 1Z | 200.7 | 70.5485
|  |  |  | 2Z | 200.7 | 70.5485
|  |  |  | 3Z | 200.7 | 70.5485
|  |  |  | 4Z | 200.7 | 70.5485
|  |  |  | 5Z | 200.7 | 70.5485
|  |  |  | 6Z | 136.7192 | 6.5677
|  |  |  | 7Z | 213.7672 | 83.6157
|  |  |  | 1Y | 70.5079 | -59.6436
|  |  |  | 2Y | 70.5079 | -59.6436
|  |  |  | 3Y | 70.5079 | -59.6436
|  |  |  | 4Y | 70.5079 | -59.6436
|  |  |  | 5Y | 70.5079 | -59.6436
|  |  |  | 6Y | 134.4887 | 4.3372
|  |  |  | 7Y | 57.4407 | -72.7108
[19 empty] | Symbol font versions |  | Plain version
[5 empty] | Multiplier (C&L) | Unit Name | Unit Multiplier | Multiplier (G&R) | Unit Name | Unit Multiplier | CMC n for L | CMC m for R | [6 empty] | Unit Name | Unit Name | Unit Name | Unit Name
|  |  |  | 1Z | 0.0001 | microH | 1e-06 | 0.0001 | milliohm | 0.001 | 0.2 | 0.002 | [6 empty] | mH | mW | microH | milliohm
|  |  |  | 2Z | 0.001 | microH | 1e-06 | 1e-06 | ohm | 1 | 2 | 0.02 | [6 empty] | mH | W | microH | ohm
|  |  |  | 3Z | 1e-05 | milliH | 0.001 | 1e-05 | ohm | 1 | 20 | 0.2 | [6 empty] | mH | W | milliH | ohm
|  |  |  | 4Z | 0.0001 | milliH | 0.001 | 0.0001 | ohm | 1 | 200 | 2 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 5Z | 0.001 | milliH | 0.001 | 0.001 | ohm | 1 | 2000 | 20 | [6 empty] | mH | ohm | milliH | ohm
|  |  |  | 6Z | 1e-05 | H | 1 | 1e-05 | kilohm | 1000 | 20000 | 200 | [6 empty] | H | kW | H | kilohm
|  |  |  | 7Z | 0.0001 | H | 1 | 0.0001 | kilohm | 1000 | 200000 | 2000 | [6 empty] | H | kW | H | kilohm
|  |  |  | 1Y | 0.0001 | microF | 1e-06 | 0.0001 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 2Y | 1e-05 | microF | 1e-06 | 1e-05 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 3Y | 1e-06 | microF | 1e-06 | 1e-06 | milliS | 0.001 | [8 empty] | mF | mS | microF | milliS
|  |  |  | 4Y | 0.0001 | nF | 1e-09 | 0.0001 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 5Y | 1e-05 | nF | 1e-09 | 1e-05 | microS | 1e-06 | [8 empty] | nF | mS | nF | microS
|  |  |  | 6Y | 0.001 | pF | 1e-12 | 0.001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
|  |  |  | 7Y | 0.0001 | pF | 1e-12 | 0.0001 | nS | 1e-09 | [8 empty] | pF | nS | pF | nS
Note on M9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on N9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on O9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on P9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on Q9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on R9: Andrew Corney: capacitor remeasured - use same resistor values as for 22/01/04
Note on V18: k.jones: AH2700A check of 1 M at 1.5915 k. AI VII, p.42
Note on M41: k.jones: The n/f or m/f factor for CMC calculation.

## Uncertainty
UNCERTAINTY |  |  |  |  | Information for calculating expanded uncertainties |  |  |  | Parameters could be dependent on component |  |  |  |  | This is a data table with B9 as the column input cell. | [12 empty] | CMC
GEUM uncertainty calculation template |  |  |  |  | k-factor from degrees of freedom | [8 empty] | This must be re-calculated with new data sets. | [12 empty] | L uncert = Q[0.000014L , 0.001R, n/f ] |  | R uncert = Q[0.000014R,0.0000001L , m/f ]
The component label is the key variable for the table |  |  |  |  | a | 1.96 |  | Model for uncertainty calculation. | Label | R alpha | u_R alpha | L/C alpha | u_L/C alpha | Label | L/C uncertainty | k | R/G uncertainty | k | Range | L/C | mult | n | R/G | mult | m | f | L in microH, R in ohm | CMC/U | L in microH, R in ohm | CMC/U
[5 empty] | b | 2.372 |  | Base on formula for Universal bridge....ie assign an uncertainty to each term. |  | 0.4%/degree |  |  |  |  | 7.0355e-05 | 2.069 | 0.024639 | 2.1612 | [8 empty] | microH |  | ohm
The 95% level of confidence expanded uncertainties are calculated using the Welch-Satterthwaite |  |  |  |  | c | 1.2737 |  |  | A1 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A1 | 7.0355e-05 | 2.069 | 0.024639 | 2.1612 | 3Z | 0.9942 | 0.001 | 20 | 5.6915 | 1 | 0.2 | 1592.22 | 0.019594 | 0.2785 | 0.00017892 | 0.0072614
formula to combine standard uncertainties. The individual degrees of freedom are derived either |  |  |  |  | d | 7.0962 |  |  | A2 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A2 | 0.00012681 | 2.0919 | 0.03629 | 2.1612 | 3Z | 1.9808 | 0.001 | 20 | 8.3825 | 1 | 0.2 | 1592.31 | 0.031576 | 0.24899 | 0.00026227 | 0.0072269
from the number of measurements used to derive the uncertainty or an estimate | [8 empty] | A3 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A3 | 0.00018622 | 2.105 | 0.04616 | 2.1611 | 3Z | 2.9697 | 0.001 | 20 | 10.6619 | 1 | 0.2 | 1592.41 | 0.044721 | 0.24016 | 0.00035531 | 0.0076974
of the percentage uncertainty in the uncertainty. | [8 empty] | A4 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A4 | 0.00024566 | 2.1114 | 0.054756 | 2.1611 | 3Z | 3.9494 | 0.001 | 20 | 12.647 | 1 | 0.2 | 1592.51 | 0.058094 | 0.23648 | 0.00045067 | 0.0082305
Component label | A1 | [7 empty] | A5 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A5 | 0.00030557 | 2.115 | 0.062413 | 2.161 | 3Z | 4.9333 | 0.001 | 20 | 14.415 | 1 | 0.2 | 1592.6 | 0.071664 | 0.23453 | 0.00054761 | 0.008774
Universal bridge reactance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 |  |  |  | A6 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A6 | 0.00036551 | 2.1173 | 0.069713 | 2.161 | 3Z | 5.916 | 0.001 | 20 | 16.1005 | 1 | 0.2 | 1592.68 | 0.085303 | 0.23338 | 0.00064542 | 0.0092582
| Standard | for | 99399 | L/C dial setting |  |  |  |  | A7 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A7 | 0.00042519 | 2.1188 | 0.076827 | 2.1609 | 3Z | 6.8929 | 0.001 | 20 | 17.7431 | 1 | 0.2 | 1592.75 | 0.098918 | 0.23265 | 0.00074336 | 0.0096758
Source of uncertainty | Uncertainties |  |  | Degrees of |  |  |  |  | A8 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A8 | 0.00048461 | 2.12 | 0.083388 | 2.1609 | 3Z | 7.8653 | 0.001 | 20 | 19.2576 | 1 | 0.2 | 1592.85 | 0.11249 | 0.23212 | 0.00084088 | 0.010084
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment |  | A9 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A9 | 0.00054385 | 2.1208 | 0.089693 | 2.1609 | 3Z | 8.8345 | 0.001 | 20 | 20.7131 | 1 | 0.2 | 1592.92 | 0.12603 | 0.23174 | 0.00093827 | 0.010461
ac/dc error of standards | 0.2485 | 0.061751 | 30 | 5.5556 | 0.00068637 | 5ppm considered two sigma limit, essentially relying on dc values of resistors |  |  | A10 | 0.004 | 0.0002 | 2e-05 | 1e-06 | A10 | 0.00060321 | 2.1214 | 0.09609 | 2.1608 | 3Z | 9.8051 | 0.001 | 20 | 22.19 | 1 | 0.2 | 1591.29 | 0.13962 | 0.23146 | 0.0010362 | 0.010784
dial linearity | 1 | 1 | 30 | 5.5556 | 0.18 | 1 ppm of full scale |  |  | B1 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B1 | 0.0076874 | 2.042 | 0.0050478 | 2.1613 | 2Z | 99.1589 | 1e-06 | 2 | 1.1661 | 1 | 0.02 | 1591.73 | 0.0022058 | 0.28694 | 2.2863e-05 | 0.0045293
amplifier/ratio | 1.6898 | 2.8554 | 30 | 5.5556 | 1.4676 | based on scatter of results in 1984, increased in 2004 and 2005 | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. |  | B2 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B2 | 0.013529 | 2.063 | 0.0075276 | 2.1613 | 2Z | 199.0329 | 1e-06 | 2 | 1.7389 | 1 | 0.02 | 1591.79 | 0.0035167 | 0.25994 | 3.3863e-05 | 0.0044985
temperature coefficient of resistors | 0.1988 | 0.039521 | 50 | 2 | 0.00078094 | depends on matching of Vishay components |  |  | B3 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B3 | 0.019545 | 2.078 | 0.0095498 | 2.1613 | 2Z | 298.3009 | 1e-06 | 2 | 2.2061 | 1 | 0.02 | 1591.84 | 0.0048873 | 0.25006 | 4.4739e-05 | 0.0046848
temperature coefficient of capacitor | 0.099399 | 0.0098802 | 30 | 5.5556 | 1.7571e-05 | 2 ppm per degree for capacitor |  |  | B4 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B4 | 0.025714 | 2.0836 | 0.012083 | 2.1613 | 2Z | 396.5866 | 1e-06 | 2 | 2.7912 | 1 | 0.02 | 1591.91 | 0.0063401 | 0.24656 | 5.7076e-05 | 0.0047237
component stability | 0.49699 | 0.247 | 50 | 2 | 0.030505 | some doubt on stability of capacitor in particular, increased in 2004 | Increase from 2 to 5 ppm to account for short-term instability of 1000 pF capacitor |  | B5 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B5 | 0.031708 | 2.09 | 0.013742 | 2.1613 | 2Z | 494.5412 | 1e-06 | 2 | 3.1744 | 1 | 0.02 | 1591.99 | 0.0077195 | 0.24345 | 6.7665e-05 | 0.0049241
L/C error due to phase error in R/G | 0.91056 | 0.82912 | - | 13 | 0.05288 | 16 ppm of resistance dial setting, see comment for resistance, reduced to 5 ppm after 2004 cal | Link to conductance setting, previously just 2 counts |  | B6 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B6 | 0.037781 | 2.0945 | 0.015314 | 2.1613 | 2Z | 593.5961 | 1e-06 | 2 | 3.5375 | 1 | 0.02 | 1591.64 | 0.0091189 | 0.24136 | 7.8321e-05 | 0.0051145
resolution | 1.2561 | 1.5778 | 50 | 2 | 1.2447 | this might be range and frequency dependent | Previous was 2, use latest value from MSLT.E.023.005, 200/f |  | B7 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B7 | 0.043794 | 2.0978 | 0.01684 | 2.1613 | 2Z | 691.4347 | 1e-06 | 2 | 3.8901 | 1 | 0.02 | 1592.21 | 0.010508 | 0.23994 | 8.8908e-05 | 0.0052795
| 0 | 0 | 1 | 5000 | 0 |  |  |  | B8 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B8 | 0.04978 | 2.1007 | 0.01815 | 2.1613 | 2Z | 789.2294 | 1e-06 | 2 | 4.1925 | 1 | 0.02 | 1592.28 | 0.011884 | 0.23874 | 9.9155e-05 | 0.0054632
| Sum of variance | 6.6204 |  | Sum of u^4/v | 2.9772 |  |  |  | B9 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B9 | 0.055816 | 2.1028 | 0.019522 | 2.1613 | 2Z | 887.4941 | 1e-06 | 2 | 4.5095 | 1 | 0.02 | 1592.34 | 0.013278 | 0.23788 | 0.00010964 | 0.005616
| u combined (counts) | 2.573 |  | effective v | 14.7221 |  |  |  | B10 | 0.004 | 0.0002 | 2e-05 | 1e-06 | B10 | 0.061804 | 2.1045 | 0.02085 | 2.1612 | 2Z | 984.9506 | 1e-06 | 2 | 4.8163 | 1 | 0.02 | 1592.38 | 0.01466 | 0.2372 | 0.00012002 | 0.0057565
|  |  |  | k | 2.1292 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 5.4785 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ For all frequencies. Additional noise typical at 50 Hz needs to be added | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Phase angle assumed small, ie no dial interaction. | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3/ Consistent with 1984 calibration, the IEEE paper and recent complement checks | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
4/ Updated following 2004/5 UB calibration, see AUDIREGC | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Universal bridge resistance uncertainty |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | for | 569099 | conductance dial setting |  | all references to the conductance dial are now to cell D34 (previously separate lookups back to results page) | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v |  | Latest comment | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
G/R error due to phase error in L/C | 15.9038 | 252.9321 | - | 13 | 4921.1277 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | Previous was 16, use latest value of 5 from MSLT.E.023.005 | Now return to 16 to stay consistent with CMCS | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
dial linearity | 10 | 100 | 30 | 5.5556 | 1800 |  | Previous value was 2, now decreased to 1 ppm following MSLT.E023.005 | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ac/dc error of standards | 14.2275 | 202.421 | 30 | 5.5556 | 7375.3703 | 5ppm considered two sigma limit, essentially relying on dc values of resistors | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
resolution | 12.5611 | 157.7807 | 30 | 5.5556 | 4481.0541 | this might be dependent on frequency and range, included below, increased from 0 to 1 in 2006, now 2000/f |  | Increase to 5 for 2011 measurements. | [10 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
amplifier/ratio | 9.6747 | 93.5995 | 30 | 5.5556 | 1576.9557 | based on scatter of results in 1984, increased following 2004 and 2005 calibration | Previous value was 7 ppm, now increased to 17 ppm to acount for errors on RHS of bridge. | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient of resistors | 11.382 | 129.5495 | 50 | 2 | 8391.5324 | depends on matching of Vishay components | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
component stability | 11.382 | 129.5495 | 30 | 5.5556 | 3020.9517 | stability of resistors only | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 50 | 2 | 0 |  | Previous value was 1, now increased to 2 ppm following MSLT.E023.005 | [11 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| 0 | 0 | 1 | 5000 | 0 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Sum of variance | 1065.8323 |  | Sum of u^4/v | 31566.9919 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u combined (counts) | 32.6471 |  | effective v | 35.9869 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  |  |  | k | 2.027 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Assumptions |  |  |  | expanded u (counts) | 66.1772 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1/ This scatter is consistent with Thompson's phase standards. | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2/ Updated following 200/54 UB calibration see AUDIREGC | [18 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
[19 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Resistance uncertainty for Sullivan inductor |  |  |  |  | k=a+b/v+c/v^2+d/v^3 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Standard | [17 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Source of uncertainty | Uncertainties |  |  | Degrees of | [14 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
ubridge uncertainty | 32.6471 | 1065.8323 | - | 35.9869 | 31566.9919 | ubridge count uncertainty | [12 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
short term stability | 0 | 0 | 30 | 5.5556 | 0 | [13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
temperature coefficient | 1139.6199 | 1298733.4239 | 20 | 12.5 | 134936680512.975 | assume u_t = 0.5 degree
auxiliary balance senstivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget
| Sum of variance | 1299799.2562 |  | Sum of u^4/v | 134936712079.9669
| u combined (counts) | 1140.0874 |  | effective v | 12.5205
|  |  |  | k | 2.1612
Assumptions |  |  |  | expanded u (counts) | 2463.9447
1/ |  |  |  |  | 0.024639 | ohm
Inductance uncertainty for Sullivan inductor |  |  |  |  | k=a+b/v+c/v^2+d/v^3
Standard
Source of uncertainty | Uncertainties |  |  | Degrees of
| u (counts) | Variances | % Accuracy | Freedom, v | u^4/v
ubridge uncertainty | 2.573 | 6.6204 | - | 14.7221 | 2.9772
temperature coefficient | 0.99523 | 0.99049 | 30 | 5.5556 | 0.17659 | assume 0.5 deg at one sigma and look up the temp coefficient
short term stability | 1.988 | 3.9521 | 30 | 5.5556 | 2.8114 | 20 ppm of dial setting
auxiliary balance sensitivity | 0 | 0 | 30 | 5.5556 | 0 | in Ubridge budget
| Sum of variance | 11.563 |  | Sum of u^4/v | 5.9651
| u combined (count) | 3.4004 |  | effective v | 22.414
|  |  |  | k | 2.069
Assumptions |  |  |  | expanded u (count) | 7.0355
1/ |  |  |  |  | 7.0355e-05 | milliH
Note on B40: k.jones: Modified as for L in keeping with CMCs
Note on B73: k.jones: 2nd order formula is rss of alpha*u_t and u_t*u_alpha.

## Report
checked by: |  |  | humidity | 43.0735 | %RH
|  |  | temperature | 19.485 |  |  |  |  | 2009-07-03
[6 empty] | temperature | 19.485 |  | S21740 | 2009
| dial code | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor
|  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W)
| A1 | 1 | 994.1995 | 0.070355 | 2.069 | 5.6915 | 0.024639 | 2.1612 | 1 | 994.0646 | 0.065681 | 2.0503 | 5.6775 | 0.024545 | 2.1614 | 1 | 994.0646 | 0.065681 | 2.0503 | 5.6775 | 0.024545 | 2.1614
| A2 | 2 | 1980.7775 | 0.12681 | 2.0919 | 8.3825 | 0.03629 | 2.1612 | 2 | 1980.5079 | 0.12156 | 2.065 | 8.3615 | 0.036147 | 2.1615 | 2 | 1980.5079 | 0.12156 | 2.065 | 8.3615 | 0.036147 | 2.1615
| A3 | 3 | 2969.7059 | 0.18622 | 2.105 | 10.6619 | 0.04616 | 2.1611 | 3 | 2969.2816 | 0.17973 | 2.0714 | 10.6351 | 0.045976 | 2.1615 | 3 | 2969.2816 | 0.17973 | 2.0714 | 10.6351 | 0.045976 | 2.1615
| A4 | 4 | 3949.4225 | 0.24566 | 2.1114 | 12.647 | 0.054756 | 2.1611 | 4 | 3948.8537 | 0.23785 | 2.0741 | 12.6152 | 0.054536 | 2.1615 | 4 | 3948.8537 | 0.23785 | 2.0741 | 12.6152 | 0.054536 | 2.1615
| A5 | 5 | 4933.3499 | 0.30557 | 2.115 | 14.415 | 0.062413 | 2.161 | 5 | 4932.5965 | 0.29641 | 2.0755 | 14.3784 | 0.062158 | 2.1615 | 5 | 4932.5965 | 0.29641 | 2.0755 | 14.3784 | 0.062158 | 2.1615
| A6 | 6 | 5915.9671 | 0.36551 | 2.1173 | 16.1005 | 0.069713 | 2.161 | 6 | 5915.0491 | 0.35499 | 2.0763 | 16.0588 | 0.069423 | 2.1615 | 6 | 5915.0491 | 0.35499 | 2.0763 | 16.0588 | 0.069423 | 2.1615
| A7 | 7 | 6892.8931 | 0.42519 | 2.1188 | 17.7431 | 0.076827 | 2.1609 | 7 | 6891.8207 | 0.41329 | 2.0767 | 17.6964 | 0.076503 | 2.1615 | 7 | 6891.8207 | 0.41329 | 2.0767 | 17.6964 | 0.076503 | 2.1615
| A8 | 8 | 7865.3283 | 0.48461 | 2.12 | 19.2576 | 0.083388 | 2.1609 | 8 | 7864.1515 | 0.47135 | 2.0771 | 19.2063 | 0.08303 | 2.1615 | 8 | 7864.1515 | 0.47135 | 2.0771 | 19.2063 | 0.08303 | 2.1615
| A9 | 9 | 8834.5327 | 0.54385 | 2.1208 | 20.7131 | 0.089693 | 2.1609 | 9 | 8833.2017 | 0.52924 | 2.0773 | 20.6573 | 0.089303 | 2.1615 | 9 | 8833.2017 | 0.52924 | 2.0773 | 20.6573 | 0.089303 | 2.1615
| A10 | 10 | 9805.0675 | 0.60321 | 2.1214 | 22.19 | 0.09609 | 2.1608 | 10 | 9803.6322 | 0.58722 | 2.0775 | 22.1296 | 0.095668 | 2.1615 | 10 | 9803.6322 | 0.58722 | 2.0775 | 22.1296 | 0.095668 | 2.1615
|  | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
|  |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor |  |  | uncertainty | factor |  | uncertainty | factor
|  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W) |  |  | (mH) | (mH) |  | (W) | (W)
| B1 | 1 | 99.1589 | 0.0076874 | 2.042 | 1.1661 | 0.0050478 | 2.1613 | 1 | 99.1564 | 0.006623 | 2.046 | 1.1635 | 0.0050299 | 2.1615 | 1 | 99.1564 | 0.006623 | 2.046 | 1.1635 | 0.0050299 | 2.1615
| B2 | 2 | 199.0329 | 0.013529 | 2.063 | 1.7389 | 0.0075276 | 2.1613 | 2 | 199.029 | 0.012292 | 2.0617 | 1.7351 | 0.0075008 | 2.1615 | 2 | 199.029 | 0.012292 | 2.0617 | 1.7351 | 0.0075008 | 2.1615
| B3 | 3 | 298.3009 | 0.019545 | 2.078 | 2.2061 | 0.0095498 | 2.1613 | 3 | 298.2944 | 0.018139 | 2.0688 | 2.2012 | 0.0095158 | 2.1615 | 3 | 298.2944 | 0.018139 | 2.0688 | 2.2012 | 0.0095158 | 2.1615
| B4 | 4 | 396.5866 | 0.025714 | 2.0836 | 2.7912 | 0.012083 | 2.1613 | 4 | 396.5757 | 0.023993 | 2.0715 | 2.7851 | 0.01204 | 2.1615 | 4 | 396.5757 | 0.023993 | 2.0715 | 2.7851 | 0.01204 | 2.1615
| B5 | 5 | 494.5412 | 0.031708 | 2.09 | 3.1744 | 0.013742 | 2.1613 | 5 | 494.5359 | 0.029828 | 2.0732 | 3.1676 | 0.013693 | 2.1615 | 5 | 494.5359 | 0.029828 | 2.0732 | 3.1676 | 0.013693 | 2.1615
| B6 | 6 | 593.5961 | 0.037781 | 2.0945 | 3.5375 | 0.015314 | 2.1613 | 6 | 593.6203 | 0.03574 | 2.0743 | 3.5304 | 0.015262 | 2.1615 | 6 | 593.6203 | 0.03574 | 2.0743 | 3.5304 | 0.015262 | 2.1615
| B7 | 7 | 691.4347 | 0.043794 | 2.0978 | 3.8901 | 0.01684 | 2.1613 | 7 | 691.4345 | 0.041582 | 2.075 | 3.8823 | 0.016783 | 2.1615 | 7 | 691.4345 | 0.041582 | 2.075 | 3.8823 | 0.016783 | 2.1615
| B8 | 8 | 789.2294 | 0.04978 | 2.1007 | 4.1925 | 0.01815 | 2.1613 | 8 | 789.2346 | 0.047424 | 2.0755 | 4.1843 | 0.018089 | 2.1615 | 8 | 789.2346 | 0.047424 | 2.0755 | 4.1843 | 0.018089 | 2.1615
| B9 | 9 | 887.4941 | 0.055816 | 2.1028 | 4.5095 | 0.019522 | 2.1613 | 9 | 887.4889 | 0.053297 | 2.0758 | 4.5009 | 0.019457 | 2.1615 | 9 | 887.4889 | 0.053297 | 2.0758 | 4.5009 | 0.019457 | 2.1615
| B10 | 10 | 984.9506 | 0.061804 | 2.1045 | 4.8163 | 0.02085 | 2.1612 | 10 | 984.942 | 0.059123 | 2.0761 | 4.8072 | 0.020781 | 2.1615 | 10 | 984.942 | 0.059123 | 2.0761 | 4.8072 | 0.020781 | 2.1615

## Historical
| 1995-09-29 |  |  |  | 2000-08-11 |  |  |  | 2005-02-25 |  |  |  | 2009-07-03 |  |  |  | 2014-10-09 |  |  |  | 2020-11-11
| L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR | L | uL | R | uR
A1 | 994.03 | 0.054572 | 5.7182 | 0.024721 | 994.07 | 0.054568 | 5.6995 | 0.02464 | 993.9916 | 0.11571 | 5.6916 | 0.024611 | 994.0646 | 0.065681 | 5.6775 | 0.024545 | 994.1995 | 0.070355 | 5.6915 | 0.024639 | 994.0168 | 0.063213 | 5.686 | 0.033719
A2 | 1980.3588 | 0.10366 | 8.4219 | 0.036409 | 1990.1204 | 0.10418 | 8.4939 | 0.03672 | 1980.4119 | 5.8869 | 8.3817 | 0.03664 | 1980.5079 | 0.12156 | 8.3615 | 0.036147 | 1980.7775 | 0.12681 | 8.3825 | 0.03629 | 1980.5213 | 0.098535 | 8.3745 | 0.049591
A3 | 2968.9979 | 0.15407 | 10.7121 | 0.04631 | 2973.7887 | 0.15432 | 10.7324 | 0.046397 | 2968.8226 | 2.0202 | 10.661 | 0.046165 | 2969.2816 | 0.17973 | 10.6351 | 0.045976 | 2969.7059 | 0.18622 | 10.6619 | 0.04616 | 2969.4166 | 0.13829 | 10.6533 | 0.063054
A4 | 3948.4356 | 0.20423 | 12.7063 | 0.054931 | 3948.9357 | 0.20425 | 12.6663 | 0.054758 | 3948.7118 | 0.179 | 12.6462 | 0.054673 | 3948.8537 | 0.23785 | 12.6152 | 0.054536 | 3949.4225 | 0.24566 | 12.647 | 0.054756 | 3949.1493 | 0.17901 | 12.6381 | 0.074789
A5 | 4931.9739 | 0.25468 | 14.4818 | 0.062607 | 4932.724 | 0.25472 | 14.4375 | 0.062415 | 4932.4416 | 0.21648 | 14.4139 | 0.062315 | 4932.5965 | 0.29641 | 14.3784 | 0.062158 | 4933.3499 | 0.30557 | 14.415 | 0.062413 | 4933.1344 | 0.22044 | 14.4077 | 0.085248
A6 | 5914.202 | 0.30511 | 16.1744 | 0.069924 | 5915.1822 | 0.30515 | 16.1228 | 0.069701 | 5914.8613 | 0.25412 | 16.0991 | 0.069599 | 5915.0491 | 0.35499 | 16.0588 | 0.069423 | 5915.9671 | 0.36551 | 16.1005 | 0.069713 | 5915.8773 | 0.26214 | 16.0935 | 0.095221
A7 | 6890.9693 | 0.35527 | 17.8224 | 0.077048 | 6892.0394 | 0.35532 | 17.7658 | 0.076804 | 6891.6899 | 0.2924 | 17.7417 | 0.0767 | 6891.8207 | 0.41329 | 17.6964 | 0.076503 | 6892.8931 | 0.42519 | 17.7431 | 0.076827 | 6892.98 | 0.3038 | 17.737 | 0.10494
A8 | 7863.1258 | 0.40521 | 19.3421 | 0.083619 | 7864.406 | 0.40527 | 19.2813 | 0.083356 | 7864.0478 | 0.33083 | 19.2569 | 0.08325 | 7864.1515 | 0.47135 | 19.2063 | 0.08303 | 7865.3283 | 0.48461 | 19.2576 | 0.083388 | 7865.6403 | 0.34534 | 19.2525 | 0.11391
A9 | 8832.1617 | 0.455 | 20.8018 | 0.089929 | 8833.702 | 0.45507 | 20.7375 | 0.089651 | 8833.3752 | 0.36924 | 20.7133 | 0.089546 | 8833.2017 | 0.52924 | 20.6573 | 0.089303 | 8834.5327 | 0.54385 | 20.7131 | 0.089693 | 8835.463 | 0.38682 | 20.7128 | 0.12254
A10 | 9802.568 | 0.50486 | 22.2832 | 0.096333 | 9804.3182 | 0.50495 | 22.2153 | 0.096039 | 9803.9928 | 0.40859 | 22.196 | 0.095955 | 9803.6322 | 0.58722 | 22.1296 | 0.095668 | 9805.0675 | 0.60321 | 22.19 | 0.09609 | 9806.5859 | 0.42841 | 22.2151 | 0.13133
B1 | 99.162 | 0.0057279 | 1.1706 | 0.0050607 | 99.158 | 0.0057257 | 1.1675 | 0.0050472 | 99.1611 | 0.0061381 | 1.1672 | 0.0050456 | 99.1564 | 0.006623 | 1.1635 | 0.0050299 | 99.1589 | 0.0076874 | 1.1661 | 0.0050478 | 99.1417 | 0.0072135 | 1.1656 | 0.0068951
B2 | 199.028 | 0.010709 | 1.7457 | 0.0075468 | 199.028 | 0.010706 | 1.741 | 0.0075267 | 199.0264 | 0.0099933 | 1.7402 | 0.007523 | 199.029 | 0.012292 | 1.7351 | 0.0075008 | 199.0329 | 0.013529 | 1.7389 | 0.0075276 | 199.0005 | 0.011164 | 1.7382 | 0.010279
B3 | 298.294 | 0.015774 | 2.2147 | 0.0095743 | 298.278 | 0.015772 | 2.2088 | 0.0095488 | 298.2935 | 0.01444 | 2.2078 | 0.0095441 | 298.2944 | 0.018139 | 2.2012 | 0.0095158 | 298.3009 | 0.019545 | 2.2061 | 0.0095498 | 298.2492 | 0.015328 | 2.2051 | 0.013039
B4 | 396.5528 | 0.020872 | 2.8022 | 0.012114 | 396.5508 | 0.020871 | 2.7977 | 0.012095 | 396.5635 | 0.018253 | 2.7933 | 0.012075 | 396.5757 | 0.023993 | 2.7851 | 0.01204 | 396.5866 | 0.025714 | 2.7912 | 0.012083 | 396.5236 | 0.019807 | 2.79 | 0.016498
B5 | 494.4916 | 0.025901 | 3.1871 | 0.013778 | 494.4936 | 0.025899 | 3.1818 | 0.013755 | 494.5144 | 0.022373 | 3.1765 | 0.013732 | 494.5359 | 0.029828 | 3.1676 | 0.013693 | 494.5412 | 0.031708 | 3.1744 | 0.013742 | 494.4727 | 0.024016 | 3.173 | 0.018765
B6 | 593.5215 | 0.030991 | 3.5497 | 0.015346 | 593.5255 | 0.030989 | 3.5435 | 0.015319 | 594.1446 | 0.12518 | 3.5335 | 0.015276 | 593.6203 | 0.03574 | 3.5304 | 0.015262 | 593.5961 | 0.037781 | 3.5375 | 0.015314 | 593.512 | 0.028287 | 3.5351 | 0.0209
B7 | 691.3553 | 0.036026 | 3.9065 | 0.016888 | 691.3543 | 0.036025 | 3.8996 | 0.016858 | 691.3874 | 0.030514 | 3.8931 | 0.01683 | 691.4345 | 0.041582 | 3.8823 | 0.016783 | 691.4347 | 0.043794 | 3.8901 | 0.01684 | 691.3425 | 0.032543 | 3.8929 | 0.023014
B8 | 789.146 | 0.041052 | 4.2106 | 0.018203 | 789.142 | 0.041049 | 4.2027 | 0.018169 | 789.1783 | 0.034488 | 4.1957 | 0.018138 | 789.2346 | 0.047424 | 4.1843 | 0.018089 | 789.2294 | 0.04978 | 4.1925 | 0.01815 | 789.1311 | 0.03675 | 4.1965 | 0.024808
B9 | 887.3628 | 0.046106 | 4.5293 | 0.019581 | 887.3698 | 0.046104 | 4.5205 | 0.019543 | 887.3962 | 0.038437 | 4.513 | 0.01951 | 887.4889 | 0.053297 | 4.5009 | 0.019457 | 887.4941 | 0.055816 | 4.5095 | 0.019522 | 887.3627 | 0.041013 | 4.5146 | 0.026688
B10 | 984.7705 | 0.05112 | 4.8376 | 0.020913 | 984.7775 | 0.051118 | 4.8279 | 0.020872 | 984.818 | 0.042525 | 4.8219 | 0.020845 | 984.942 | 0.059123 | 4.8072 | 0.020781 | 984.9506 | 0.061804 | 4.8163 | 0.02085 | 984.7941 | 0.045247 | 4.8226 | 0.028509
b10
| L | uL | R | uR
1995-09-29 | 1995-09-29 L: 984.7705; 1995-09-29 uL: 0.05112; 1995-09-29 R: 4.8376; 1995-09-29 uR: 0.020913
2000-08-11 | 1995-09-29 L: 984.7775; 1995-09-29 uL: 0.051118; 1995-09-29 R: 4.8279; 1995-09-29 uR: 0.020872
2005-02-25 | 1995-09-29 L: 984.818; 1995-09-29 uL: 0.042525; 1995-09-29 R: 4.8219; 1995-09-29 uR: 0.020845
2009-07-03 | 1995-09-29 L: 984.942; 1995-09-29 uL: 0.059123; 1995-09-29 R: 4.8072; 1995-09-29 uR: 0.020781
2014-10-09 | 1995-09-29 L: 984.9506; 1995-09-29 uL: 0.061804; 1995-09-29 R: 4.8163; 1995-09-29 uR: 0.02085
2020-11-11 | 1995-09-29 L: 984.7941; 1995-09-29 uL: 0.045247; 1995-09-29 R: 4.8226; 1995-09-29 uR: 0.028509

## Previous History
[6 empty] | checked by: | Tom been here
| 2005-02-15 | S21586 |  |  | 2005-02-15 |  |  |  |  | 2000-08-11 | S21384 | [5 empty] | 2005-02-25
| 9171/1962 | [8 empty] | 9171/1962 | [6 empty] | dial code | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
Label | Range 2Z | [7 empty] | Label | Range 2Z | [9 empty] | uncertainty | factor |  | uncertainty | factor
B10 | 10 | 984646 | 4821626 | [5 empty] | B10 | 10 | 984617 | 4827668 | [6 empty] | (mH) | (mH)
B9 | 9 | 887241 | 4512760 | [5 empty] | B9 | 9 | 887225 | 4520316 |  |  |  |  | A1 | 1 | 993.9916 | 0.11571 | 2.1008 | 5.6916 | 0.024611 | 2.1611
B8 | 8 | 789040 | 4195500 | [5 empty] | B8 | 8 | 789013 | 4202536 |  |  |  |  | A2 | 2 | 1980.4119 | 5.8869 | 2.1615 | 8.3817 | 0.03664 | 2.1473
B7 | 7 | 691266 | 3892830 | [5 empty] | B7 | 7 | 691241 | 3899426 |  |  |  |  | A3 | 3 | 2968.8226 | 2.0202 | 2.1599 | 10.661 | 0.046165 | 2.1591
B6 | 6 | 593428 | 3543346 |  | 594040 | 3533330 |  |  | B6 | 6 | 593428 | 3543346 |  |  |  |  | A4 | 4 | 3948.7118 | 0.179 | 2.1114 | 12.6462 | 0.054673 | 2.1614
B5 | 5 | 494427 | 3176330 | [5 empty] | B5 | 5 | 494412 | 3181644 |  |  |  |  | A5 | 5 | 4932.4416 | 0.21648 | 2.1362 | 14.4139 | 0.062315 | 2.1614
B4 | 4 | 396493 | 2793140 | [5 empty] | B4 | 4 | 396485 | 2797540 |  |  |  |  | A6 | 6 | 5914.8613 | 0.25412 | 2.1586 | 16.0991 | 0.069599 | 2.1614
B3 | 3 | 298240 | 2207660 | [5 empty] | B3 | 3 | 298228 | 2208660 |  |  |  |  | A7 | 7 | 6891.6899 | 0.2924 | 2.1754 | 17.7417 | 0.0767 | 2.1614
B2 | 2 | 198990 | 1740140 | [5 empty] | B2 | 2 | 198994 | 1740961 |  |  |  |  | A8 | 8 | 7864.0478 | 0.33083 | 2.1889 | 19.2569 | 0.08325 | 2.1614
B1 | 1 | 99142 | 1167100 | [5 empty] | B1 | 1 | 99140 | 1167446 |  |  |  |  | A9 | 9 | 8833.3752 | 0.36924 | 2.2006 | 20.7133 | 0.089546 | 2.1614
| 0 | [8 empty] | 0 | [6 empty] | A10 | 10 | 9803.9928 | 0.40859 | 2.2066 | 22.196 | 0.095955 | 2.1615
| 9173/1962 | [8 empty] | 9173/1962
| Range 3Z | [8 empty] | Range 3Z
A10 | 10 | 980229 | 2219460 | [5 empty] | A10 | 10 | 980273 | 2221419
A9 | 9 | 883184 | 2071200 | [5 empty] | A9 | 9 | 883227 | 2073654 | [5 empty] | Tap | Inductance | Expanded | Coverage | Resistance | Expanded | Coverage
A8 | 8 | 786268 | 1925567 | [5 empty] | A8 | 8 | 786313 | 1928038 | [7 empty] | uncertainty | factor |  | uncertainty | factor
A7 | 7 | 689049 | 1774060 | [5 empty] | A7 | 7 | 689092 | 1776493 | [6 empty] | (mH) | (mH) |  | (W) | (W)
A6 | 6 | 591383 | 1609810 | [5 empty] | A6 | 6 | 591422 | 1612197 |  |  |  |  | B1 | 1 | 99.1611 | 0.0061381 | 2.0413 | 1.1672 | 0.0050456 | 2.1615
A5 | 5 | 493158 | 1441300 | [5 empty] | A5 | 5 | 493192 | 1443680 |  |  |  |  | B2 | 2 | 199.0264 | 0.0099933 | 2.07 | 1.7402 | 0.007523 | 2.1615
A4 | 4 | 394802 | 1264540 | [5 empty] | A4 | 4 | 394829 | 1266570 |  |  |  |  | B3 | 3 | 298.2935 | 0.01444 | 2.0766 | 2.2078 | 0.0095441 | 2.1615
A3 | 3 | 296830 | 1066030 | [5 empty] | A3 | 3 | 297330 | 1073180 |  |  |  |  | B4 | 4 | 396.5635 | 0.018253 | 2.1065 | 2.7933 | 0.012075 | 2.1615
A2 | 2 | 198006 | 838110 | [5 empty] | A2 | 2 | 198979 | 849340 |  |  |  |  | B5 | 5 | 494.5144 | 0.022373 | 2.1154 | 3.1765 | 0.013732 | 2.1615
A1 | 1 | 99381 | 569116 | [5 empty] | A1 | 1 | 99390 | 569920 |  |  |  |  | B6 | 6 | 594.1446 | 0.12518 | 2.1432 | 3.5335 | 0.015276 | 2.1615
| 0 | -1 | -10 | [6 empty] | 0 | [6 empty] | B7 | 7 | 691.3874 | 0.030514 | 2.1339 | 3.8931 | 0.01683 | 2.1615
[17 empty] | B8 | 8 | 789.1783 | 0.034488 | 2.143 | 4.1957 | 0.018138 | 2.1615
2000-08-11 | S21384 | [15 empty] | B9 | 9 | 887.3962 | 0.038437 | 2.1527 | 4.513 | 0.01951 | 2.1615
9171/1962 | [16 empty] | B10 | 10 | 984.818 | 0.042525 | 2.1556 | 4.8219 | 0.020845 | 2.1615
Range 2Z
10 | 984617 | 4827668 |  |  | 984617 | 4827668
9 | 887225 | 4520316 |  |  | 887225 | 4520316
8 | 789013 | 4202536 |  |  | 789013 | 4202536
7 | 691241 | 3899426 |  |  | 691241 | 3899426
6 | 593428 | 3543346 |  |  | 593428 | 3543346
5 | 494412 | 3181644 |  |  | 494412 | 3181644
4 | 396485 | 2797540 |  |  | 396485 | 2797540
3 | 298228 | 2208660 |  |  | 298228 | 2208660
2 | 198994 | 1740961 |  |  | 198994 | 1740961
1 | 99140 | 1167446 |  |  | 99140 | 1167446
0
9173/1962
Range 3Z
10 | 980273 | 2221419 |  |  | 980273 | 2221419
9 | 883227 | 2073654 |  |  | 883227 | 2073654
8 | 786313 | 1928038 |  |  | 786313 | 1928038
7 | 689092 | 1776493 |  |  | 689092 | 1776493
6 | 591422 | 1612197 |  |  | 591422 | 1612197
5 | 493192 | 1443680 |  |  | 493192 | 1443680
4 | 394829 | 1266570 |  |  | 394829 | 1266570
3 | 297330 | 1073180 |  |  | 297330 | 1073180
2 | 198979 | 849340 |  |  | 198979 | 849340
1 | 99390 | 569920 |  |  | 99390 | 569920
0
1995-09-29 | [7 empty] | 1995-09-29 | [7 empty] | 2005-02-23
9171/1962 | [7 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage |  | 9171/1962 |  |  |  | L % | R %
Range 2Z | [8 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor |  | Range 2Z |  |  |  |  | 3 Month Drift
10 | 984610 | 4837350 | [5 empty] | 1 | 994.03 | 0.054572 | 2.1524 | 5.7182 | 0.024721 | 2.1615 |  | 10 | 984646 | 4821626 |  | 0 | 0
9 | 887218 | 4529080 | [5 empty] | 2 | 1980.3588 | 0.10366 | 2.2108 | 8.4219 | 0.036409 | 2.1615 |  | 9 | 887241 | 4512760 |  | 0 | 0
8 | 789017 | 4210450 | [5 empty] | 3 | 2968.9979 | 0.15407 | 2.2298 | 10.7121 | 0.04631 | 2.1615 |  | 8 | 789040 | 4195500 |  | 0 | 0
7 | 691242 | 3906360 | [5 empty] | 4 | 3948.4356 | 0.20423 | 2.2381 | 12.7063 | 0.054931 | 2.1615 |  | 7 | 691266 | 3892830 |  | 0 | 0
6 | 593424 | 3549500 | [5 empty] | 5 | 4931.9739 | 0.25468 | 2.2428 | 14.4818 | 0.062607 | 2.1615 |  | 6 | 594040 | 3533330 |  | 1.1602 | -3.18
5 | 494410 | 3186950 | [5 empty] | 6 | 5914.202 | 0.30511 | 2.2456 | 16.1744 | 0.069924 | 2.1615 |  | 5 | 494427 | 3176330 |  | 0 | 0
4 | 396487 | 2802060 | [5 empty] | 7 | 6890.9693 | 0.35527 | 2.2475 | 17.8224 | 0.077048 | 2.1615 |  | 4 | 396493 | 2793140 |  | 0 | 0
3 | 298244 | 2214550 | [5 empty] | 8 | 7863.1258 | 0.40521 | 2.2488 | 19.3421 | 0.083619 | 2.1615 |  | 3 | 298240 | 2207660 |  | 0 | 0
2 | 198994 | 1745600 | [5 empty] | 9 | 8832.1617 | 0.455 | 2.2499 | 20.8018 | 0.089929 | 2.1615 |  | 2 | 198990 | 1740140 |  | 0 | 0
1 | 99144 | 1170560 | [5 empty] | 10 | 9802.568 | 0.50486 | 2.2506 | 22.2832 | 0.096333 | 2.1615 |  | 1 | 99142 | 1167100 |  | 0 | 0
0 | [15 empty] | 0
9173/1962 | [16 empty] | 9173/1962
Range 3Z | [16 empty] | Range 3Z
10 | 980098 | 2228210 | [13 empty] | 10 | 980229 | 2219460 |  | 0 | 0
9 | 883073 | 2080080 | [13 empty] | 9 | 883184 | 2071200 |  | 0 | 0
8 | 786185 | 1934120 | [5 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage |  | 8 | 786268 | 1925567 |  | 0 | 0
7 | 688985 | 1782150 | [6 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor |  | 7 | 689049 | 1774060 |  | 0 | 0
6 | 591324 | 1617360 | [5 empty] | 1 | 99.162 | 0.0057279 | 2.0996 | 1.1706 | 0.0050607 | 2.1615 |  | 6 | 591383 | 1609810 |  | 0 | 0
5 | 493117 | 1448110 | [5 empty] | 2 | 199.028 | 0.010709 | 2.1618 | 1.7457 | 0.0075468 | 2.1615 |  | 5 | 493158 | 1441300 |  | 0 | 0
4 | 394779 | 1270570 | [5 empty] | 3 | 298.294 | 0.015774 | 2.1888 | 2.2147 | 0.0095743 | 2.1615 |  | 4 | 394802 | 1264540 |  | 0 | 0
3 | 296852 | 1071150 | [5 empty] | 4 | 396.5528 | 0.020872 | 2.1975 | 2.8022 | 0.012114 | 2.1615 |  | 3 | 296830 | 1066030 |  | 0 | 0
2 | 198003 | 842150 | [5 empty] | 5 | 494.4916 | 0.025901 | 2.2074 | 3.1871 | 0.013778 | 2.1615 |  | 2 | 198006 | 838110 |  | 0 | 0
1 | 99384 | 571790 | [5 empty] | 6 | 593.5215 | 0.030991 | 2.2143 | 3.5497 | 0.015346 | 2.1615 |  | 1 | 99381 | 569116 |  | 0 | 0
0 | [7 empty] | 7 | 691.3553 | 0.036026 | 2.219 | 3.9065 | 0.016888 | 2.1615 |  | 0
[8 empty] | 8 | 789.146 | 0.041052 | 2.2231 | 4.2106 | 0.018203 | 2.1615
1991-05-30 | [7 empty] | 9 | 887.3628 | 0.046106 | 2.2261 | 4.5293 | 0.019581 | 2.1615 |  |  | 2000-08-11
9171/1962 | [7 empty] | 10 | 984.7705 | 0.05112 | 2.2284 | 4.8376 | 0.020913 | 2.1615
Range 2Z | L | R | [14 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage
10 | 1004838 | 4788256 | [15 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor
9 | 904811 | 4485090 | [14 empty] | 1 | 994.07 | 0.054568 | 2.1525 | 5.6995 | 0.02464 | 2.1615
8 | 803682 | 4172580 | [14 empty] | 2 | 1990.1204 | 0.10418 | 2.2107 | 8.4939 | 0.03672 | 2.1615
7 | 703119 | 3873825 | [14 empty] | 3 | 2973.7887 | 0.15432 | 2.2298 | 10.7324 | 0.046397 | 2.1615
6 | 602695 | 3522505 | [14 empty] | 4 | 3948.9357 | 0.20425 | 2.2382 | 12.6663 | 0.054758 | 2.1615
5 | 501222 | 3164915 | [14 empty] | 5 | 4932.724 | 0.25472 | 2.2429 | 14.4375 | 0.062415 | 2.1615
4 | 401408 | 2783495 | [14 empty] | 6 | 5915.1822 | 0.30515 | 2.2457 | 16.1228 | 0.069701 | 2.1615
3 | 301525 | 220245 | [14 empty] | 7 | 6892.0394 | 0.35532 | 2.2475 | 17.7658 | 0.076804 | 2.1615
2 | 200942 | 1734921 | [14 empty] | 8 | 7864.406 | 0.40527 | 2.2489 | 19.2813 | 0.083356 | 2.1615
1 | 99878 | 1163301 | [14 empty] | 9 | 8833.702 | 0.45507 | 2.2499 | 20.7375 | 0.089651 | 2.1615
0 | [16 empty] | 10 | 9804.3182 | 0.50495 | 2.2507 | 22.2153 | 0.096039 | 2.1615
9173/1962
Range 3Z | L | R
10 | 1000042 | 2185630 | [14 empty] | Tap | Inductance | Uncertainty | Coverage | Resistance | Uncertainty | Coverage
9 | 900119 | 2043361 | [15 empty] | (mH) | (mH) | Factor | (W) | (W) | Factor
8 | 800504 | 1903129 | [14 empty] | 1 | 99.158 | 0.0057257 | 2.0999 | 1.1675 | 0.0050472 | 2.1615
7 | 700510 | 1756857 | [14 empty] | 2 | 199.028 | 0.010706 | 2.1621 | 1.741 | 0.0075267 | 2.1615
6 | 600413 | 1596577 | [14 empty] | 3 | 298.278 | 0.015772 | 2.189 | 2.2088 | 0.0095488 | 2.1615
5 | 500100 | 1431956 | [14 empty] | 4 | 396.5508 | 0.020871 | 2.1977 | 2.7977 | 0.012095 | 2.1615
4 | 399841 | 1258189 | [14 empty] | 5 | 494.4936 | 0.025899 | 2.2076 | 3.1818 | 0.013755 | 2.1615
3 | 300168 | 1062034 | [14 empty] | 6 | 593.5255 | 0.030989 | 2.2145 | 3.5435 | 0.015319 | 2.1615
2 | 199991 | 836012 | [14 empty] | 7 | 691.3543 | 0.036025 | 2.2191 | 3.8996 | 0.016858 | 2.1615
1 | 100176 | 568494 | [14 empty] | 8 | 789.142 | 0.041049 | 2.2233 | 4.2027 | 0.018169 | 2.1615
0 | [16 empty] | 9 | 887.3698 | 0.046104 | 2.2262 | 4.5205 | 0.019543 | 2.1615
[17 empty] | 10 | 984.7775 | 0.051118 | 2.2286 | 4.8279 | 0.020872 | 2.1615
1984-02-16
9171/1962 |  |  | R % | L %
Range 2Z | L | R | 3 Month Drift
10 | 1005990 | 4806400 | 0.012772 | 0.0038745
9 | 905840 | 4507200 | 0.016598 | 0.0038435
8 | 804560 | 4198900 | 0.021209 | 0.0036923
7 | 704300 | 3900100 | 0.022794 | 0.0056735
6 | 604250 | 3550300 | 0.026489 | 0.0087071
5 | 503200 | 3193100 | 0.029865 | 0.0133
4 | 401510 | 2810300 | 0.032272 | 0.00085954
3 | 301630 | 2224200 | 3.0484 | 0.0011778
2 | 201130 | 1756160 | 0.04092 | 0.0031626
1 | 100040 | 1180150 | 0.048306 | 0.005479
0 | 358 | 5860 | 3.3835 | 3.3835
9173/1962
Range 3Z | L | R
10 | 1000616 | 2191510 | 0.0090781 | 0.0019409
9 | 900587 | 2049680 | 0.010431 | 0.0017583
8 | 800915 | 1910750 | 0.013495 | 0.0017363
7 | 701003 | 1766150 | 0.017803 | 0.0023795
6 | 601187 | 1605900 | 0.019643 | 0.004356
5 | 500064 | 1441700 | 0.022868 | -0.00024358
4 | 399832 | 1267980 | 0.026126 | -7.616e-05
3 | 300175 | 1071200 | 0.028951 | 7.8901e-05
2 | 199984 | 844290 | 0.033174 | -0.00011843
1 | 100176 | 574990 | 0.038225 | 0
0 | 30 | 602 | 3.3835 | 3.3835
[6 empty] | 1984-02-16 | 1005990 | 905840 | 804560 | 704300 | 604250 | 503200 | 401510 | 100040 | 358 | 1000616 | 900587 | 800915 | 701003 | 601187 | 500064 | 399832 | 300175 | 199984 | 100176 | 30
[6 empty] | 1991-05-30 | 1004838 | 904811 | 803682 | 703119 | 602695 | 501222 | 401408 | 99878 |  | 1000042 | 900119 | 800504 | 700510 | 600413 | 500100 | 399841 | 300168 | 199991 | 100176
[6 empty] | 1995-09-29 | 984610 | 887218 | 789017 | 691242 | 593424 | 494410 | 396487 | 99144 |  | 980098 | 883073 | 786185 | 688985 | 591324 | 493117 | 394779 | 296852 | 198003 | 99384
[6 empty] | 2000-08-11 | 984617 | 887225 | 789013 | 691241 | 593428 | 494412 | 396485 | 99140 |  | 980273 | 883227 | 786313 | 689092 | 591422 | 493192 | 394829 | 297330 | 198979 | 99390
[6 empty] | 2005-02-15 | 984646 | 887241 | 789040 | 691266 | 594040 | 494427 | 396493 | 99142 |  | 980229 | 883184 | 786268 | 689049 | 591383 | 493158 | 394802 | 296830 | 198006 | 99381 | -1
[6 empty] | 2009-07-03 | 984766 | 887330 | 789093 | 691310 | 593513 | 494446 | 396503 | 99136 |  | 980190 | 883164 | 786276 | 689060 | 591400 | 493172 | 394815 | 296875 | 198015 | 99388
[6 empty] | 1995-09-29 | 0.97987 | 0.98056 | 0.98175 | 0.98311 | 0.98462 | 0.98641 | 0.98774 | 0.99265 |  | 0.98006 | 0.98106 | 0.98211 | 0.98355 | 0.98486 | 0.98604 | 0.98734 | 0.98895 | 0.99006 | 0.99209
[6 empty] | 2000-08-11 | 0.97988 | 0.98056 | 0.98175 | 0.98311 | 0.98462 | 0.98641 | 0.98774 | 0.99261 |  | 0.98023 | 0.98123 | 0.98227 | 0.9837 | 0.98503 | 0.98619 | 0.98747 | 0.99055 | 0.99494 | 0.99215
|  | 984766 | 980190 |  |  | 2005-02-15 | 0.97991 | 0.98058 | 0.98178 | 0.98314 | 0.98564 | 0.98644 | 0.98776 | 0.99263 |  | 0.98019 | 0.98119 | 0.98222 | 0.98364 | 0.98496 | 0.98612 | 0.9874 | 0.98888 | 0.99007 | 0.99206
|  | 887330 | 883164 |  |  | 2009-07-03 | 0.98002 | 0.98068 | 0.98185 | 0.9832 | 0.98477 | 0.98648 | 0.98778 | 0.99257 |  | 0.98015 | 0.98116 | 0.98223 | 0.98365 | 0.98499 | 0.98615 | 0.98743 | 0.98903 | 0.99012 | 0.99213
|  | 789093 | 786276 |  |  | 1995-09-29 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  | 691310 | 689060 |  |  | 2000-08-11 | 1 | 1 | 0.99999 | 1 | 1 | 1 | 0.99999 | 0.99996 |  | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0002 | 1.0001 | 1.0016 | 1.0049 | 1.0001
|  | 593513 | 591400 |  |  | 2005-02-15 | 1 | 1 | 1 | 1 | 1.001 | 1 | 1 | 0.99998 |  | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 0.99993 | 1 | 0.99997
|  | 494446 | 493172 |  |  | 2009-07-03 | 1.0002 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1 | 0.99992 |  | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1
|  | 396503 | 394815
|  | 298239 | 296875
|  | 198991 | 198015 |  |  |  | 34971
|  | 99136 | 99388
500