## Read Me
resolve use/value of 'c' correction term (5 ppm issue)
resolved, derived on 'Resistors' page as mean of corrections from ESI boxes.
On 7 May 2009 am preparing for Fitzroy inductor calibration.
apply temperature correction to SR104 for completeness
use March 09 value for SR104, prior to BIPM correction
attempt full cal in August
completed in November
Started at 24/01/23: whole calibration of UB | sheet color changed to after editing
confirmed and finished, color
Sheets before "ReadMe" are not in use
New sheet from this year 2023
Changes in 2026 | 3458A readings now entering from tab 'out_E3458A_Res_2026-06-11' with raw data in tab 'E3458A_Res_2026-06-11.'
K16 in 'Calibration' is the manual entry of the SR104 value gvien in 'SR104'.
GR1000B excluded from capacitance calculation.

## Calibration
Calibration of Universal Bridge | [15 empty] | This block selects data from the csv Python output
A. Measurement of bridge resistors using dc.
Date: | 2023-01-31 |  |  | Enter data into light blue areas
DMM: | 3458A S/N MY45055679 | [14 empty] | esi SR1010 BOX S/N: 937003
DMM settings: | 4 wire ohms | [5 empty] | Data collected in 'E3458A_Res_2023-01-31.csv' and 'E3458A_Res_2023-02-01.csv', combined to "UBcalibrationJune2023_OD.xlsx" | [12 empty] | reading | sd
[6 empty] | GMH5I | using MeasureR.py to gather HP3458 readings over GPIB | [6 empty] | Tue Jan 31 16:05:39 2023 | R4A | 10000.5803 | 0.0015774 | sr104 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328 | 10000.0076
Observation No | Measurement | OCOMP ON? | DELAY | Mean of 20 | sd of 20 | temp | sd ppm | Observation No | sensor R | Temp | R to add to 20 value |  |  | Tue Jan 31 11:25:10 2023 | Sensor SR104 | 9971.2359 | 0.019186 | sens_sr104 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186 | 9971.2359
1 | SR104 | on | 0 | 10000.0076 | 0.0017328 | 21.03 | 0.17328 | 2 | 9971.2359 | 20.1228 | 0.00041598 |  |  | Tue Jan 31 11:11:00 2023 | SR104 | 10000.0076 | 0.0017328 | ep100k | Thu Jun 11 10:22:55 2026 | 10000.1244 | 0.0019637 | 10000.1244
3 | 100k esi parallel | on | 0 | 10000.1244 | 0.0019637 | 21.39 | 0.19637 |  |  | 20.1228 |  |  |  | Tue Jan 31 12:06:30 2023 | 100k esi parallel | 10000.1244 | 0.0019637 | esp100k | Thu Jun 11 10:35:06 2026 | 100000.9992 | 0.015814 | 100000.9992
4 | 100k esi series-para | on | 0 | 100000.9992 | 0.015814 | 21.4 | 0.15814 | [6 empty] | Wed Feb 1 17:09:44 2023 | 100k esi series-para | 100000.9992 | 0.015814 | er10_100k | Thu Jun 11 10:45:56 2026 | 100000.7244 | 0.014969 | 100000.7244
5 | 100k esi R10 | on | 0 | 100000.7244 | 0.014969 | 21.47 | 0.14969 | [6 empty] | Tue Jan 31 13:26:56 2023 | 100k esi R10 | 100000.7244 | 0.014969 | es100k | Thu Jun 11 11:07:10 2026 | 1000012.626 | 3.2902 | 1000012.626
6 | 100k esi series | on | 0 | 1000012.626 | 3.2902 | 21.4 | 3.2901 | [6 empty] | Tue Jan 31 17:03:53 2023 | 100k esi series | 1000012.626 | 3.2902 | g1 | Thu Jun 11 14:56:11 2026 | 99998.2436 | 0.016354 | 99998.2436
7 | G1 | on | 0 | 99998.2436 | 0.016354 | 21.08 | 0.16355 | [6 empty] | Tue Jan 31 16:11:03 2023 | R4B | 99999.5235 | 0.014609 | g2 | Thu Jun 11 15:04:25 2026 | 99998.793 | 0.014888 | 99998.793
8 | G2 | on | 0 | 99998.793 | 0.014888 | 21.08 | 0.14888 | [6 empty] | Tue Jan 31 14:48:45 2023 | G1 | 99998.2436 | 0.016354 | R4A | Thu Jun 11 15:10:57 2026 | 10000.5803 | 0.0015774 | 10000.5803
9 | R4A | on | 0 | 10000.5803 | 0.0015774 | 21.35 | 0.15773 | [6 empty] | Tue Jan 31 14:57:29 2023 | R2(correct name:G2) | 99998.793 | 0.014888 | R4B | Thu Jun 11 15:22:08 2026 | 99999.5235 | 0.014609 | 99999.5235
10 | R4B | on | 0 | 99999.5235 | 0.014609 | 21.43 | 0.14609 |  | SR104 cal val | 10000.0027 | at 20 |  |  | Tue Jan 31 16:11:03 2023 | R4C | 1000073.316 | 0.83842 | R4C | Thu Jun 11 15:32:49 2026 | 1000073.316 | 0.83842 | 1000073.316
11 | R4C | on | 0 | 1000073.316 | 0.83842 | 21.26 | 0.83835 |  |  | see SR104 worksheet | [7 empty] | sr104 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328 | 10000.0076
[18 empty] | sens_sr104 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186 | 9971.2359
Analysis
Resistor | Nominal Value (ohms) | ppm of nominal resistance | ppm of nominal conductance | mean DMM reading (ohms) | stdev (ohms) | DMM correction (ohms) | Value of Resistor (ohms) | Traceability - last cal date | Reference |  |  |  | DMM corrn ppm | stdev ppm | U ppm
SR104 | 10000 |  |  | 10000.0076 | 0.0017328 | -0.004454 | 10000.0031 | 2026-06-19 | See SR104 worksheet |  |  |  | -0.4454 | 0.17328 | 0.078 | sheet:ResultsFrom ValuesUsed; cell I18
100k esi parallel | 10000 | 11.9986 |  | 10000.1244 | 0.0019637 |  | 10000.12 | [6 empty] | 0.19637
100k esi series-para | 100000 | 12.2734 |  | 100000.9992 | 0.015814 | 0.22814 | 100001.2273 | Note: see ESI model SR1010 handbook for series/parallel formula |  |  |  |  | 2.2814 | 0.15814
100k esi R10 | 100000 |  |  | 100000.7244 | 0.014969 | [8 empty] | 0.14969
100k esi series | 1000000 | 11.9986 |  | 1000012.626 | 3.2902 | -0.6274 | 1000011.9986 | [5 empty] | -0.62739 | 3.2901
G1 | 100000 |  | 15.2828 | 99998.2436 | 0.016354 |  | 99998.4717 | [6 empty] | 0.16355
G2 | 100000 |  | 9.7891 | 99998.793 | 0.014888 |  | 99999.0211 | [6 empty] | 0.14888
R4A | 10000 | 57.5886 |  | 10000.5803 | 0.0015774 |  | 10000.5759 | [6 empty] | 0.15773
R4B | 100000 | -2.4836 |  | 99999.5235 | 0.014609 |  | 99999.7516 | [6 empty] | 0.14609 |  | 1.08e-07
R4C | 1000000 | 72.6886 |  | 1000073.316 | 0.83842 |  | 1000072.6886 | [6 empty] | 0.83835
B. Measurement of 1000 pF GR capacitor S/N 2248
|  |  |  | PRT | UB1000
|  | C | U | Traceability
|  | ppm | ppm | Last cal date | Reference
value of GR1000A |  | 29.15 | 0.44 | 2026-07-19 | Prediction in E.005.04, p.33 | [6 empty] | Calculation of temp by the Temp sensors
value of GR1000B |  | 12.4 | 1.1 | 2026-07-19 | Prediction in E.005.04, p.33 | [6 empty] | From E009-011 P23
Measurement with AH2700A
Date: | 2026-06-15
|  |  |  | PRT | UB1000
Observation No. | Measurement | C pF | G nS | data logger | temperature (deg C)
12 | GR1000A | 1000.0333 | 0.21189 | 23..0065
13 | GR 1000B | 999.9995 | 0.2257 | 23.0065
14 | UB 1000 | 1000.1614 | 0.16223 |  | 20.81 | deg C
Meter Correction
ppm
GR1000A |  |  | -4.15
GR 1000B |  |  | 12.86
UB 1000 |  | 1000.1572 | -4.15
Summary of results for pasting into AudioRegC
C_1 | R_4A | R_4B | R_4C | G_1 | G_2
157.25 | 57.5886 | -2.4836 | 72.6886 | 15.2828 | 9.7891
Note on H17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 10 times bigger than 2017
Note on C46: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: GR1000B now of doubtful stability. Still following long-term drift but likely more influenced by humidity?
Note on D52: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Now using only the GR1000A correction.

## SR104
Hi Keith,
According to ​ResistorHistory.xlsx​, the last measured value for SR104r was:
Date: 09/06/2025
Dev from nominal: +0.079 ppm
Exp Uncert: 0.117 ppm (k=2)
Although this fell a bit below the usual linear trend-line (possibly a bad measurement).
Alternatively, I can give you a predicted value for today (std unc = +/- 30 days with 8 dof):
Date: 19/06/2026
Dev from nominal: +0.269 ppm
Exp Uncert: 0.029 ppm (k=2)
Cheers,
Tim.
From: Keith Jones <keith.jones@measurement.govt.nz>
Sent: Friday, June 19, 2026 9:20 AM
To: Tim Lawson <tim.lawson@measurement.govt.nz>
Subject: SR104 Serial 109006
Hi Tim,
Could you please give me a pointer to the most recent value of this 10 k resistor? It is needed for traceability of the Universal bridge.
Thanks,
Keith

## out_E3458A_Res_2026-06-11
| Resistor Fluke 8508A-7000K Four-wire shorting PCB 410590-A me]asured by HP3458A with ID number 5 | Thu Jun 11 09:47:26 2026 | 3.45e-05 | 9.61e-06
| Resistor SR104 StdR (SN:109006) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328
| Resistor SR04TS(temperature Sensor) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186
| Resistor SR1010--100k Parallel with PC 101 (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:22:55 2026 | 10000.1244 | 0.0019637
| Resistor SR1010--100k Series Parallel with PC 102(as above SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:35:06 2026 | 100000.9992 | 0.015814
| Resistor SR1010--100k R10 only (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:45:56 2026 | 100000.7244 | 0.014969
| Resistor SR1010--100k R1-R10 inSeries (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:07:10 2026 | 1000012.626 | 3.2902
| Resistor SR1010_1k parallel_R1-R10 with Pc 101 (SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:21:29 2026 | 99.996 | 9.96e-05
| Resistor SR1010_1k series parallel_R1-R9 with PC 102(SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:34:01 2026 | 999.9598 | 0.00018332
| Resistor SR1010_1k R10 only (SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:44:36 2026 | 999.9722 | 0.00014671
| Resistor SR1010_1k series R1-R10(SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:56:07 2026 | 9999.6007 | 0.014931
| Resistor UB BNC BPO coax zero me]asured by HP3458A with ID number 5 | Thu Jun 11 13:21:35 2026 | -3.67e-07 | 9.08e-06
| Resistor Vishay 100 (RHK02) me]asured by HP3458A with ID number 5 | Thu Jun 11 13:45:11 2026 | 2.04e-05 | 8.34e-05
| Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5 | Thu Jun 11 13:52:44 2026 | 2e+36 | 1.41e+37
| Resistor Vishay 100 (RHK02)-right me]asured by HP3458A with ID number 5 | Thu Jun 11 14:00:58 2026 | 100.0152 | 9.92e-05
| Resistor 1e4 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:22:52 2026 | 10000.3143 | 0.0017302
| Resistor Vishay 10k (UBCAL10k)-right me]asured by HP3458A with ID number 5 | Thu Jun 11 14:30:33 2026 | 10000.3213 | 0.009968
| Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5 | Thu Jun 11 14:38:08 2026 | 10000.3147 | 0.001471
| Resistor Vishay 100k me]asured by HP3458A with ID number 5 | Thu Jun 11 14:46:29 2026 | 100002.594 | 0.030633
| Resistor UB G1 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:56:11 2026 | 99998.2436 | 0.016354
| Resistor UB G2 me]asured by HP3458A with ID number 5 | Thu Jun 11 15:04:25 2026 | 99998.793 | 0.014888
| Resistor UB R4A me]asured by HP3458A with ID number 5 | Thu Jun 11 15:10:57 2026 | 1000073.316 | 0.83842
| Resistor UB R2B me]asured by HP3458A with ID number 5 | Thu Jun 11 15:22:08 2026 | 10000.5803 | 0.0015774
| Resistor UB R4C me]asured by HP3458A with ID number 5 | Thu Jun 11 15:32:49 2026 | 99999.5235 | 0.014609
| Resistor coaxial zero me]asured by HP3458A with ID number 5 | Thu Jun 11 15:41:13 2026 | 2.97e-07 | 8.23e-06
| Resistor Fluke 8508A-7000K Four-wire shorting PCB 410590-A me]asured by HP3458A with ID number 5 | Thu Jun 11 09:47:26 2026 | 3.45e-05 | 9.61e-06
| Resistor SR104 StdR (SN:109006) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328
| Resistor SR04TS(temperature Sensor) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186
| Resistor SR1010--100k Parallel with PC 101 (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:22:55 2026 | 10000.1244 | 0.0019637
| Resistor SR1010--100k Series Parallel with PC 102(as above SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:35:06 2026 | 100000.9992 | 0.015814
| Resistor SR1010--100k R10 only (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:45:56 2026 | 100000.7244 | 0.014969
| Resistor SR1010--100k R1-R10 inSeries (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:07:10 2026 | 1000012.626 | 3.2902
| Resistor SR1010_1k parallel_R1-R10 with Pc 101 (SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:21:29 2026 | 99.996 | 9.96e-05
| Resistor SR1010_1k series parallel_R1-R9 with PC 102(SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:34:01 2026 | 999.9598 | 0.00018332
| Resistor SR1010_1k R10 only (SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:44:36 2026 | 999.9722 | 0.00014671
| Resistor SR1010_1k series R1-R10(SN:923001) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:56:07 2026 | 9999.6007 | 0.014931
| Resistor UB BNC BPO coax zero me]asured by HP3458A with ID number 5 | Thu Jun 11 13:21:35 2026 | -3.67e-07 | 9.08e-06
| Resistor Vishay 100 (RHK02) me]asured by HP3458A with ID number 5 | Thu Jun 11 13:45:11 2026 | 2.04e-05 | 8.34e-05
| Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5 | Thu Jun 11 13:52:44 2026 | 2e+36 | 1.41e+37
| Resistor Vishay 100 (RHK02)-right me]asured by HP3458A with ID number 5 | Thu Jun 11 14:00:58 2026 | 100.0152 | 9.92e-05
| Resistor 1e4 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:22:52 2026 | 10000.3143 | 0.0017302
| Resistor Vishay 10k (UBCAL10k)-right me]asured by HP3458A with ID number 5 | Thu Jun 11 14:30:33 2026 | 10000.3213 | 0.009968
| Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5 | Thu Jun 11 14:38:08 2026 | 10000.3147 | 0.001471
| Resistor Vishay 100k me]asured by HP3458A with ID number 5 | Thu Jun 11 14:46:29 2026 | 100002.594 | 0.030633
| Resistor UB G1 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:56:11 2026 | 99998.2436 | 0.016354
| Resistor UB G2 me]asured by HP3458A with ID number 5 | Thu Jun 11 15:04:25 2026 | 99998.793 | 0.014888
| Resistor UB R4A me]asured by HP3458A with ID number 5 | Thu Jun 11 15:10:57 2026 | 1000073.316 | 0.83842
| Resistor UB R2B me]asured by HP3458A with ID number 5 | Thu Jun 11 15:22:08 2026 | 10000.5803 | 0.0015774
| Resistor UB R4C me]asured by HP3458A with ID number 5 | Thu Jun 11 15:32:49 2026 | 99999.5235 | 0.014609
| Resistor coaxial zero me]asured by HP3458A with ID number 5 | Thu Jun 11 15:41:13 2026 | 2.97e-07 | 8.23e-06
sr104 | Resistor SR104 StdR (SN:109006) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328 | 10000.0076
sens_sr104 | Resistor SR04TS(temperature Sensor) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186 | 9971.2359
ep100k | Resistor SR1010--100k Parallel with PC 101 (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:22:55 2026 | 10000.1244 | 0.0019637 | 10000.1244
esp100k | Resistor SR1010--100k Series Parallel with PC 102(as above SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:35:06 2026 | 100000.9992 | 0.015814 | 100000.9992
er10_100k | Resistor SR1010--100k R10 only (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:45:56 2026 | 100000.7244 | 0.014969 | 100000.7244
es100k | Resistor SR1010--100k R1-R10 inSeries (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:07:10 2026 | 1000012.626 | 3.2902 | 1000012.626
g1 | Resistor UB G1 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:56:11 2026 | 99998.2436 | 0.016354 | 99998.2436
g2 | Resistor UB G2 me]asured by HP3458A with ID number 5 | Thu Jun 11 15:04:25 2026 | 99998.793 | 0.014888 | 99998.793
R4A | Resistor UB R4A me]asured by HP3458A with ID number 5 | Thu Jun 11 15:10:57 2026 | 1000073.316 | 0.83842 | 1000073.316
R4B | Resistor UB R2B me]asured by HP3458A with ID number 5 | Thu Jun 11 15:22:08 2026 | 10000.5803 | 0.0015774 | 10000.5803
R4C | Resistor UB R4C me]asured by HP3458A with ID number 5 | Thu Jun 11 15:32:49 2026 | 99999.5235 | 0.014609 | 99999.5235
sr104 | Resistor SR104 StdR (SN:109006) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:00:21 2026 | 10000.0076 | 0.0017328 | 10000.0076
sens_sr104 | Resistor SR04TS(temperature Sensor) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:09:23 2026 | 9971.2359 | 0.019186 | 9971.2359
ep100k | Resistor SR1010--100k Parallel with PC 101 (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:22:55 2026 | 10000.1244 | 0.0019637 | 10000.1244
esp100k | Resistor SR1010--100k Series Parallel with PC 102(as above SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:35:06 2026 | 100000.9992 | 0.015814 | 100000.9992
er10_100k | Resistor SR1010--100k R10 only (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 10:45:56 2026 | 100000.7244 | 0.014969 | 100000.7244
es100k | Resistor SR1010--100k R1-R10 inSeries (SN:937003) me]asured by HP3458A with ID number 5 | Thu Jun 11 11:07:10 2026 | 1000012.626 | 3.2902 | 1000012.626
g1 | Resistor UB G1 me]asured by HP3458A with ID number 5 | Thu Jun 11 14:56:11 2026 | 99998.2436 | 0.016354 | 99998.2436
g2 | Resistor UB G2 me]asured by HP3458A with ID number 5 | Thu Jun 11 15:04:25 2026 | 99998.793 | 0.014888 | 99998.793
R4A | Resistor UB R4A me]asured by HP3458A with ID number 5 | Thu Jun 11 15:10:57 2026 | 1000073.316 | 0.83842 | 1000073.316
R4B | Resistor UB R2B me]asured by HP3458A with ID number 5 | Thu Jun 11 15:22:08 2026 | 10000.5803 | 0.0015774 | 10000.5803
R4C | Resistor UB R4C me]asured by HP3458A with ID number 5 | Thu Jun 11 15:32:49 2026 | 99999.5235 | 0.014609 | 99999.5235

## E3458A_Res_2026-06-11.
Auto Calibration
HP3458A
mains frequency = | 50
temperature = | 32.6
ACAL starting
ACAL finished
Time taken = | 16.1124 | minutes
temperature = | 32.6
Cal done
Thu Jun 11 09:27:26 2026
end of the measurement
Resistor Fluke 8508A-7000K Four-wire shorting PCB 410590-A me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.6
Range?: | 0.001 | Item?: | Fluke 8508A-7000K Four-wire shorting PCB 410590-A
Trail1 | 4.71e-05 | Trail2 | 3.22e-05
at | Thu Jun 11 09:47:26 2026 | Fluke 8508A-7000K Four-wire shorting PCB 410590-A | reading is:
Resistance_mean: | 3.45e-05
Resistance_stdevp: | 9.61e-06
Time taken = | 6.1372 | minutes
----------------
Resistor SR104 StdR (SN:109006) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.6
Range?: | 10000 | Item?: | SR104 StdR (SN:109006)
Trail1 | 10000.0073 | Trail2 | 10000.0073
at | Thu Jun 11 10:00:21 2026 | SR104 StdR (SN:109006) | reading is:
Resistance_mean: | 10000.0076
Resistance_stdevp: | 0.0017328
Time taken = | 6.139 | minutes
----------------
Resistor SR04TS(temperature Sensor) me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.7
Range?: | 10000 | Item?: | SR04TS(temperature Sensor)
Trail1 | 9971.2011 | Trail2 | 9971.2037
at | Thu Jun 11 10:09:23 2026 | SR04TS(temperature Sensor) | reading is:
Resistance_mean: | 9971.2359
Resistance_stdevp: | 0.019186
Time taken = | 6.1358 | minutes
----------------
Resistor SR1010--100k Parallel with PC 101 (SN:937003) me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.7
Range?: | 10000 | Item?: | SR1010--100k Parallel with PC 101 (SN:937003)
Trail1 | 10000.1264 | Trail2 | 10000.1274
at | Thu Jun 11 10:22:55 2026 | SR1010--100k Parallel with PC 101 (SN:937003) | reading is:
Resistance_mean: | 10000.1244
Resistance_stdevp: | 0.0019637
Time taken = | 6.1381 | minutes
----------------
Resistor SR1010--100k Series Parallel with PC 102(as above SN:937003) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 100000 | Item?: | SR1010--100k Series Parallel with PC 102(as above SN:937003)
Trail1 | 100001.0166 | Trail2 | 100001.0058
at | Thu Jun 11 10:35:06 2026 | SR1010--100k Series Parallel with PC 102(as above SN:937003) | reading is:
Resistance_mean: | 100000.9992
Resistance_stdevp: | 0.015814
Time taken = | 6.3468 | minutes
----------------
Resistor SR1010--100k R10 only (SN:937003) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 100000 | Item?: | SR1010--100k R10 only (SN:937003)
Trail1 | 100000.7323 | Trail2 | 100000.7407
at | Thu Jun 11 10:45:56 2026 | SR1010--100k R10 only (SN:937003) | reading is:
Resistance_mean: | 100000.7244
Resistance_stdevp: | 0.014969
Time taken = | 6.3402 | minutes
----------------
Resistor SR1010--100k R1-R10 inSeries (SN:937003) me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.7
Range?: | 1000000 | Item?: | SR1010--100k R1-R10 inSeries (SN:937003)
Trail1 | 1000012.01 | Trail2 | 1000012.483
at | Thu Jun 11 11:07:10 2026 | SR1010--100k R1-R10 inSeries (SN:937003) | reading is:
Resistance_mean: | 1000012.626
Resistance_stdevp: | 3.2902
Time taken = | 4.3014 | minutes
----------------
Resistor SR1010_1k parallel_R1-R10 with Pc 101 (SN:923001) me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.7
Range?: | 100 | Item?: | SR1010_1k parallel_R1-R10 with Pc 101 (SN:923001)
Trail1 | 99.9962 | Trail2 | 99.996
at | Thu Jun 11 11:21:29 2026 | SR1010_1k parallel_R1-R10 with Pc 101 (SN:923001) | reading is:
Resistance_mean: | 99.996
Resistance_stdevp: | 9.96e-05
Time taken = | 6.1382 | minutes
----------------
Resistor SR1010_1k series parallel_R1-R9 with PC 102(SN:923001) me]asured by HP3458A with ID number 5
mains frequency = | 50.1
temperature = | 32.7
Range?: | 1000 | Item?: | SR1010_1k series parallel_R1-R9 with PC 102(SN:923001)
Trail1 | 999.9601 | Trail2 | 999.96
at | Thu Jun 11 11:34:01 2026 | SR1010_1k series parallel_R1-R9 with PC 102(SN:923001) | reading is:
Resistance_mean: | 999.9598
Resistance_stdevp: | 0.00018332
Time taken = | 6.1373 | minutes
----------------
Resistor SR1010_1k R10 only (SN:923001) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 1000 | Item?: | SR1010_1k R10 only (SN:923001)
Trail1 | 999.9727 | Trail2 | 999.9726
at | Thu Jun 11 11:44:36 2026 | SR1010_1k R10 only (SN:923001) | reading is:
Resistance_mean: | 999.9722
Resistance_stdevp: | 0.00014671
Time taken = | 6.1369 | minutes
----------------
Resistor SR1010_1k series R1-R10(SN:923001) me]asured by HP3458A with ID number 5
mains frequency = | 50.1
temperature = | 32.7
Range?: | 10000 | Item?: | SR1010_1k series R1-R10(SN:923001)
Trail1 | 9999.6024 | Trail2 | 9999.6045
at | Thu Jun 11 11:56:07 2026 | SR1010_1k series R1-R10(SN:923001) | reading is:
Resistance_mean: | 9999.6007
Resistance_stdevp: | 0.014931
Time taken = | 6.1377 | minutes
----------------
Resistor UB BNC BPO coax zero me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.6
Range?: | 0.001 | Item?: | UB BNC BPO coax zero
Trail1 | -3.95e-06 | Trail2 | 6.5e-06
at | Thu Jun 11 13:21:35 2026 | UB BNC BPO coax zero | reading is:
Resistance_mean: | -3.67e-07
Resistance_stdevp: | 9.08e-06
Time taken = | 6.1378 | minutes
----------------
Resistor Vishay 100 (RHK02) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 100 | Item?: | Vishay 100 (RHK02)
Trail1 | 0.00013765 | Trail2 | -6.04e-05
at | Thu Jun 11 13:45:11 2026 | Vishay 100 (RHK02) | reading is:
Resistance_mean: | 2.04e-05
Resistance_stdevp: | 8.34e-05
Time taken = | 6.1412 | minutes
----------------
Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 10000 | Item?: | Vishay 10k (UBCAL10k)
Trail1 | -0.001079 | Trail2 | 7.49e-05
at | Thu Jun 11 13:52:44 2026 | Vishay 10k (UBCAL10k) | reading is:
Resistance_mean: | 2e+36
Resistance_stdevp: | 1.41e+37
Time taken = | 6.0661 | minutes
----------------
Resistor Vishay 100 (RHK02)-right me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 100 | Item?: | Vishay 100 (RHK02)-right
Trail1 | 100.0153 | Trail2 | 100.0152
at | Thu Jun 11 14:00:58 2026 | Vishay 100 (RHK02)-right | reading is:
Resistance_mean: | 100.0152
Resistance_stdevp: | 9.92e-05
Time taken = | 6.1391 | minutes
----------------
Resistor 1e4 me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 10000 | Item?: | 10000
Trail1 | 10000.315 | Trail2 | 10000.3179
at | Thu Jun 11 14:22:52 2026 | 10000 | reading is:
Resistance_mean: | 10000.3143
Resistance_stdevp: | 0.0017302
Time taken = | 6.1372 | minutes
----------------
Resistor Vishay 10k (UBCAL10k)-right me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: |  | Item?: | Vishay 10k (UBCAL10k)-right
Trail1 | 10000.3177 | Trail2 | 10000.3305
at | Thu Jun 11 14:30:33 2026 | Vishay 10k (UBCAL10k)-right | reading is:
Resistance_mean: | 10000.3213
Resistance_stdevp: | 0.009968
Time taken = | 6.3395 | minutes
----------------
Resistor Vishay 10k (UBCAL10k) me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 10000 | Item?: | Vishay 10k (UBCAL10k)
Trail1 | 10000.3167 | Trail2 | 10000.3137
at | Thu Jun 11 14:38:08 2026 | Vishay 10k (UBCAL10k) | reading is:
Resistance_mean: | 10000.3147
Resistance_stdevp: | 0.001471
Time taken = | 6.1396 | minutes
----------------
Resistor Vishay 100k me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.7
Range?: | 100000 | Item?: | Vishay 100k
Trail1 | 100002.617 | Trail2 | 100002.6311
at | Thu Jun 11 14:46:29 2026 | Vishay 100k | reading is:
Resistance_mean: | 100002.594
Resistance_stdevp: | 0.030633
Time taken = | 6.3419 | minutes
----------------
Resistor UB G1 me]asured by HP3458A with ID number 5
mains frequency = | 50.1
temperature = | 32.8
Range?: | 100000 | Item?: | UB G1
Trail1 | 99998.2681 | Trail2 | 99998.2741
at | Thu Jun 11 14:56:11 2026 | UB G1 | reading is:
Resistance_mean: | 99998.2436
Resistance_stdevp: | 0.016354
Time taken = | 6.341 | minutes
----------------
Resistor UB G2 me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.8
Range?: | 100000 | Item?: | UB G2
Trail1 | 99998.8067 | Trail2 | 99998.8133
at | Thu Jun 11 15:04:25 2026 | UB G2 | reading is:
Resistance_mean: | 99998.793
Resistance_stdevp: | 0.014888
Time taken = | 6.3435 | minutes
----------------
Resistor UB R4A me]asured by HP3458A with ID number 5
mains frequency = | 50.1
temperature = | 32.8
Range?: | 1000000 | Item?: | UB R4A
Trail1 | 1000073.202 | Trail2 | 1000073.794
at | Thu Jun 11 15:10:57 2026 | UB R4A | reading is:
Resistance_mean: | 1000073.316
Resistance_stdevp: | 0.83842
Time taken = | 4.3023 | minutes
----------------
Resistor UB R2B me]asured by HP3458A with ID number 5
mains frequency = | 49.9
temperature = | 32.8
Range?: | 10000 | Item?: | UB R2B
Trail1 | 10000.583 | Trail2 | 10000.5816
at | Thu Jun 11 15:22:08 2026 | UB R2B | reading is:
Resistance_mean: | 10000.5803
Resistance_stdevp: | 0.0015774
Time taken = | 6.144 | minutes
----------------
Resistor UB R4C me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.8
Range?: | 100000 | Item?: | UB R4C
Trail1 | 99999.5552 | Trail2 | 99999.54
at | Thu Jun 11 15:32:49 2026 | UB R4C | reading is:
Resistance_mean: | 99999.5235
Resistance_stdevp: | 0.014609
Time taken = | 6.3436 | minutes
----------------
Resistor coaxial zero me]asured by HP3458A with ID number 5
mains frequency = | 50
temperature = | 32.8
Range?: | 1e-12 | Item?: | coaxial zero
Trail1 | 5.99e-06 | Trail2 | -7.44e-06
at | Thu Jun 11 15:41:13 2026 | coaxial zero | reading is:
Resistance_mean: | 2.97e-07
Resistance_stdevp: | 8.23e-06
Time taken = | 6.1335 | minutes
----------------
end of the measurement

## Short, Open, Phase
Calibration of Universal Bridge
Shorts and Opens
Date: | 2026-06-15
Frequency | 1590
| Shorts |  |  |  | error, ppm | error, ppm
O bservation No. | Range | L | R |  | L | R | Temp
21 | 7Z | -2e-06 | -5e-07 | coaxial | -2 | -0.5 | 20.97
20 | 6Z | -1e-06 | -1.2e-06 | coaxial | -1 | -1.2 | 20.98
19 | 5Z | 0 | -8.6e-06 | coaxial | 0 | -8.6 | 20.97
18 | 4Z | -1e-06 | -5e-07 | coaxial | -1 | -0.5 | 21.12
17 | 3Z | -2e-06 | -7e-07 | coaxial | -2 | -0.7 | 21.07
16 | 2Z | -3e-06 | -7e-07 | coaxial | -3 | -0.7 | 21.04 | flying lead zeros used with ESI
15 | 1Z | -1.8e-05 | -1.8e-06 | coaxial | -18 | -1.8 | 21.12 | boxes
O bservation No. | Opens | C | G |  | C | G
28 | 7Y | 6e-06 | 4.5e-06 | coaxial | 6 | 4.5 | 21.04
27 | 6Y | 0 | 1.7e-06 | coaxial | 0 | 1.7 | 21.06
26 | 5Y | 0 | 4e-07 | coaxial | 0 | 0.4 | 21.05
25 | 4Y | 0 | 9e-07 | coaxial | 0 | 0.9 | 21.23
24 | 3Y | 0 | 9e-07 | coaxial | 0 | 0.9 | 21.26
23 | 2Y | 0 | -1e-07 | coaxial | 0 | -0.1 | 21.36
22 | 1Y | 0 | -1e-07 | coaxial | 0 | -0.1 | 21.5
Phase angle standards | [5 empty] | Y |  |  |  | Z
O bservation No. | Range | Frequency | Standard | value (pF) | Range | C | G |  | Range | L | R
29/30 | 6Y | 1594 | 100k No 4 | 0.014 | 6 | 2.7e-05 | 1.0002 |  | 6Z | -3.9e-05 | 0.99977
31/32 | 6Y | 1594 | 100k No 4 | 0.014 |  | na | na |  | 6Z | na | na
33/34 | 5Y | 1594 | 10k No 4 | 0.028 | 5 | -1.2e-05 | 0.99926 |  | 5Z | 0 | 1.0007
35/36 | 5Y | 1594 | 10k No 4 | 0.028 |  | na | na |  | 5Z | na | na
37/38 | 4Y | 1594 | 1k No 4 | 0.01 | 4 | -1.7e-05 | 1.0023 |  | 4Z | 8e-06 | 0.99768
39/40 | 4Y | 1594 | 1k No 4 | 0.01 |  | na | na |  | 4Z | na | na
41/42 | 3Y | 1594 | 100 ohm No 2 | 0 | 3 | -4e-06 | 0.99662 |  | 3Z | -3e-06 | 1.0033
43/44 | 3Y | 1594 | 100 ohm No 2 | 0 |  | na | na |  | 3Z | na | na
|  | Phase error, pF |  |  | Phase error ppm of fsd |  | phase error
Range | Full scale pF | Y ranges | Z ranges | Frequency | Y ranges | Z ranges | of ratio G1/G2
6 | 1000 | 0.013 | -0.025 | 1594 | 13 | -25 | -6
6 | 1000 | #VALUE! | #VALUE! | 1594 | #VALUE! | #VALUE! | #VALUE!
5 | 10000 | -0.148 | 0.028 | 1594 | -14.8 | 2.8 | -6
5 | 10000 | #VALUE! | #VALUE! | 1594 | #VALUE! | #VALUE! | #VALUE!
4 | 100000 | -1.71 | 0.81 | 1594 | -17.1 | 8.1 | -4.5
4 | 100000 | #VALUE! | #VALUE! | 1594 | #VALUE! | #VALUE! | #VALUE!
3 | 1000000 | -4 | -3 | 1594 | -4 | -3 | -3.5
3 | 1000000 | #VALUE! | #VALUE! | 1594 | #VALUE! | #VALUE! | #VALUE!
|  |  |  | mean | -5.725 | -4.275 | -5
|  |  |  | stdev | 13.7277 | 14.5413 | 1.2247
similar to previous results, retain the 16 ppm uncertainty allowance.
Phase angle standards |  | With the C value taken by E4890A @15/02/2023
Phase angle standards | [5 empty] | Y |  |  |  | Z
O bservation No. | Range | Frequency | Standard | value (pF) | Range | C | G |  | Range | L | R
29/30 | 6Y | 1594 | 100k#2 | 0.037 | 6 | -2e-05 | 0.99793 | 20.96C | 6Z | -1.8e-05 | 1.0019 | 21.01C
31/32 | 6Y | 1594 | 100k#4 | 0.033 |  | -2.3e-05 | 1.0001 | 21.17C | 6Z | -1.4e-05 | 0.99973 | 20.94C
33/34 | 5Y | 1594 | 10k#3 | 0.211 | 5 | -3.8e-05 | 1.0003 | 20.95C | 5Z | 4e-06 | 0.99955 | 20.98C
35/36 | 5Y | 1594 | 10k#4 | 0.199 |  | -3.9e-05 | 0.99911 | 20.99C | 5Z | 4e-06 | 1.0008 | 20.94C
37/38 | 4Y | 1594 | 1k#1 | 1.475 | 4 | -1.8e-05 | 1.0014 | 21.09C | 4Z | 1e-05 | 0.99861 | 21.04C
39/40 | 4Y | 1594 | 1k#4 | 1.723 |  | -1.7e-05 | 1.0023 | 20.96C | 4Z | 1e-05 | 0.9977 | 20.96C
41/42 | 3Y | 1594 | 100#2 | 19.238 | 3 | -6e-06 | 0.99658 | 21.16C | 3Z | -2e-06 | 1.0034 | 21.02C
43/44 | 3Y | 1594 | 100#5 | 18.122 |  | -7e-06 | 0.99693 | 21.14C | 3Z | -1e-06 | 1.0031 | 20.96C
|  | Phase error, pF |  |  | Phase error ppm of fsd |  | phase error
Range | Full scale pF | Y ranges | Z ranges | Frequency | Y ranges | Z ranges | of ratio G1/G2
6 | 1000 | -0.057 | 0.019 | 1594 | -57 | 19 | -19
| 1000 | -0.056 | 0.019 | 1594 | -56 | 19 | -18.5
5 | 10000 | -0.591 | 0.251 | 1594 | -59.1 | 25.1 | -17
| 10000 | -0.589 | 0.239 | 1594 | -58.9 | 23.9 | -17.5
4 | 100000 | -3.275 | 2.475 | 1594 | -32.75 | 24.75 | -4
| 100000 | -3.423 | 2.723 | 1594 | -34.23 | 27.23 | -3.5
3 | 1000000 | -25.238 | 17.238 | 1594 | -25.238 | 17.238 | -4
| 1000000 | -25.122 | 17.122 | 1594 | -25.122 | 17.122 | -4
|  |  |  | mean | -43.5425 | 21.6675 | -10.9375
|  |  |  | stdev | 15.546 | 3.995 | 7.5755
2023 comment II
Y ranges higher phase error with the calculation C value by LCR E4890A
re-measured Thompsons keep them away from any electronics apparatuses as the table from row 94-row 105
calculated results as Row 109-Row 119
Phase angle standards | [6 empty] | with New UB reading @pm, 15/02/2023
Phase angle standards | [5 empty] | Y |  |  |  | Z
O bservation No. | Range | Frequency | Standard | value (pF) | Range | C | G |  | Range | L | R
45/46 | 6Y | 1594 | 100k#2 | 0.0163 | 6 | 2.1e-05 | 0.99803 | 21.68C | 6Z | -3.4e-05 | 1.002 | 21.86C
47/48 | 6Y | 1594 | 100k#4 | 0.014 |  | 2.1e-05 | 1.0002 | 21.87C | 6Z | -3.1e-05 | 0.99977 | 21.71C
49/50 | 5Y | 1594 | 10k#3 | 0.031 | 5 | -1.1e-05 | 1.0004 | 21.72C | 5Z | 3e-06 | 0.99957 | 21.59C
51/52 | 5Y | 1594 | 10k#4 | 0.028 |  | -1.1e-05 | 0.99922 | 21.50C | 5Z | 3e-06 | 1.0008 | 21.88C
53/54 | 4Y | 1594 | 1k#1 | 0 | 4 | -1.8e-05 | 1.0014 | 21.80C | 4Z | 1e-05 | 0.99862 | 21.52C
55/56 | 4Y | 1594 | 1k#4 | 0.01 |  | 9e-06 | 0.99772 | 21.77C | 4Z | -1.1e-05 | 1.0023 | 21.57C
57/58 | 3Y | 1594 | 100#2 | 0 | 3 | -5e-06 | 0.99657 | 21.57C | 3Z | -2e-06 | 1.0034 | 21.47C
59/60 | 3Y | 1594 | 100#5 | 0 |  | -7e-06 | 0.99693 | 21..57C | 3Z | -1e-06 | 1.0031 | 21.73C
|  | Phase error, pF |  |  | Phase error ppm of fsd |  | phase error
Range | Full scale pF | Y ranges | Z ranges | Frequency | Y ranges | Z ranges | of ratio G1/G2
6 | 1000 | 0.0047 | -0.0177 | 1594 | 4.7 | -17.7 | -6.5
| 1000 | 0.007 | -0.017 | 1594 | 7 | -17 | -5
5 | 10000 | -0.141 | 0.061 | 1594 | -14.1 | 6.1 | -4
| 10000 | -0.138 | 0.058 | 1594 | -13.8 | 5.8 | -4
4 | 100000 | -1.8 | 1 | 1594 | -18 | 10 | -4
| 100000 | 0.89 | -1.09 | 1594 | 8.9 | -10.9 | -1
3 | 1000000 | -5 | -2 | 1594 | -5 | -2 | -3.5
| 1000000 | -7 | -1 | 1594 | -7 | -1 | -4
|  |  |  | mean | -4.6625 | -3.3375 | -4
|  |  |  | stdev | 10.4447 | 10.7427 | 1.5353
Note on A7: Minyu Zhang: O bservation??
Note on E64: Minyu Zhang: reading taken @15/02/23 Rp Cp by E4980A

## Resistors
Calibration of Universal Bridge
Calibration of check resistors using DC
Date: | 2023-02-20 |  |  | Enter data into light blue areas
DMM: |  |  |  | DC resistance data linked to docx page
DMM settings: | NPLC 20, 4 wire ohms | [5 empty] | Correction
GMH51
Observation No | Measurement | OCOMP ON? | DELAY | Mean | sd | sd ppm | temp
61/62 | SR104 StdsR SN:109006 | on |  | 10000.0148 | 0.0018672 | 0.18672 | 21.4 | taken:21/02/2023 | DMM3458MRS.py | 3458A S/N MY45055679
63/64 | SR104 Temp_sensor SN109006 |  |  | 9984.9603 | 0.002167 | 0.21703 | 21.4
65/66 | ESI 1k S NetworkBars | on |  | 9999.6176 | 0.002539 | 0.25391 | 21.38 | Temp. by GMH5I from Notebook: "Minyu II"
67/68 | ESI 1k S//P NetworkBars | on |  | 999.9599 | 0.00034004 | 0.34006 | 21.35
69/70 | ESI R10 only | on |  | 999.972 | 0.00026674 | 0.26675 | 21.59 |  |  |  | 999.972
71/72 | ESI 1k R1 only | on |  | 999.9604 | 0.00026334 | 0.26336 | 21.75
73/74 | ESI 1k P// only | on |  | 99.996 | 0.00019027 | 1.9028 | 21.77
75/76 | ESI 100 Ohm S network bars SN:123009 | on |  | 999.9663 | 0.00022752 | 0.22753 | 21.58
77/78 | ESI 100 Ohm R1 only | on |  | 99.9977 | 0.00015645 | 1.5645 | 21.5
79/80 | eSI 100 Ohm P// network bars | on |  | 9.9996 | 1.8983e-05 | 1.8984 | 21.58
81/82 | ESI 10 Ohm R1 only SN:A1-1752030 | on |  | 10.0001 | 1.6918e-05 | 1.6918 | 21.49
83/84 | ESI 10 Ohm S network bars | on |  | 100.0009 | 0.00015356 | 1.5356 | 21.68
| 10Ohm // | on |  | 1 | 1.0187e-05 | 10.1872 | 21.04
85/86 | 100k Vishay | on |  | 100003.3199 | 0.042298 | 0.42296 | 21.5
87/88 | 10k Vishay | on |  | 10000.3233 | 0.0081507 | 0.81504 | 21.37
89/90 | 100 Vishay | on |  | 100.0148 | 0.00019899 | 1.9896 | 21.46 | [6 empty] | sensor R | Temp | R to add to 20 value
| 1M Vishay | On |  | 1000001.7175 | 0.45496 | 0.45496 | 21.3
[14 empty] | sensor R | Temp | R to add to 20 value
[14 empty] | 9971.4967 | 20.1489 | 0.00050326
Analysis
Resistor | Nominal Value (ohms) | ppm of nominal resistance | ppm of nominal conductance | mean DMM reading (ohms) | stdev (ohms) | DMM correction (ohms) | Value of Resistor (ohms) | Traceability - last cal date | Reference |  |  |  | DMM Corrn ppm
SR104 | 10000 |  |  | 10000.0148 | 0.0018672 | -0.015391 | 9999.9994 | 2026-06-19 | See SR104 worksheet |  |  |  | -1.5391
| 100000 |  |  |  |  | 0.22814 | DMM correction at 100k from "Calibration" worksheet | [5 empty] | 2.2814 | 2005 values
1k ESI, series | 10000 | -39.7787 |  | 9999.6176 | 0.002539 |  | 9999.6022 | [6 empty] | 9999.7443 | -14.2111
1k ESI S/P | 1000 | -40.9892 |  | 999.9599 | 0.00034004 | -0.00092042 | 999.959 | Note: see ESI model SR1010 handbook for series/parallel formula |  |  |  |  | -0.92046 | 999.9745 | -15.5089
1k ESI, R10 | 1000 |  |  | 999.972 | 0.00026674
1k ESI, R1 | 1000 | -40.5635 |  | 999.9604 | 0.00026334 |  | 999.9594 | [6 empty] | 999.9761 | -16.6632
100 ohm ESI, series | 1000 | -34.5819 |  | 999.9663 | 0.00022752 |  | 999.9654 | [6 empty] | 999.9723 | -6.8814
1k ESI, parallel | 100 |  |  | 99.996 | 0.00019027 | 3.9034e-05 | 99.996 | [5 empty] | 0.39036 | 99.9974 | -14.2111
100 ohm ESI, R1 | 100 | -22.2788 |  | 99.9977 | 0.00015645 |  | 99.9978 | [6 empty] | 99.9975 | 2.5093
100 ohm ESI, parallel | 10 | -34.5819 |  | 9.9996 | 1.8983e-05 | 1.2396e-05 | 9.9997 | [5 empty] | 1.2397 | 9.9997 | -6.8814
10 ohm ESI, R1 | 10 | 9.6918 |  | 10.0001 | 1.6918e-05 |  | 10.0001 | [6 empty] | 10.0001 | 2.3924
10 ohm ESI, series | 100 |  |  | 100.0009 | 0.00015356 |  | 100.001 | [6 empty] | 99.9993 | 16.9445
10 ohm ESI, para | 1 |  |  | 1 | 4.2298e-07 |  | 1 | [6 empty] | 0.99999 | 16.9445
100k Vishay | 100000 |  |  | 100003.3199 | 0.042298 |  | 100003.5481 | [6 empty] | 100002.938 | 6.1007
10 k Vishay | 10000 |  |  | 10000.3233 | 0.0081507 |  | 10000.3079 | [6 empty] | 10000.3196 | -1.169
100 ohm Vishay | 0 |  |  | 100.0148 | 0.00019899 |  | 100.0148 | [6 empty] | 100.012 | 27.7414
New build up to 1M Vishay on 31 July 2014
SR104 | 10000 |  |  | 10000.0123 | 0.0025697 | -0.013698 | 9999.9986 | Wed Mar 15 09:54:32 2023 | See SR104 worksheet
100k esi parallel | 10000 |  |  | 10000.1341 | 0.0018379 | -0.013698 | 10000.1204
100k esi series | 1000000 |  |  | 1000015.8203 | 0.54195 | -3.776 | 1000012.0443
1M Vishay | 1000000 |  |  | 1000001.7175 | 0.45496 |  | 999997.9415 | [6 empty] | sensor R | Temp | R to add to 20 value
[14 empty] | 9969.188 | 19.918 | -0.00028216
0.01
Date: | 2014-07-22 | [6 empty] | -0.0212
[5 empty] | Y |  |  | Z
Observation No | Range | Nominal value ohms | Resistor | Frequency | C | G | Range | L | R | temp | DC resistance | Corrected UB G | Corrected UB R | Gac -Gdc ppm | Rac - Rdc ppm |  | Ratio error of G1/G2 | phase error of G1/G2
117/118 | 7 | 1000000 | Vishay | 1592 | 0.000208 | 1.0001 | 21.23 | -0.000212 | 0.99992 | 21.12 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -2
115/116 |  | 1000000 | Vishay | 160 | 8.4e-05 | 1 | 21.23 | -7.2e-05 | 0.99994 | 21.12 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | 6
113/114 | 6 | 100000 | Vishay | 1592 | 0.000625 | 0.99997 | 21.29 | -0.000636 | 1 | 21.16 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -5.5
111/112 |  | 100000 | Vishay | 160 | 0.000625 | 0.99996 | 21.29 | -0.000666 | 1 | 21.16 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -20.5
190/110 | 5 | 10000 | Vishay | 1592 | 5.1e-05 | 1 | 21.61 | -6.2e-05 | 0.99997 | 21.5 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -5.5
107/108 |  | 10000 | Vishay | 160 | 5.1e-05 | 1 | 21.61 | -5.9e-05 | 0.99997 | 21.5 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -4
105/106 | 4 | 1000 | ESI 1000 R1 | 1592 | -0.000201 | 1.0001 | 21.36 | 0.000191 | 0.99989 | 21.35 | 0.99996 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -5
103/104 |  | 1000 | ESI 1000 R1 | 160 | -0.000203 | 1.0001 | 21.36 | 0.000196 | 0.9999 | 21.35 | 0.99996 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -3.5
101/102 | 3 | 100 | ESI 100 R1 | 1592 | -0.002461 | 1 | 21.32 | 0.002455 | 0.99995 | 21.29 | 0.99998 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -3
99/100 |  | 100 | ESI 100 R1 | 160 | -0.002489 | 1.0001 | 21.32 | 0.002482 | 0.99993 | 21.29 | 0.99998 | #N/A | #N/A | #N/A | #N/A |  | #N/A | -3.5
97/98 | 2 | 10 | ESI 10 R1 | 1592 | -0.000306 | 1 | 21.34 | 0.000312 | 0.99998 | 21.12 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | 3
95/96 |  | 10 | ESI 10 R1 | 160 | -3.3e-05 | 1 | 21.34 | 0.000282 | 1 | 21.12 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | 124.5
93/94 | 1 | 1 | ESI 10 para | 1592 | -1.5e-05 | 1.0001 | 21.35 | 0.000395 | 0.99997 | 21.25 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | 190
91/92 |  | 1 | ESI 10 para | 160 | 1.9e-05 | 1 | 21.3 | 0.00035 | 0.99995 | 21.26 | 1 | #N/A | #N/A | #N/A | #N/A |  | #N/A | 184.5
[13 empty] | average: | #N/A | #N/A | average: | #N/A | 37.625
[13 empty] | std dev: | #N/A | #N/A | std dev: | #N/A | 79.325
| test_c | 10.97
[14 empty] | #N/A | average +/- error
Using test_c |  | correction to | correction to
|  | L/C dial | R/G dial
Range | [12 empty] | Comment, 2014 calibration
|  | (ppm) | (ppm) | [9 empty] | The 'c' term has increased significantly(from 2 to 11 ppm) but the standard deviation in the ac-dc differences has decreased.
1Z |  | 216.0195 | 74.0523 | [9 empty] | c' was measured as 5.7 ppm in 2005, so it is not clear any repair is required, set to 10.97 ppm.
2Z |  | 216.0195 | 74.0523 | [9 empty] | This suggests the calibration method is performing well with more careful use of the HP3458 (eg measuring frequency each time for NPLC)
3Z |  | 216.0195 | 74.0523 | [9 empty] | The increase in 'c' is a concern
4Z |  | 216.0195 | 74.0523 | [9 empty] | G1/G2 ratio scatter has increased.
5Z |  | 216.0195 | 74.0523 | [9 empty] | Scatter in Gac-Gdc, and Rac-Rdc, is 8 ppm, down from 17.5 ppm.
6Z |  | 155.9473 | 13.9801 | [9 empty] | Cautiously leave the uncertainty at 17 ppm as in 2009.
7Z |  | 231.1195 | 89.1523
1Y |  | 78.9023 | -63.0649 | [9 empty] | Comment, 2009 calibration
2Y |  | 78.9023 | -63.0649 | [9 empty] | The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
3Y |  | 78.9023 | -63.0649 | [9 empty] | The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
4Y |  | 78.9023 | -63.0649 | [9 empty] | However, the scatter in Gac-Gdc, and Rac-Rdc, is 17.5 ppm with an average error of 2 ppm in G, or -2 in R. Note the 2005 version has an error in the 100 k dc value
5Y |  | 78.9023 | -63.0649 | [9 empty] | The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
6Y |  | 138.9745 | -2.9927 | [9 empty] | errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
7Y |  | 63.8023 | -78.1649 | [9 empty] | of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 2.0 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 17.5
Hence, increase "amplifier ratio" in the uncertainty budget to 17 ppm.
Comment, 2005 calibration
The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
However, the scatter in Gac-Gdc, and Rac-Rdc, is 13.8 ppm with an average error of 5.7 ppm in G, or -5.7 in R.
The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 5.7 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 13.9589
Hence, increase "amplifier ratio" in the uncertainty budget to 14 ppm.
Comment, 2004 calibration
The uncertainty component "amplifier ratio" arises from the scatter in G1/G2 ratio error.
The historic 1-sigma estimate for this is 2.5 ppm, which is slightly conservative. Reduce to 2.0 ppm.
However, the scatter in Gac-Gdc, and Rac-Rdc, is 6.3 ppm with an average error of 5.2 ppm in G, or -5.2 in R.
The cause of this scatter and offset has not been established despite investigation, but it is believed to mainly be due to
errors on the RHS of the bridge, rather than the amplifier ratio. It can be regarded as scatter in the apparent value
of R_4, or the ratio of the RHS transformers on ranges 1, 2, 3, and 4.
Apply a correction "c" (see worksheet Corrections) of 5.2 ppm.
The most convenient way to account for the RHS scatter is to regard it as being lumped in with the "amplifier ratio".
[13 empty] | The rss of the true amplifier ratio scatter and RHS scatter is: |  |  |  | 6.2642
Hence, increase "amplifier ratio" in the uncertainty budget to 7 ppm.
Set 'test_c' to give a zero average +/- error.
Note on A8: Minyu: Load from the docx sheet and not books copies
Note on Q59: Keith Jones: Range 7 is a new addition to the calibration process. Values not used for setting c.

## L and C
Calibration of Universal Bridge
Measurements on MSL inductors and capacitors
Date: | 2023-03-08
| Inductors | [15 empty] | 2004-02-18
Observation No. | Range | Nominal value microhenries | GR inductor S/N | Frequency | L | R | Lo | Ro | temp | Full scale inductance (uH) | Full scale resistance (ohms) | Corrected L dial | Corrected R dial | L, millihenires | R, ohms |  | L, millihenires | R, ohms | ppm shift in L
128/129 | 1Z | 100 | 18290 | 1000 | 1.0003 | 0.086209 | 0.001427 | 0.004331 | 20.96 | 100 | 1 | 0.99904 | 0.081884 | 0.099904 | 0.081884 |  | 0.099907 | 0.081795 | 27.0632
130/131 | 2Z | 1000 | 18142 | 1000 | 0.99875 | 0.086265 | 2.2e-05 | 0.0009011 | 21.05 | 1000 | 10 | 0.99894 | 0.085371 | 0.99894 | 0.85371 |  | 0.99901 | 0.85334 | 68.12
132/133 | 3Z | 10000 | 18413 | 1000 | 0.99975 | 0.08522 | 1e-06 | 0.0003253 | 20.76 | 10000 | 100 | 0.99997 | 0.084901 | 9.9997 | 8.4901 |  | 9.9997 | 8.4806 | 6.9997
134/135 | 4Z | 100000 | 1879 | 1000 | 0.9996 | 0.083491 | 0 | 1.07e-05 | 20.8 | 100000 | 1000 | 0.99982 | 0.083487 | 99.9821 | 83.487 |  | 99.9815 | 83.1956 | -5.9986
136/137 | 5Z | 1000000 | 18430 | 1000 | 1.002 | 0.064334 | 0 | -1.3e-06 | 20.91 | 1000000 | 10000 | 1.0023 | 0.06434 | 1002.2605 | 643.4016 |  | 1002.2554 | 641.3039 | -5.0762
138/139 | 5Z | 1000000 | 18430 | 107 | 0.99972 | 0.063789 | 0 | -1.3e-06 | 2096 | 1000000 | 10000 | 0.99994 | 0.063795 | 999.936 | 637.9509 |  | 999.935 | 635.4936 | -1.001
Date: | 2014-08-01
| Capacitors | [15 empty] | Value | Last cal date | Reference
Observation No. | Range | Nominal value pF | Capacitor | Frequency | C | G |  |  |  | Full scale capacitance (pF) | Full scale conductance (uS) | Corrected C dial | G dial | C, (pF) | tan delta (urad) | Capacitance error ppm |  | Traceability
122 | 4Y | 100000 | Sullivan 57709/1960 | 1000 | 1.0014 | 0.0055362 | 21 |  |  | 100000 | 1000 | 1.0015 | 0.0055359 | 100146.3012 | 8810.5803 | #DIV/0!
123 | 3Y | 1000000 | Sullivan 57708/1960 | 1000 | 1.0036 | 0.011207 | 20.97 |  |  | 1000000 | 10000 | 1.0037 | 0.011206 | 1003717.1893 | 17834.733 | #DIV/0!
119 | 6Y | 1000 | GR 1000A | 1592 | 0.99987 | -1.4e-05 | 23.605 |  |  | 1000 | 10 | 1 | -1.4e-05 | 1000.013 | -13.996 | -16.1925 | 1000.0292 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
120 | 6Y | 1000 | GR1000B | 1592 | 0.99986 | -1.3e-05 | 23.605 |  |  | 1000 | 10 | 0.99999 | -1.3e-05 | 999.995 | -12.9963 | -17.4453 | 1000.0124 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
121 | 7Y | 100 | GR100 | 1592 | 0.99992 | 0.000259 | 23.605 |  |  | 100 | 1 | 0.99998 | 0.00025898 | 99.9982 | 258.9065 | -54.7609 | 100.0037 | 2014-07-11 | Y:\Ongoing\Farad\CalcCap\Cap2014\Simple_buildup_July2014.xls
124 | 3Y | 1000000 | GR box 3024 10x100nF | 1000 | 0.98236 | -0.031129 | 21.26 |  |  | 1000000 | 10000 | 0.98244 | -0.031127 | 982393.057 | -49540.5361 | #DIV/0!
125 | 4Y | 100000 | GR box 3024 10x10nF | 1000 | 0.99747 | -0.0087993 | 21.23 |  |  | 100000 | 1000 | 0.99754 | -0.0087987 | 99705.9164 | -14003.6377 | #DIV/0!
126 | 5Y | 10000 | GR box 3024 10x1nF | 1000 | 1.0054 | 0.0002107 | 21.2 |  |  | 10000 | 100 | 1.0055 | 0.00021069 | 10006.0594 | 335.3183 | #DIV/0!
127 | 5Y | 0 | GR box 3024 dials zero | 1000 | 0.004845 | 1.97e-05 | 21.21 |  |  | 10000 | 100 | 0.0048454 | 1.9699e-05 | 48.4538
box zero subtracted
note that the AH2700A is two terminal-pair and expect lead corrections of the order of 50 ppm for 1 microfarad load.
With correction term "c"
|  |  | correction to L/C dial | correction to R/G dial
Range
|  |  | (ppm) | (ppm)
| 1Z |  | 216.0195 | 74.0523
| 2Z |  | 216.0195 | 74.0523
| 3Z |  | 216.0195 | 74.0523
| 4Z |  | 216.0195 | 74.0523
| 5Z |  | 216.0195 | 74.0523
| 6Z |  | 155.9473 | 13.9801
| 7Z |  | 231.1195 | 89.1523
| 1Y |  | 78.9023 | -63.0649
| 2Y |  | 78.9023 | -63.0649
| 3Y |  | 78.9023 | -63.0649
| 4Y |  | 78.9023 | -63.0649
| 5Y |  | 78.9023 | -63.0649
| 6Y |  | 138.9745 | -2.9927
| 7Y |  | 63.8023 | -78.1649
Note on T10: Keith Jones: corrected transcription error in old spreadsheet

## Results
DC calibration of resistors
Name | Value | Uncertainty |  |  |  |  | ppm from nominal | nominal
| ohm | ohm
true_sr104_t | 10000.0018 | 0.001 |  |  |  |  | 0.18167 | 10000
true_ep_100k | 10000.1294 | 0.0056281 |  |  |  |  | 12.9427 | 10000
true_esp_100k | 100001.3118 | 0.056362 |  |  |  |  | 13.1182 | 100000
true_es_100k | 1000012.9427 | 0.56281 |  |  |  |  | 12.9427 | 1000000
true_g1 | 99998.3734 | 0.063032 |  |  |  |  | -16.2659 | 100000
true_g2 | 99998.9742 | 0.062308 |  |  |  |  | -10.2576 | 100000
true_r4a | 10000.5358 | 0.0069259 |  |  |  |  | 53.5786 | 10000
true_r4b | 99999.4273 | 0.063662 |  |  |  |  | -5.7269 | 100000
true_r4c | 1000073.3126 | 1.9596 |  |  |  |  | 73.3126 | 1000000
true_sr104b_t | 10000.0029 | 0.001 |  |  |  |  | 0.28701 | 10000
true_es_1k | 9999.6057 | 0.0033065 |  |  |  |  | -39.434 | 10000
true_esp_1k | 999.9594 | 0.00033346 |  |  |  |  | -40.6445 | 1000
true_er1_1k | 999.9598 | 0.00052254 |  |  |  |  | -40.2188 | 1000
true_es_100 | 999.9658 | 0.00050543 |  |  |  |  | -34.2372 | 1000
true_ep_1k | 99.9961 | 3.3065e-05 |  |  |  |  | -39.434 | 100
true_ep_100 | 9.9997 | 5.0543e-06 |  |  |  |  | -34.2372 | 10
true_er1_100 | 99.9978 | 0.00024854 |  |  |  |  | -21.9341 | 100
true_er1_10 | 10.0001 | 2.5925e-05 |  |  |  |  | 10.0365 | 10
true_es_10 | 100.001 | 0.00024673 |  |  |  |  | 9.9011 | 100
true_v_100k | 100003.5119 | 0.074726 |  |  |  |  | 35.1186 | 100000
true_v_10k | 10000.3114 | 0.0084214 |  |  |  |  | 31.1357 | 10000
true_v_100 | 100.0148 | 0.0002773 |  |  |  |  | 148.3014 | 100
true_sr104c_t | 10000.0029 | 0.001 |  |  |  |  | 0.28701 | 10000
true_epc_100k | 10000.1222 | 0.0030123 |  |  |  |  | 12.2196 | 10000
true_esc_100k | 1000012.2196 | 0.30123 |  |  |  |  | 12.2196 | 1000000
true_v_1M | 999998.1168 | 0.76905 |  |  |  |  | -1.8832 | 1000000
Phase angle corrections from Thompson set
Label | Range | Freq/Hz | F | H | ppm C | ppm L
thomp_6a | 6 | 1592 | 3.7e-14 | 1.5064e-07 | 37 | 0.015064
thomp_5a | 5 | 1592 | 1.8e-14 | -6.8085e-06 | 1.8 | -6.8085
thomp_4a | 4 | 1592 | 1.71e-12 | -1.0099e-06 | 17.1 | -10.0991
thomp_3a | 3 | 1592 | 6e-12 | -3e-08 | 6 | -3
Gain Factor
Range | prod / ppm | factor / ppm | Freq / Hz
7 | -12.6021 | -6.5936 | 1592
7 | -9.751 | -3.7425 | 160
6 | -8.0502 | -2.0417 | 1592
6 | -4.6005 | 1.4079 | 160
5 | -6.7503 | -0.74184 | 1592
5 | -5.2503 | 0.75816 | 160
4 | -6.3557 | -0.34729 | 1592
4 | -3.3051 | 2.7034 | 160
3 | -7.7509 | -1.7424 | 1592
3 | -5.5518 | 0.45668 | 160
2 | 11.4994 | 17.5078 | 1592
2 | 23.1498 | 29.1583 | 160
1 | 23.1482 | 29.1567 | 1592
1 | -17.8012 | -11.7928 | 160
UB measured values
Label | G / siemen | R / ohm |  |  |  |  | 1/G | R | GxR | DC R | rel to 1/G | rel to R
v_1M_1592 | 9.9998e-07 | 1000002.2943 |  |  |  |  | 1000015.4789 | 1000002.2943 | 0.99999 | 999998.1168 | 17.3621 | 4.1774
v_1M_160 | 9.9997e-07 | 1000025.2961 |  |  |  |  | 1000032.7785 | 1000025.2961 | 0.99999 | 999998.1168 | 34.6618 | 27.1793
v_100k_1592 | 9.9999e-06 | 100001.068 |  |  |  |  | 100001.4761 | 100001.068 | 1 | 100003.5119 | -20.3574 | -24.4379
v_100k_160 | 9.9997e-06 | 100002.788 |  |  |  |  | 100002.5061 | 100002.788 | 1 | 100003.5119 | -10.0572 | -7.2385
v_10k_1592 | 9.9997e-05 | 10000.2859 |  |  |  |  | 10000.3007 | 10000.2859 | 1 | 10000.3114 | -1.0634 | -2.5442
v_10k_160 | 9.9997e-05 | 10000.3009 |  |  |  |  | 10000.2857 | 10000.3009 | 1 | 10000.3114 | -2.5633 | -1.0441
er1_1k_1592 | 0.0010001 | 999.9461 |  |  |  |  | 999.9468 | 999.9461 | 1 | 999.9598 | -13.0036 | -13.6953
er1_1k_160 | 0.001 | 999.9558 |  |  |  |  | 999.9504 | 999.9558 | 1 | 999.9598 | -9.404 | -3.9943
er1_100_1592 | 0.0099999 | 100.001 |  |  |  |  | 100.0013 | 100.001 | 1 | 99.9978 | 35.0072 | 31.5251
er1_100_160 | 0.01 | 99.9995 |  |  |  |  | 99.9994 | 99.9995 | 1 | 99.9978 | 15.6076 | 16.5239
er1_10_1592 | 0.1 | 10.0004 |  |  |  |  | 10 | 10.0004 | 1 | 10.0001 | -8.6637 | 26.3552
er1_10_160 | 0.099999 | 10.0006 |  |  |  |  | 10.0001 | 10.0006 | 1.0001 | 10.0001 | -4.8639 | 53.4565
ep_10_1592 | 1 | 1 |  |  |  |  | 0.99996 | 1 | 1.0001 | 1 | -45.4249 | 12.8899
ep_10_160 | 0.99998 | 0.99999 |  |  |  |  | 1 | 0.99999 | 0.99998 | 1 | 6.2711 | -17.3116
Inductor set
Label | L / henry | R / ohm |  |  |  |  | L / microhenry
gr100uh_1k | 9.9904e-05 | 0.081882 |  |  |  |  | 99.9039
gr1000uh_1k | 0.00099894 | 0.85369 |  |  |  |  | 998.943
gr10000uh_1k | 0.0099996 | 8.49 |  |  |  |  | 9999.6418
gr100000uh_1k | 0.099982 | 83.4858 |  |  |  |  | 99981.9145
gr100000uh_100 | 1.0023 | 643.3923 |  |  |  |  | 1002258.6678
Capacitor set
Label | C / farad | G / siemen |  |  |  |  | C / pF
su100000_1k | 1.0015e-07 | 5.5359e-06 |  |  |  |  | 100149.1198
su1000000_1k | 1.0037e-06 | 0.00011206 |  |  |  |  | 1003745.4391
gr1000a_1592 | 1e-09 | -1.4e-10 |  |  |  |  | 1000.0403
gr1000b_1592 | 1e-09 | -1.3e-10 |  |  |  |  | 1000.0223
gr100_1592 | 1e-10 | 2.5898e-10 |  |  |  |  | 100.0005
gr_10x100_1k | 9.8247e-07 | -0.00031128 |  |  |  |  | 982469.1617
gr_10x10n_1k | 9.9757e-08 | -8.7989e-06 |  |  |  |  | 99757.1778
gr_10x1n_1k | 1.0055e-08 | 2.1069e-08 |  |  |  |  | 10054.7962
gr_0_1k | 4.8455e-11 | 1.9699e-09 |  |  |  |  | 48.4552

## csv_data
Resistor cal set a | Value | Stdev | dof | Comment
true_sr104_20 | 10000.0027 | 0.000145 | 50 | cal_cert interpretation
sens_sr104 | 9971.2359 | 0.019186 | 19
sr104 | 10000.0076 | 0.0017328 | 19
ep_100k | 10000.1244 | 0.0019637 | 19
esp_100k | 100000.9992 | 0.015814 | 19
er10_100k | 100000.7244 | 0.014969 | 19
es_100k | 1000012.626 | 3.2902 | 19
g1 | 99998.2436 | 0.016354 | 19
g2 | 99998.793 | 0.014888 | 19
r4a | 10000.5803 | 0.0015774 | 19
r4b | 99999.5235 | 0.014609 | 19
r4c | 1000073.316 | 0.83842 | 19
Zeros | dial a | dial b
zero7z_1592 | -2e-06 | -5e-07
zero6z_1592 | -1e-06 | -1.2e-06
zero5z_1592 | 0 | -8.6e-06
zero4z_1592 | -1e-06 | -5e-07
zero3z_1592 | -2e-06 | -7e-07
zero2z_1592 | -3e-06 | -7e-07
zero1z_1592 | -1.8e-05 | -1.8e-06
zero7y_1592 | 6e-06 | 4.5e-06
zero6y_1592 | 0 | 1.7e-06
zero5y_1592 | 0 | 4e-07
zero4y_1592 | 0 | 9e-07
zero3y_1592 | 0 | 9e-07
zero2y_1592 | 0 | -0.1 |  | flying leads
zero1y_1592 | 0 | -0.1 |  | flying leads
zero7z_160 | 0 | 0 |  | not measured at 160 Hz!
zero6z_160 | 0 | 0 |  | not measured at 160 Hz!
zero5z_160 | 0 | 0 |  | not measured at 160 Hz!
zero4z_160 | 0 | 0 |  | not measured at 160 Hz!
zero3z_160 | 0 | 0 |  | not measured at 160 Hz!
zero2z_160 | 0 | 0 |  | not measured at 160 Hz!
zero1z_160 | 0 | 0 |  | not measured at 160 Hz!
zero7y_160 | 0 | 0 |  | not measured at 160 Hz!
zero6y_160 | 0 | 0 |  | not measured at 160 Hz!
zero5y_160 | 0 | 0 |  | not measured at 160 Hz!
zero4y_160 | 0 | 0 |  | not measured at 160 Hz!
zero3y_160 | 0 | 0 |  | not measured at 160 Hz!
zero2y_160 | 0 | 0 |  | not measured at 160 Hz!
zero1y_160 | 0 | 0 |  | not measured at 160 Hz!
Thompson ac Cal | dial a | dial b | cap pF
thomp_100k_4_z_1592 | na | na | 0.014
thomp_100k_4_y_1592 | na | na | 0.014
thomp_10k_4_z_1592 | na | na | 0.028
thomp_10k_4_y_1592 | 0 | na | 0.028
thomp_1k_4_z_1592 | na | na | 0.01
thomp_1k_4_y_1592 | na | na | 0.01
thomp_100_2_z_1592 | -3e-06 | 1.0033 | 0
thomp_100_2_y_1592 | -4e-06 | 0.99662 | 0
thomp_100k_2_z_160 | 0 | 0 | 0.014 | 160 Hz not measured!
thomp_100k_4_y_160 | 0 | 0 | 0.014 | 161 Hz not measured!
thomp_10k_3_z_160 | 0 | 0 | 0.028 | 162 Hz not measured!
thomp_10k_4_y_160 | 0 | 0 | 0.028 | 163 Hz not measured!
thomp_1k_4_z_160 | 0 | 0 | 0.01 | 164 Hz not measured!
thomp_1k_4_y_160 | 0 | 0 | 0.01 | 165 Hz not measured!
thomp_100_2_z_160 | 0 | 0 | 0 | 166 Hz not measured!
thomp_100_2_y_160 | 0 | 0 | 0 | 167 Hz not measured!
Resistor cal set b | Value | Stdev | dof | Comment
sens_sr104b | 9984.9603 | 0.002167 | 19 | assume true value of sr104 as in set1, just different temperature
sr104b | 10000.0148 | 0.0018672 | 19
es_1k | 9999.6176 | 0.002539 | 19
esp_1k | 999.9599 | 0.00034004 | 19
er10_1k | 999.972 | 0.00026674 | 19
er1_1k | 999.9604 | 0.00026334 | 19
es_100 | 999.9663 | 0.00022752 | 19
ep_1k | 99.996 | 0.00019027 | 19
er1_100 | 99.9977 | 0.00015645 | 19
ep_100 | 9.9996 | 1.8983e-05 | 19
er1_10 | 10.0001 | 1.6918e-05 | 19
es_10 | 100.0009 | 0.00015356 | 19
v_100k | 100003.3199 | 0.042298 | 19
v_10k | 10000.3233 | 0.0081507 | 19
v_100 | 100.0148 | 0.00019899 | 19
Resistor cal set c | Value | Stdev | dof | Comment
sr104c | 10000.0148 | 0.002167 | 19 | not measured, use the sr104b value
sens_sr104c | 9984.9603 | 0.0018672 | 19
ep_100k | 10000.1341 | 0.0018379 | 19
es_100k | 1000015.8203 | 0.54195 | 19
v_1M | 1000001.7175 | 0.45496 | 19
Resistor ac Cal | Y dial a | Y dial b | Z dial a | Z dial b | Range | Frequency
v_1M_1592 | 0.000208 | 1.0001 | -0.000212 | 0.99992 | 7 | 1592
v_1M_160 | 8.4e-05 | 1 | -7.2e-05 | 0.99994 | 7 | 160
v_100k_1592 | 0.000625 | 0.99997 | -0.000636 | 1 | 6 | 1592
v_100k_160 | 0.000625 | 0.99996 | -0.000666 | 1 | 6 | 160
v_10k_1592 | 5.1e-05 | 1 | -6.2e-05 | 0.99997 | 5 | 1592
v_10k_160 | 5.1e-05 | 1 | -5.9e-05 | 0.99997 | 5 | 160
er1_1k_1592 | -0.000201 | 1.0001 | 0.000191 | 0.99989 | 4 | 1592
er1_1k_160 | -0.000203 | 1.0001 | 0.000196 | 0.9999 | 4 | 160
er1_100_1592 | -0.002461 | 1 | 0.002455 | 0.99995 | 3 | 1592
er1_100_160 | -0.002489 | 1.0001 | 0.002482 | 0.99993 | 3 | 160
er1_10_1592 | -0.000306 | 1 | 0.000312 | 0.99998 | 2 | 1592
er1_10_160 | -3.3e-05 | 1 | 0.000282 | 1 | 2 | 160
ep_10_1592 | -1.5e-05 | 1.0001 | 0.000395 | 0.99997 | 1 | 1592
ep_10_160 | 1.9e-05 | 1 | 0.00035 | 0.99995 | 1 | 160
Manual entry
Resistor ac Cal zeros | Y dial a | Y dial b | Z dial a | Z dial b | Range | Frequency | Zeros from Minyu II, 26/3/23
v_1M_1592 | 5e-06 | 1.1e-06 | -1e-06 | -9e-07 | 7 | 1592 | 00000-1 | 00000-11 | -1e-06 | -9e-07 | 7Z, 1595
v_1M_160 | 3e-06 | 1.4e-06 | -1e-06 | -9e-07 | 7 | 160 | 00000-1 | 00000-11 | -1e-06 | -9e-07 | 7Z, 161
v_100k_1592 | 5e-06 | 1.1e-06 | -1e-06 | -8e-07 | 6 | 1592 | 5 | 11 | 5e-06 | 1.1e-06 | 7Y, 1595
v_100k_160 | 3e-06 | 1.4e-06 | 0 | -1e-06 | 6 | 160 | 3 | 14 | 3e-06 | 1.4e-06 | 7Y, 161
v_10k_1592 | -1e-06 | 3e-07 | -1e-06 | -2.3e-06 | 5 | 1592 | 00000-1 | 0000-192 | -1e-06 | -8e-07 | 6Z, 1595
v_10k_160 | 2e-06 | 1e-07 | -1e-06 | -2.2e-06 | 5 | 160 | 0 | 0000-190 | 0 | -1e-06 | 6Z, 161
er1_1k_1592 | 0 | -1e-07 | -1e-06 | -8e-07 | 4 | 1592 | 0 | 11 | 0 | 1.1e-06 | 6Y, 1595
er1_1k_160 | 0 | -3e-07 | 7e-06 | -7e-07 | 4 | 160 | 0 | 12 | 0 | 1.2e-06 | 6Y, 161
er1_100_1592 | 0 | -3e-07 | -1e-06 | -5e-07 | 3 | 1592 | 00000-1 | 0000-177 | -1e-06 | -2.3e-06 | 5Z, 1594
er1_100_160 | 1e-06 | -4e-07 | 5e-06 | -1e-07 | 3 | 160 | 00000-1 | 0000-178 | -1e-06 | -2.2e-06 | 5Z, 161
er1_10_1592 | 0 | -1e-07 | -1e-06 | 2.1e-06 | 2 | 1592 | 0000-19 | 3 | -1e-06 | 3e-07 | 5Y, 1595
er1_10_160 | 1e-06 | -2e-07 | 0 | -4e-06 | 2 | 160 | 2 | 1 | 2e-06 | 1e-07 | 5Y, 161
ep_10_1592 | 0 | -1e-07 | -2.2e-05 | 9.9e-06 | 1 | 1592 | 0000-19 | 0000-192 | -1e-06 | -8e-07 | 4Z, 1594
ep_10_160 | 0 | -3e-07 | 0 | 1.99e-05 | 1 | 160 | 7 | 00000-13 | 7e-06 | -7e-07 | 4Z, 161
[7 empty] | 0 | 000000-1 | 0 | -1e-07 | 4Y, 1594
[7 empty] | 0 | 00000-17 | 0 | -3e-07 | 4Y, 161
[7 empty] | 00000-1 | 000-1995 | -1e-06 | -5e-07 | 3Z, 1594
[7 empty] | 5 | 000000-1 | 5e-06 | -1e-07 | 3Z, 161
[7 empty] | 0 | 00000-17 | 0 | -3e-07 | 3Y, 1595
[7 empty] | 1 | 00000-16 | 1e-06 | -4e-07 | 3Y, 161
[7 empty] | 00000-1 | 21 | -1e-06 | 2.1e-06 | 2Z, 1597
[7 empty] | 0 | 0000-160 | 0 | -4e-06 | 2Z, 161
[7 empty] | 0 | 00000-19 | 0 | -1e-07 | 2Y, 1596
[7 empty] | 1 | 00000-18 | 1e-06 | -2e-07 | 2Y, 161
[7 empty] | 000-178 | 00000X-1 | -2.2e-05 | 9.9e-06 | 1Z, 1592
[7 empty] | 0 | 000020-1 | 0 | 1.99e-05 | 1Z, 161
[7 empty] | 0 | 00000-19 | 0 | -1e-07 | 1Y, 1597
[7 empty] | 0 | 00000-17 | 0 | -3e-07 | 1Y, 161
Inductors | dial a | dial b | dial a zero | dial b zero | Range | Frequency
gr100uh_1k | 1.0003 | 0.086209 | 0.001427 | 0.004331 | 1Z | 1000
gr1000uh_1k | 0.99875 | 0.086265 | 2.2e-05 | 0.0009011 | 2Z | 1000
gr10000uh_1k | 0.99975 | 0.08522 | 1e-06 | 0.0003253 | 3Z | 1000
gr100000uh_1k | 0.9996 | 0.083491 | 0 | 1.07e-05 | 4Z | 1000
gr100000uh_100 | 1.002 | 0.064334 | 0 | -1.3e-06 | 5Z | 1000
Capacitors | dial a | dial b | dial a zero | dial b zero | Range | Frequency
su100000_1k | 1.0014 | 0.0055362 | 0 | 0 | 4Y | 1000 | range zeros not measured
su1000000_1k | 1.0036 | 0.011207 | 0 | 0 | 3Y | 1000
gr1000a_1592 | 0.99987 | -1.4e-05 | 0 | 0 | 6Y | 1592
gr1000b_1592 | 0.99986 | -1.3e-05 | 0 | 0 | 6Y | 1592
gr100_1592 | 0.99992 | 0.000259 | 0 | 0 | 7Y | 1592
gr_10x100_1k | 0.98236 | -0.031129 | 0 | 0 | 3Y | 1000
gr_10x10n_1k | 0.99747 | -0.0087993 | 0 | 0 | 4Y | 1000
gr_10x1n_1k | 1.0054 | 0.0002107 | 0 | 0 | 5Y | 1000
gr_0_1k | 0.004845 | 1.97e-05 | 0 | 0 | 5Y | 1000

## Corrections
This table is the usual form for calculating spreadsheet corrections
Component | Value
(ppm)
C_1 | 157.25
R_4A | 57.5886
R_4B | -2.4836
R_4C | 72.6886
G_1 | 15.2828
G_2 | 9.7891 |  | Correction term d is
assumed negligible.
Z3=(R4/G2)(aG1+ßjwC1)(1+c+jd) |  |  | Correction term c is
where Z3=R+jwC |  | correction_c | 10.97 | see 'Resistors' sheet, Comment block N74
Y4=(G3/G2)(aG1+ßjwC1)(1-c-jd) |  |  | (G3=1/R4*)
where Y4=G+jwC
R4A is used for ranges 5 to 1 inclusive
R4B is used on range 6 only
R4C is used on range 7 only
Without correction term "c" |  |  |  |  | With correction term "c"
|  | correction to L/C dial | correction to R/G dial |  |  | correction to L/C dial | correction to R/G dial
Range |  | b correction | a correction |  | Range | b correction | a correction
|  | (ppm) | (ppm) |  |  | (ppm) | (ppm)
1Z |  | 205.0495 | 63.0823 |  | 1Z | 216.0195 | 74.0523
2Z |  | 205.0495 | 63.0823 |  | 2Z | 216.0195 | 74.0523
3Z |  | 205.0495 | 63.0823 |  | 3Z | 216.0195 | 74.0523
4Z |  | 205.0495 | 63.0823 |  | 4Z | 216.0195 | 74.0523
5Z |  | 205.0495 | 63.0823 |  | 5Z | 216.0195 | 74.0523
6Z |  | 144.9773 | 3.0101 |  | 6Z | 155.9473 | 13.9801
7Z |  | 220.1495 | 78.1823 |  | 7Z | 231.1195 | 89.1523
1Y |  | 89.8723 | -52.0949 |  | 1Y | 78.9023 | -63.0649
2Y |  | 89.8723 | -52.0949 |  | 2Y | 78.9023 | -63.0649
3Y |  | 89.8723 | -52.0949 |  | 3Y | 78.9023 | -63.0649
4Y |  | 89.8723 | -52.0949 |  | 4Y | 78.9023 | -63.0649
5Y |  | 89.8723 | -52.0949 |  | 5Y | 78.9023 | -63.0649
6Y |  | 149.9445 | 7.9773 |  | 6Y | 138.9745 | -2.9927
7Y |  | 74.7723 | -67.1949 |  | 7Y | 63.8023 | -78.1649

## Uncertainties
Universal Bridge Uncertainties compared with uncertainties prior to Aug-Nov 2009 calibration
Note: This sheet is not protected.
Assumptions |  | Changes in uncertainties
1) The uncertainties apply to ranges 1 to 6 |  |  |  |  | Revert phase interaction to 16, otherwise
2) The uncertainties apply for frequencies from 160 Hz to 1592 Hz |  |  |  |  | July 2014 uncertainties
are same as for 2009
| Previous uncertainties | Feb 2004 uncertainties | Feb 2005 uncertainties | Aug 2009 uncertainties | July 2014 uncertainties
| Standard | Standard | Standard | Standard | Standard
Source of uncertainty | Uncertainties | Uncertainties | Uncertainties | Uncertainties | Uncertainties | Degrees of
| u (ppm of full dial scale) | u (ppm of full dial scale) | u (ppm of full dial scale) |  |  | Freedom, v | Previous Comment | Feb 2004 Comment | Feb 2005 Comment | Aug 2009 comment
INDUCTANCE
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here |  | problem is similar, increase is not clearly significant
component stability | 2 | 5 | 5 | 5 | 5 | 2 | some doubt on stability of capacitor in particular | the 1000 pF GR capacitor can change 5 to 10 ppm short term
dial linearity | 1 | 1 | 1 | 1 | 1 | 5.5556 | 1 ppm of full scale
inductance error due to phase error in resistance | 16 | 5 | 5 | 5 | 16 | 13 | 16 ppm of resistance dial setting, see comment for resistance | If no correction applied. See comment E53, Short, Open, Phase.
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be range and frequency dependent
temperature coefficient of capacitor | 2 | 2 | 2 | 2 | 2 | 5.5556 | 2 ppm per degree for capacitor
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
RESISTANCE of an Inductor
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 2 | 2 | 2 | 2 | 5.5556 | stability of resistors only
dial linearity | 2 | 2 | 2 | 1 | 1 | 5.6 |  |  |  | match to inductance linearity
resistance error due to phase error in inductance | 16 | 5 | 5 | 5 | 5 | 13 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | If no correction applied. See comment E53, Short, Open, Phase.
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be dependent on frequency and range, included below.
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
CAPACITANCE
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 5 | 5 | 5 | 5 | 2 | some doubt on stability of capacitor in particular | the 1000 pF GR capacitor can change 5 to 10 ppm short term
dial linearity | 1 | 1 | 1 | 1 | 1 | 5.5556 | 1 ppm of full scale
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be range and frequency dependent
temperature coefficient of capacitor | 2 | 2 | 2 | 2 | 2 | 5.5556 | 2 ppm per degree for capacitor
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components
CONDUCTANCE of a Capacitor
ac/dc error of standards | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 5.5556 | 5ppm considered two sigma limit, essentially relying on dc values of resistors
amplifier/ratio | 2.5 | 7 | 14 | 17 | 17 | 5.5556 | based on scatter of results in 1984 | unexplained errors on RHS of bridge included here
component stability | 2 | 2 | 2 | 2 | 2 | 5.5556 | stability of resistors only
conductance error due to phase error in capacitance | 16 | 16 | 16 | 16 | 16 | 13 | based on standard deviation of 1984 determinations, 160 counts (16 ppm) for C at full scale | If no correction applied. See comment E53, Short, Open, Phase. |  | Phase error same as for inductor
dial linearity | 2 | 2 | 2 | 1 | 1 | 5.6 |  |  |  | match to capacitance linearity
resolution | 1 | 1 | 1 | 1 | 1 | 2 | this might be dependent on frequency and range, included below.
temperature coefficient of resistors | 2 | 2 | 2 | 2 | 2 | 2 | depends on matching of Vishay components

## ESI_check
ESI box summary for comparison with MI bridge values
see Y:\Ongoing\OHM\ESI_Boxes\BuildupESIboxes2014.xls, ~10 Feb 2014
Box/step | MI(ppm) | HP3458
1 | -73.6095
10 | -6.4394 | 9.5564 | 15.9957
100 | -32.076 | -34.5819 | -2.5058
1000 | -31.7342 | -39.7787 | -8.0445
10000 | 17.9648
100000 | 8.7165 | 11.9986 | 3.2821
Note that humidity/temperature/time will account for the differences.
View as a 'common sense' check only.
To be comfirmed, if it's needed
| Latest one is "BuildupESIboxes2016", which modified at 01/11/22 | G:\Shared drives\MSL - Electricity - Ongoing\OHM\ESI_Boxes

## Original readings
mains frequency = | 49.9854618E+00 | Note; | 2023-01-31 | BY DMM3458A
temperature = | 33.3
Item?: | SR104, SN:109006
Result
Time: | Tue Jan 31 10:57:54 2023
A_mean: | 10000.019534
A_stdevp: | 0.004134559608655788
B_mean: | 10000.0179305
B_stdevp: | 0.0013878778762882844
mains frequency = | 49.9900916E+00
temperature = | 33.3
Item?: | SR104, SN:109006, resistor ternimals repeat
Result
Time: | Tue Jan 31 11:11:00 2023
A_mean: | 10000.019602000002
A_stdevp: | 0.002021613737990394
B_mean: | 10000.017241500002
B_stdevp: | 0.0017046292234726512
mains frequency = | 49.9957409E+00
temperature = | 33.3
Item?: | SR104, SN:109006 Temp semsor reading_1
Result
Time: | Tue Jan 31 11:21:09 2023
A_mean: | 9981.0262778
A_stdevp: | 0.004442466226086462
B_mean: | 9981.0249082
B_stdevp: | 0.004513280142210357
mains frequency = | 49.9925106E+00
temperature = | 33.3
Item?: | SR104, SN:109006 Temp semsor reading_2
Result
Time: | Tue Jan 31 11:25:10 2023
A_mean: | 9981.08561885
A_stdevp: | 0.004515444715204005
B_mean: | 9981.0833095
B_stdevp: | 0.004518600918942526
mains frequency = | 49.9921380E+00
temperature = | 33.4
Item?: | esi SR1010 sn: 937003 in parallel PC101 sn: 939007
Result
Time: | Tue Jan 31 12:06:30 2023
A_mean: | 10000.143969000002
A_stdevp: | 0.0055625570607072015
B_mean: | 10000.144852000001
B_stdevp: | 0.0052697124660064385
mains frequency = | 50.0339570E+00
temperature = | 33.3
Item?: | esi SR1010 sn: 937003 in parallel/series SPC102 sn: 940008
Result
Time: | Tue Jan 31 13:08:30 2023
A_mean: | 10000.139812500001
A_stdevp: | 0.010778800770701237
B_mean: | 10000.1401705
B_stdevp: | 0.009663476259740686
mains frequency = | 50.0876297E+00
temperature = | 33.3
Item?: | R10 ONLY in esi SR1010 sn: 937003
Result
Error due to function accidentally typo
A_mean: | 100001.12315999999
A_stdevp: | 0.008322980236262902
B_mean: | 100000.97967
B_stdevp: | 0.008322980236262902
mains frequency = | 50.0655055E+00
temperature = | 33.3
Item?: | R10 ONLY in esi SR1010 sn: 937003 repeat
Result
Time: | Tue Jan 31 13:26:56 2023
A_mean: | 100001.08276500001
A_stdevp: | 0.011271308474733231
B_mean: | 100000.94443499998
B_stdevp: | 0.0116006022996121
mains frequency = | 49.9997594E+00
temperature = | 33.3
Item?: | R10 ONLY in esi SR1010 sn: 937003 No Ground
Result
Time: | Tue Jan 31 13:30:41 2023
A_mean: | 100001.08064999999
A_stdevp: | 0.011976182502723285
B_mean: | 100000.93784
B_stdevp: | 0.008410857643272722
mains frequency = | 50.0135650E+00
temperature = | 33.4
Item?: | UB impedance 4wire cable zero with coaxial short
Result
Time: | Tue Jan 31 14:33:54 2023
A_mean: | 0.00228932024475
A_stdevp: | 1.3839761719688043e-05
B_mean: | 0.0022892898923500002
B_stdevp: | 2.2927334223598846e-05
mains frequency = | 49.9679828E+00
temperature = | 33.6
Item?: | UB G1
Result
Time: | Tue Jan 31 14:48:45 2023
A_mean: | 99998.2863975
A_stdevp: | 0.015152377569737014
B_mean: | 99998.1814875
B_stdevp: | 0.013347610026660316
mains frequency = | 49.9970898E+00
temperature = | 33.6
Item?: | UB G2
Result
Time: | Tue Jan 31 14:57:29 2023
A_mean: | 99998.885913
A_stdevp: | 0.00845639907261579
B_mean: | 99998.7823265
B_stdevp: | 0.009348035131469411
mains frequency = | 49.9823930E+00
temperature = | 33.8
Item?: | UB R4C
Result
Time: | Tue Jan 31 15:59:26 2023
A_mean: | 1000075.4752
A_stdevp: | 0.9836337893827924
B_mean: | 1000075.73225
B_stdevp: | 1.779940267336856
mains frequency = | 49.9863151E+00
temperature = | 33.8
Item?: | UB R4A
Result
Time: | Tue Jan 31 16:05:39 2023
A_mean: | 10000.548465499998
A_stdevp: | 0.008200356523020575
B_mean: | 10000.551211
B_stdevp: | 0.006637910974320766
mains frequency = | 50.0325542E+00
temperature = | 33.8
Item?: | UB R4B
Result
Time: | Tue Jan 31 16:11:03 2023
A_mean: | 99999.33132550001
A_stdevp: | 0.009978601564664652
B_mean: | 99999.235396
B_stdevp: | 0.016064037115282145
mains frequency = | 49.9933075E+00
temperature = | 33.9
Item?: | esi SR1010 sn: 937003 in series p2 used a jupmer
Result
Time: | Tue Jan 31 16:56:06 2023
A_mean: | 1000015.10485
A_stdevp: | 0.2846176151336055
B_mean: | 999466.2641550001
B_stdevp: | 2455.073229902235
mains frequency = | 49.9263398E+00
temperature = | 33.8
Item?: | esi SR1010 sn: 937003 R1-R10 in series
Result
Time: | Tue Jan 31 17:00:42 2023
A_mean: | 1000016.3737
A_stdevp: | 1.1955658911324463
B_mean: | 1000016.1350499999
B_stdevp: | 0.6583370177523764
mains frequency = | 49.9913317E+00
temperature = | 33.8
Item?: | R1-R10 another repeat 21.41C
Result
Time: | Tue Jan 31 17:03:53 2023
A_mean: | 1000015.3794
A_stdevp: | 0.6090915928084638
B_mean: | 1000015.3623499997
B_stdevp: | 0.595737338351875
mains frequency = | 50.0904
temperature = | 33.5
Item?: | esi sr1010 sn:939007 p//ssspc102 sn:94008
Result
Time: | Wed Feb 1 17:09:44 2023
A_mean: | 100001.2749
A_stdevp: | 0.055109
B_mean: | 100001.1199
B_stdevp: | 0.027799

## docx
Calibration Log
This is a first time to track the calibration of dc resistors for the universal bridge by copying in the output of Python scripts for operating the HP3458A. | readings of Step 16 in 2023 Starts from row 2111
ACAL ALL
Do this at the beginning of every day.
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 49.9868902E+00 Hz
Temperature = 32.8
ACAL ALL starting
ACAL finished |  | .
Time taken = 16.0816833337 minutes
Temperature = 33.0
Mon Jul 07 10:02:35 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
Measure SR104 10k
Using DIN lead to connect
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104 temperature sensor
Trial 9.968847642E+03
0 9.968848812E+03
1 9.968849556E+03
2 9.968848774E+03
3 9.968851888E+03
4 9.968849861E+03
5 9.968850095E+03
6 9.968851475E+03
7 9.968851193E+03
8 9.968850259E+03
9 9.968848787E+03
10 9.968852180E+03
11 9.968852252E+03
12 9.968850656E+03
13 9.968852848E+03
14 9.968853640E+03
15 9.968853569E+03
16 9.968852335E+03
17 9.968853911E+03
18 9.968853262E+03
19 9.968851806E+03
Result
Mon Jul 07 11:11:43 2014
9968.8514
0.0016853
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104
Trial 9.999973446E+03
0 9.999974730E+03
1 9.999975014E+03
2 9.999974547E+03
3 9.999974136E+03
4 9.999972150E+03
5 9.999972246E+03
6 9.999974452E+03
7 9.999971986E+03
8 9.999972802E+03
9 9.999974782E+03
10 9.999975478E+03
11 9.999974745E+03
12 9.999972656E+03
13 9.999971060E+03
14 9.999973176E+03
15 9.999973711E+03
16 9.999973923E+03
17 9.999974118E+03
18 9.999973191E+03
19 9.999972615E+03
Result
Mon Jul 07 11:22:27 2014
9999.9736
0.001206
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? SR104 repeat
Trial 9.999973826E+03
0 9.999973572E+03
1 9.999974375E+03
2 9.999972206E+03
3 9.999975232E+03
4 9.999973291E+03
5 9.999972529E+03
6 9.999973886E+03
7 9.999972752E+03
8 9.999973631E+03
9 9.999974736E+03
10 9.999971425E+03
11 9.999975183E+03
12 9.999974039E+03
13 9.999973422E+03
14 9.999972873E+03
15 9.999972462E+03
16 9.999974698E+03
17 9.999973321E+03
18 9.999974010E+03
19 9.999973134E+03
Result
Mon Jul 07 11:25:05 2014
9999.9735
0.001011
>Exit code: 0
GMH51 reads 19.81 °C in well of SR104
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.0E4
Item? UB R4A guard open
Trial 1.000051105E+04
0 1.000051677E+04
1 1.000051329E+04
2 1.000051552E+04
3 1.000051277E+04
4 1.000051393E+04
5 1.000051192E+04
6 1.000051119E+04
7 1.000051371E+04
8 1.000051637E+04
9 1.000051443E+04
10 1.000051400E+04
11 1.000051349E+04
12 1.000051429E+04
13 1.000051503E+04
14 1.000051428E+04
15 1.000051693E+04
16 1.000051126E+04
17 1.000051550E+04
18 1.000050864E+04
19 1.000051388E+04
Result
Mon Jul 07 11:43:03 2014
10000.5139
0.0020299
>Exit code: 0
GMH51 reads 19.65°C at R4A
This next run had error messag…wrong range?
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :!.0E4
Item? UB R4A, guard to low, guard to case in both instances
Trial 1.000051184E+04
0 1.000050288E+04
1 1.000054114E+04
2 1.000051742E+04
3 1.000050780E+04
4 1.000027609E+04
5 1.000047247E+04
6 1.000068019E+04
7 1.000052316E+04
8 1.000054238E+04
9 1.000060744E+04
10 1.000046489E+04
11 1.000052452E+04
12 1.000051911E+04
13 1.000051009E+04
14 1.000052685E+04
15 1.000052598E+04
16 1.000051702E+04
17 1.000051372E+04
18 1.000051516E+04
19 1.000051513E+04
Result
Mon Jul 07 11:46:34 2014
10000.5152
0.072305
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, guard to low and case, correct range?
Trial 1.000052034E+04
0 1.000047879E+04
1 1.000051572E+04
2 1.000051427E+04
3 1.000051448E+04
4 1.000051173E+04
5 1.000052494E+04
6 1.000050944E+04
7 1.000051340E+04
8 1.000051265E+04
9 1.000051301E+04
10 1.000051536E+04
11 1.000051105E+04
12 1.000051434E+04
13 1.000051467E+04
14 1.000051164E+04
15 1.000051268E+04
16 1.000050846E+04
17 1.000051377E+04
18 1.000051271E+04
19 1.000051169E+04
Result
Mon Jul 07 11:49:40 2014
10000.5117
0.0084162
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E4
Item? UB R4A, guard open and to screen
Trial 1.000044113E+04
0 1.000051299E+04
1 1.000051810E+04
2 1.000050990E+04
3 1.000051338E+04
4 1.000051482E+04
5 1.000050946E+04
6 1.000051668E+04
7 1.000050979E+04
8 1.000051601E+04
9 1.000051061E+04
10 1.000051345E+04
11 1.000051100E+04
12 1.000051316E+04
13 1.000051236E+04
14 1.000051058E+04
15 1.000051413E+04
16 1.000051436E+04
17 1.000051031E+04
18 1.000051388E+04
19 1.000051327E+04
Result
Mon Jul 07 11:52:54 2014
10000.5129
0.0024255
>Exit code: 0
GMH51 19.60 °C at R4A.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, after lunch break, now with DELAY set to 2 seconds
Trial 1.000052589E+04
0 1.000052074E+04
1 1.000051917E+04
2 1.000051844E+04
3 1.000051013E+04
4 1.000051697E+04
5 1.000051094E+04
6 1.000051846E+04
7 1.000051544E+04
8 1.000049867E+04
9 1.000051633E+04
10 1.000053928E+04
11 1.000051701E+04
12 1.000052019E+04
13 1.000051840E+04
14 1.000051891E+04
15 1.000051976E+04
16 1.000051501E+04
17 1.000052022E+04
18 1.000052817E+04
19 1.000050408E+04
Result
Mon Jul 07 13:02:18 2014
10000.5173
0.0081742
>Exit code: 0
GMH51 19.52 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2e4
Item? UB R4A, no special delay after trigger, guard still open
Trial 1.000051283E+04
0 1.000052131E+04
1 1.000051804E+04
2 1.000051886E+04
3 1.000051461E+04
4 1.000055984E+04
5 1.000046729E+04
6 1.000055989E+04
7 1.000047512E+04
8 1.000054184E+04
9 1.000049248E+04
10 1.000052779E+04
11 1.000048655E+04
12 1.000053447E+04
13 1.000048190E+04
14 1.000054823E+04
15 1.000054815E+04
16 1.000047037E+04
17 1.000052479E+04
18 1.000050109E+04
19 1.000051294E+04
Result
Mon Jul 07 13:08:38 2014
10000.5153
0.029127
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? UB R4A, no special delay, Guard still open
Trial 1.000051478E+04
0 1.000054698E+04
1 1.000051195E+04
2 1.000051439E+04
3 1.000051689E+04
4 1.000050801E+04
5 1.000052710E+04
6 1.000050310E+04
7 1.000052782E+04
8 1.000050876E+04
9 1.000051984E+04
10 1.000051878E+04
11 1.000051292E+04
12 1.000052282E+04
13 1.000051277E+04
14 1.000054627E+04
15 1.000050765E+04
16 1.000051395E+04
17 1.000052917E+04
18 1.000050868E+04
19 1.000051352E+04
Result
Mon Jul 07 13:12:05 2014
10000.5186
0.011933
>Exit code: 0
The above 2 runs after lunch seem to start with one high value that shows as a higher standard deviation.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? UB R4A, try again to see if standard deviation comes down
Trial 1.000052412E+04
0 1.000049257E+04
1 1.000051917E+04
2 1.000051294E+04
3 1.000052710E+04
4 1.000050427E+04
5 1.000051553E+04
6 1.000051450E+04
7 1.000051694E+04
8 1.000051604E+04
9 1.000051280E+04
10 1.000051749E+04
11 1.000051597E+04
12 1.000051586E+04
13 1.000051458E+04
14 1.000051526E+04
15 1.000051668E+04
16 1.000051302E+04
17 1.000052021E+04
18 1.000051263E+04
19 1.000051755E+04
Result
Mon Jul 07 13:17:35 2014
10000.5146
0.0066755
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? UB R4A, one more preliminary trial added
Trial 1.000050991E+04
Trial 2 1.000051573E+04
0 1.000051664E+04
1 1.000051542E+04
2 1.000051310E+04
3 1.000051525E+04
4 1.000051363E+04
5 1.000051658E+04
6 1.000051659E+04
7 1.000051617E+04
8 1.000051439E+04
9 1.000051410E+04
10 1.000051457E+04
11 1.000051342E+04
12 1.000051518E+04
13 1.000051140E+04
14 1.000051286E+04
15 1.000051382E+04
16 1.000051251E+04
17 1.000051281E+04
18 1.000051286E+04
19 1.000051499E+04
Result
Mon Jul 07 13:20:44 2014
10000.5143
0.0015186
>Exit code: 0
GMH51 19.55 °C
Note better behaviour above with the extra preliminary reading. Data could be reprocessed to eliminate odd initial readings…then keeping guard to low would make sense.
Now proceed with an ESI buildup to give us 100 k and 1 M values.
Start with SR104 again
Guard to lo and connected to screen is now the default
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? SR104 thermometer
Trial 9.969754446E+03
Trial 2 9.969752259E+03
0 9.969753171E+03
1 9.969755265E+03
2 9.969752744E+03
3 9.969753893E+03
4 9.969755510E+03
5 9.969754812E+03
6 9.969756225E+03
7 9.969756247E+03
8 9.969755079E+03
9 9.969754888E+03
10 9.969755070E+03
11 9.969753998E+03
12 9.969755202E+03
13 9.969754467E+03
14 9.969755483E+03
15 9.969755763E+03
16 9.969755918E+03
17 9.969758751E+03
18 9.969757845E+03
19 9.969756461E+03
Result
Mon Jul 07 13:46:06 2014
9969.7553
0.0014158
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
1Range? :1.2e4
Item? SR104 thermometer repeat
Trial 9.969757629E+03
Trial 2 9.969768426E+03
0 9.969783854E+03
1 9.969764005E+03
2 9.969757494E+03
3 9.969774507E+03
4 9.969777351E+03
5 9.969757395E+03
6 9.969778043E+03
7 9.969772016E+03
8 9.969764523E+03
9 9.969780099E+03
10 9.969762923E+03
11 9.969765847E+03
12 9.969763095E+03
13 9.969780969E+03
14 9.969768253E+03
15 9.969782385E+03
16 9.969779386E+03
17 9.969754849E+03
18 9.969784565E+03
19 9.969766940E+03
Result
Mon Jul 07 13:48:45 2014
9969.7709
0.0095478
>Exit code: 0
GMH51 19.88 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2e4
Item? SR104
Trial 9.999972622E+03
Trial 2 9.999970557E+03
0 9.999972805E+03
1 9.999973710E+03
2 9.999973138E+03
3 9.999975826E+03
4 9.999973500E+03
5 9.999971629E+03
6 9.999971824E+03
7 9.999971934E+03
8 9.999972801E+03
9 9.999971968E+03
10 9.999972916E+03
11 9.999971712E+03
12 9.999973280E+03
13 9.999974768E+03
14 9.999972485E+03
15 9.999974247E+03
16 9.999973615E+03
17 9.999973429E+03
18 9.999972278E+03
19 9.999972582E+03
Result
Mon Jul 07 13:56:58 2014
9999.973
0.0010819
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
1Range? :1.2e4
Item? SR104 repeat
Trial 9.999972281E+03
Trial 2 9.999969798E+03
0 9.999970333E+03
1 9.999973018E+03
2 9.999971955E+03
3 9.999972642E+03
4 9.999974570E+03
5 9.999970514E+03
6 9.999971585E+03
7 9.999970321E+03
8 9.999972314E+03
9 9.999972744E+03
10 9.999971399E+03
11 9.999971101E+03
12 9.999971577E+03
13 9.999971836E+03
14 9.999972996E+03
15 9.999971721E+03
16 9.999973228E+03
17 9.999971913E+03
18 9.999971466E+03
19 9.999970512E+03
Result
Mon Jul 07 13:59:39 2014
9999.9719
0.0010976
>Exit code: 0
GMH51 19.91 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2e4
Item? ESI 100 k/step in parallel
Trial 1.000004489E+04
Trial 2 1.000003969E+04
0 1.000004267E+04
1 1.000004232E+04
2 1.000004242E+04
3 1.000004483E+04
4 1.000004193E+04
5 1.000004357E+04
6 1.000004127E+04
7 1.000004281E+04
8 1.000004205E+04
9 1.000004411E+04
10 1.000004482E+04
11 1.000004401E+04
12 1.000004213E+04
13 1.000004370E+04
14 1.000004166E+04
15 1.000004122E+04
16 1.000004310E+04
17 1.000004374E+04
18 1.000004264E+04
19 1.000004363E+04
Result
Mon Jul 07 14:06:30 2014
10000.0429
0.0010864
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? ESI 100 k/step, with parallel bars and network as before
Trial 1.000004600E+04
Trial 2 1.000004386E+04
0 1.000004403E+04
1 1.000004318E+04
2 1.000004348E+04
3 1.000004129E+04
4 1.000004017E+04
5 1.000004470E+04
6 1.000004281E+04
7 1.000004242E+04
8 1.000004286E+04
9 1.000004135E+04
10 1.000004320E+04
11 1.000004179E+04
12 1.000004017E+04
13 1.000004132E+04
14 1.000004045E+04
15 1.000004219E+04
16 1.000004128E+04
17 1.000004109E+04
18 1.000004435E+04
19 1.000004015E+04
Result
Mon Jul 07 14:11:36 2014
10000.0421
0.0014255
>Exit code: 0
GMH51 20.07 °C next to ESI box
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E4
Item? ESI 100k/step in parallel, repeat (distracted by email to Laly)
Trial 1.000003918E+04
Trial 2 1.000003865E+04
0 1.000004056E+04
1 1.000004283E+04
2 1.000003867E+04
3 1.000004095E+04
4 1.000004042E+04
5 1.000004087E+04
6 1.000004012E+04
7 1.000003737E+04
8 1.000003809E+04
9 1.000003920E+04
10 1.000003750E+04
11 1.000004008E+04
12 1.000003788E+04
13 1.000003379E+04
14 1.000003932E+04
15 1.000003986E+04
16 1.000003920E+04
17 1.000003612E+04
18 1.000003831E+04
19 1.000003994E+04
Result
Mon Jul 07 14:45:30 2014
10000.0391
0.0019682
>Exit code: 0
GMH51 20.01 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step in series parallel on 1-9
Trial 1.000002318E+05
Trial 2 1.000002333E+05
0 1.000001983E+05
1 1.000002073E+05
2 1.000001877E+05
3 1.000001935E+05
4 1.000002095E+05
5 1.000002126E+05
6 1.000001944E+05
7 1.000002033E+05
8 1.000002193E+05
9 1.000001998E+05
10 1.000002189E+05
11 1.000002049E+05
12 1.000002117E+05
13 1.000002004E+05
14 1.000001984E+05
15 1.000001999E+05
16 1.000001974E+05
17 1.000002061E+05
18 1.000001741E+05
19 1.000002034E+05
Result
Mon Jul 07 14:53:59 2014
100000.202
0.010456
>Exit code: 0
GMH51 20.07 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step in s-p 1-9, repeat
Trial 1.000002125E+05
Trial 2 1.000002080E+05
0 1.000002132E+05
1 1.000002136E+05
2 1.000002040E+05
3 1.000002001E+05
4 1.000001778E+05
5 1.000002069E+05
6 1.000001998E+05
7 1.000001934E+05
8 1.000001969E+05
9 1.000002015E+05
10 1.000001930E+05
11 1.000001828E+05
12 1.000001848E+05
13 1.000001809E+05
14 1.000001838E+05
15 1.000001979E+05
16 1.000002046E+05
17 1.000002011E+05
18 1.000001961E+05
19 1.000001986E+05
Result
Mon Jul 07 14:57:05 2014
100000.1965
0.010205
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
120000
Item? ESI 100 k/step coil 10 only
Trial 9.999990254E+04
Trial 2 9.999988844E+04
0 9.999989239E+04
1 9.999988490E+04
2 9.999987999E+04
3 9.999988814E+04
4 9.999987455E+04
5 9.999987447E+04
6 9.999987710E+04
7 9.999986630E+04
8 9.999986782E+04
9 9.999988442E+04
10 9.999987215E+04
11 9.999985706E+04
12 9.999987768E+04
13 9.999987282E+04
14 9.999985879E+04
15 9.999986546E+04
16 9.999985970E+04
17 9.999986167E+04
18 9.999986196E+04
19 9.999986205E+04
Result
Mon Jul 07 15:02:38 2014
99999.872
0.010464
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? ESI 100k/step coil 10 repeat
Trial 9.999987439E+04
Trial 2 9.999987560E+04
0 9.999986246E+04
1 9.999984082E+04
2 9.999985943E+04
3 9.999984823E+04
4 9.999986350E+04
5 9.999986447E+04
6 9.999984148E+04
7 9.999984114E+04
8 9.999986801E+04
9 9.999984083E+04
10 9.999983898E+04
11 9.999984853E+04
12 9.999984752E+04
13 9.999986121E+04
14 9.999986661E+04
15 9.999985507E+04
16 9.999986677E+04
17 9.999985653E+04
18 9.999985792E+04
19 9.999986690E+04
Result
Mon Jul 07 15:04:43 2014
99999.8548
0.01038
>Exit code: 0
GMH51 20.10 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E6
Item? ESI 100k/step 1-10 in series
Trial 9.999999411E+05
Trial 2 1.000000261E+06
0 1.000000511E+06
1 1.000000485E+06
2 1.000000661E+06
3 1.000000075E+06
4 1.000000824E+06
5 1.000000551E+06
6 1.000000941E+06
7 1.000000650E+06
8 1.000000773E+06
9 1.000000617E+06
10 1.000000732E+06
11 1.000001057E+06
12 1.000001001E+06
13 1.000001072E+06
14 1.000001005E+06
15 1.000001410E+06
16 1.000001335E+06
17 1.000001110E+06
18 1.000001290E+06
19 1.000001315E+06
Result
Mon Jul 07 15:11:06 2014
1000000.8708
0.3444
>Exit code: 0
GMH51 20.11 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? ESI 100k/step 1-10 in series, repeat
Trial 1.000002540E+06
Trial 2 1.000002393E+06
0 1.000002853E+06
1 1.000002656E+06
2 1.000002477E+06
3 1.000002456E+06
4 1.000002490E+06
5 1.000002296E+06
6 1.000002549E+06
7 1.000002195E+06
8 1.000002461E+06
9 1.000002326E+06
10 1.000002483E+06
11 1.000002379E+06
12 1.000002663E+06
13 1.000002427E+06
14 1.000002362E+06
15 1.000002331E+06
16 1.000002983E+06
17 1.000002585E+06
18 1.000002991E+06
19 1.000002926E+06
Result
Mon Jul 07 15:13:35 2014
1000002.5445
0.23377
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R4B,screen connected to guard
Trial 9.999851622E+04
Trial 2 9.999852071E+04
0 9.999851887E+04
1 9.999852840E+04
2 9.999852880E+04
3 9.999851685E+04
4 9.999852670E+04
5 9.999849643E+04
6 9.999850678E+04
7 9.999854300E+04
8 9.999848452E+04
9 9.999850813E+04
10 9.999853209E+04
11 9.999851728E+04
12 9.999855155E+04
13 9.999846786E+04
14 9.999850982E+04
15 9.999852877E+04
16 9.999850358E+04
17 9.999852062E+04
18 9.999852889E+04
19 9.999851465E+04
Result
Mon Jul 07 15:50:00 2014
99998.5167
0.019288
>Exit code: 0
GMH51 19.64 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R4B, repeat
Trial 9.999851985E+04
Trial 2 9.999854826E+04
0 9.999854139E+04
1 9.999851160E+04
2 9.999851911E+04
3 9.999852385E+04
4 9.999850656E+04
5 9.999852170E+04
6 9.999849829E+04
7 9.999850536E+04
8 9.999851757E+04
9 9.999851992E+04
10 9.999851615E+04
11 9.999851062E+04
12 9.999852176E+04
13 9.999853130E+04
14 9.999853575E+04
15 9.999851119E+04
16 9.999850832E+04
17 9.999849515E+04
18 9.999851274E+04
19 9.999852954E+04
Result
Mon Jul 07 15:53:39 2014
99998.5169
0.011883
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB G1 through BPO tees
Trial 9.999821890E+04
Trial 2 9.999825721E+04
0 9.999822478E+04
1 9.999824312E+04
2 9.999822837E+04
3 9.999822644E+04
4 9.999831956E+04
5 9.999816724E+04
6 9.999827530E+04
7 9.999823707E+04
8 9.999820014E+04
9 9.999825013E+04
10 9.999820758E+04
11 9.999823278E+04
12 9.999824534E+04
13 9.999819385E+04
14 9.999823569E+04
15 9.999824552E+04
16 9.999823241E+04
17 9.999823336E+04
18 9.999821847E+04
19 9.999820986E+04
Result
Mon Jul 07 15:58:11 2014
99998.2314
0.031135
>Exit code: 0
GMH51 19.65 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E5
Item? UB G1, repeat
Trial 9.999820850E+04
Trial 2 9.999824112E+04
0 9.999823716E+04
1 9.999820791E+04
2 9.999821928E+04
3 9.999820587E+04
4 9.999822005E+04
5 9.999821540E+04
6 9.999820422E+04
7 9.999820865E+04
8 9.999819905E+04
9 9.999820661E+04
10 9.999819155E+04
11 9.999820612E+04
12 9.999819771E+04
13 9.999819191E+04
14 9.999820451E+04
15 9.999821002E+04
16 9.999821671E+04
17 9.999819260E+04
18 9.999819907E+04
19 9.999821590E+04
Result
Mon Jul 07 16:00:55 2014
99998.2075
0.011203
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E5
Item? UB R2, through tees
Trial 9.999872717E+04
Trial 2 9.999875119E+04
0 9.999875371E+04
1 9.999876633E+04
2 9.999874843E+04
3 9.999873617E+04
4 9.999874592E+04
5 9.999874689E+04
6 9.999875441E+04
7 9.999875031E+04
8 9.999875558E+04
9 9.999874941E+04
10 9.999875476E+04
11 9.999875096E+04
12 9.999876137E+04
13 9.999875080E+04
14 9.999875900E+04
15 9.999874084E+04
16 9.999874885E+04
17 9.999876419E+04
18 9.999876074E+04
19 9.999875652E+04
Result
Mon Jul 07 16:04:27 2014
99998.7528
0.0075329
>Exit code: 0
GMH51 19.65 °C
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E5
Item? UB R2, repeat
Trial 9.999873303E+04
Trial 2 9.999876652E+04
0 9.999875159E+04
1 9.999874877E+04
2 9.999874038E+04
3 9.999875179E+04
4 9.999875089E+04
5 9.999876352E+04
6 9.999873818E+04
7 9.999874177E+04
8 9.999875559E+04
9 9.999873250E+04
10 9.999874771E+04
11 9.999875003E+04
12 9.999873891E+04
13 9.999876193E+04
14 9.999873841E+04
15 9.999873134E+04
16 9.999873348E+04
17 9.999875415E+04
18 9.999875171E+04
19 9.999873934E+04
Result
Mon Jul 07 16:07:34 2014
99998.7461
0.0094139
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E6
Item? UB R4C
Trial 1.000064944E+06
Trial 2 1.000063309E+06
0 1.000063088E+06
1 1.000063522E+06
2 1.000062799E+06
3 1.000063077E+06
4 1.000062732E+06
5 1.000063157E+06
6 1.000063340E+06
7 1.000063490E+06
8 1.000063180E+06
9 1.000063129E+06
10 1.000062883E+06
11 1.000062625E+06
12 1.000062786E+06
13 1.000063017E+06
14 1.000063443E+06
15 1.000063546E+06
16 1.000063724E+06
17 1.000063227E+06
18 1.000063196E+06
19 1.000062793E+06
Result
Mon Jul 07 16:12:22 2014
1000063.1377
0.30782
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E6
Item? UB R4C, repeat
Trial 1.000065819E+06
Trial 2 1.000067191E+06
0 1.000066664E+06
1 1.000064486E+06
2 1.000063662E+06
3 1.000063648E+06
4 1.000065137E+06
5 1.000067281E+06
6 1.000068879E+06
7 1.000069332E+06
8 1.000068557E+06
9 1.000066801E+06
10 1.000065168E+06
11 1.000064379E+06
12 1.000064598E+06
13 1.000065889E+06
14 1.000067184E+06
15 1.000068094E+06
16 1.000068104E+06
17 1.000067190E+06
18 1.000067680E+06
19 1.000067224E+06
Result
Mon Jul 07 16:40:29 2014
1000066.4979
1.7619
>Exit code: 0
GMH51 19.58 °C
Next calibrate ESI set of resistors
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 50.0514113E+00 Hz
Temperature = 32.9
ACAL ALL starting
ACAL finished
Time taken = 16.2855000019 minutes
Temperature = 32.9
Thu Jul 17 13:34:55 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 32.9
Range? :1.2E4
Item? SR104 thermistor
Trial 9.971495330E+03
Trial 2 9.971496304E+03
0 9.971495573E+03
1 9.971496996E+03
2 9.971496629E+03
3 9.971495857E+03
4 9.971494343E+03
5 9.971495549E+03
6 9.971495693E+03
7 9.971495905E+03
8 9.971494767E+03
9 9.971497310E+03
10 9.971497057E+03
11 9.971501974E+03
12 9.971499281E+03
13 9.971497535E+03
14 9.971498418E+03
15 9.971495111E+03
16 9.971497004E+03
17 9.971495674E+03
18 9.971498310E+03
19 9.971495692E+03
Result
Thu Jul 17 13:45:11 2014
9971.4967
0.0017802
>Exit code: 0
GMH51 reads 20.06 degrees
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? SR104
Trial 9.999969676E+03
Trial 2 9.999969541E+03
0 9.999967457E+03
1 9.999968973E+03
2 9.999967447E+03
3 9.999970984E+03
4 9.999967750E+03
5 9.999970241E+03
6 9.999968971E+03
7 9.999970669E+03
8 9.999970184E+03
9 9.999969089E+03
10 9.999968120E+03
11 9.999970090E+03
12 9.999970392E+03
13 9.999967860E+03
14 9.999967198E+03
15 9.999968372E+03
16 9.999968818E+03
17 9.999966314E+03
18 9.999967115E+03
19 9.999968399E+03
Result
Thu Jul 17 13:52:00 2014
9999.9687
0.0013473
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.0
Range? :1.2E4
Item? SR104 repeat
Trial 9.999969196E+03
Trial 2 9.999968014E+03
0 9.999968116E+03
1 9.999969293E+03
2 9.999968547E+03
3 9.999965997E+03
4 9.999966624E+03
5 9.999967819E+03
6 9.999966747E+03
7 9.999966218E+03
8 9.999967544E+03
9 9.999967935E+03
10 9.999970729E+03
11 9.999967323E+03
12 9.999966437E+03
13 9.999967657E+03
14 9.999966065E+03
15 9.999967762E+03
16 9.999967547E+03
17 9.999967953E+03
18 9.999967937E+03
19 9.999968175E+03
Result
Thu Jul 17 13:55:30 2014
9999.9676
0.001136
>Exit code: 0
GMH51 in SR104 well 20.07 degrees
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E4
Item? ES! 1k 1-10 series
Trial 9.999645202E+03
Trial 2 9.999645817E+03
0 9.999644582E+03
1 9.999645575E+03
2 9.999644686E+03
3 9.999645225E+03
4 9.999645158E+03
5 9.999644638E+03
6 9.999644933E+03
7 9.999645746E+03
8 9.999646619E+03
9 9.999644517E+03
10 9.999646996E+03
11 9.999644402E+03
12 9.999644341E+03
13 9.999643298E+03
14 9.999644679E+03
15 9.999644022E+03
16 9.999645659E+03
17 9.999645362E+03
18 9.999644955E+03
19 9.999644741E+03
Result
Thu Jul 17 14:05:00 2014
9999.645
0.00084579
>Exit code: 0
GMH51 20.05 degrees by ESI box
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E3
Item? ESI 1k 1-9 S/P
Trial 9.999632104E+02
Trial 2 9.999630321E+02
0 9.999631495E+02
1 9.999631178E+02
2 9.999631765E+02
3 9.999628805E+02
4 9.999632166E+02
5 9.999631175E+02
6 9.999629689E+02
7 9.999629700E+02
8 9.999629707E+02
9 9.999628537E+02
10 9.999630094E+02
11 9.999630536E+02
12 9.999628324E+02
13 9.999630765E+02
14 9.999632411E+02
15 9.999629403E+02
16 9.999630600E+02
17 9.999627381E+02
18 9.999626925E+02
19 9.999629922E+02
Result
Thu Jul 17 14:14:26 2014
999.963
0.00015074
>Exit code: 0
GMH51 20.09
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E3
Item? ESI 1k R10
Trial 9.999737655E+02
Trial 2 9.999737940E+02
0 9.999736473E+02
1 9.999737959E+02
2 9.999738792E+02
3 9.999737596E+02
4 9.999736667E+02
5 9.999735997E+02
6 9.999735344E+02
7 9.999738024E+02
8 9.999737068E+02
9 9.999735789E+02
10 9.999734382E+02
11 9.999736382E+02
12 9.999733559E+02
13 9.999733443E+02
14 9.999736660E+02
15 9.999735552E+02
16 9.999735056E+02
17 9.999732994E+02
18 9.999733617E+02
19 9.999733233E+02
Result
Thu Jul 17 14:19:33 2014
999.9736
0.0001751
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E3
Item? ESI 1k R1
Trial 9.999639804E+02
Trial 2 9.999638130E+02
0 9.999636698E+02
1 9.999636799E+02
2 9.999638243E+02
3 9.999636862E+02
4 9.999638516E+02
5 9.999640541E+02
6 9.999636862E+02
7 9.999637291E+02
8 9.999639401E+02
9 9.999638077E+02
10 9.999640010E+02
11 9.999637841E+02
12 9.999638811E+02
13 9.999638383E+02
14 9.999637598E+02
15 9.999638422E+02
16 9.999636058E+02
17 9.999637151E+02
18 9.999639641E+02
19 9.999638198E+02
Result
Thu Jul 17 14:24:20 2014
999.9638
0.0001199
>Exit code: 0
GMH51 20.17 degrees.
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 1k 1-10 p
Trial 9.999619231E+01
Trial 2 9.999641379E+01
0 9.999607254E+01
1 9.999640437E+01
2 9.999643328E+01
3 9.999634847E+01
4 9.999621716E+01
5 9.999633592E+01
6 9.999662691E+01
7 9.999640136E+01
8 9.999658124E+01
9 9.999635024E+01
10 9.999617473E+01
11 9.999634126E+01
12 9.999633012E+01
13 9.999650638E+01
14 9.999662631E+01
15 9.999648935E+01
16 9.999625468E+01
17 9.999632772E+01
18 9.999617991E+01
19 9.999641203E+01
Result
Thu Jul 17 14:33:16 2014
99.9964
0.00014916
>Exit code: 0
GMH51 20.20
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 1k 1-10 p, repeat
Trial 9.999633590E+01
Trial 2 9.999622223E+01
0 9.999640251E+01
1 9.999638511E+01
2 9.999638605E+01
3 9.999638538E+01
4 9.999637363E+01
5 9.999637178E+01
6 9.999624219E+01
7 9.999628705E+01
8 9.999636234E+01
9 9.999613280E+01
10 9.999620267E+01
11 9.999646790E+01
12 9.999656748E+01
13 9.999633301E+01
14 9.999622343E+01
15 9.999637672E+01
16 9.999640555E+01
17 9.999649081E+01
18 9.999630989E+01
19 9.999636415E+01
Result
Thu Jul 17 14:36:11 2014
99.9964
0.00010128
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E3
Item? ESI 100 1-10 S
Trial 9.999593322E+02
Trial 2 9.999593652E+02
0 9.999594681E+02
1 9.999594293E+02
2 9.999593828E+02
3 9.999594268E+02
4 9.999592960E+02
5 9.999592070E+02
6 9.999593703E+02
7 9.999594132E+02
8 9.999593939E+02
9 9.999593433E+02
10 9.999595731E+02
11 9.999594137E+02
12 9.999593631E+02
13 9.999594964E+02
14 9.999594233E+02
15 9.999592255E+02
16 9.999593256E+02
17 9.999593129E+02
18 9.999592013E+02
19 9.999592254E+02
Result
Thu Jul 17 14:46:46 2014
999.9594
9.9641e-05
>Exit code: 0
GMH51 19.98
pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.1
Range? :1.2E3
Item?E SI 100 1-10 S repeat after tea
Trial 9.999591445E+02
Trial 2 9.999589284E+02
0 9.999588058E+02
1 9.999588937E+02
2 9.999588314E+02
3 9.999589860E+02
4 9.999590854E+02
5 9.999590110E+02
6 9.999590416E+02
7 9.999589177E+02
8 9.999589846E+02
9 9.999587510E+02
10 9.999586031E+02
11 9.999587845E+02
12 9.999590555E+02
13 9.999586938E+02
14 9.999590344E+02
15 9.999588307E+02
16 9.999588225E+02
17 9.999586650E+02
18 9.999587548E+02
19 9.999587591E+02
Result
Thu Jul 17 15:49:57 2014
999.9589
0.00014271
>Exit code: 0
GMH51 19.72
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E2
Item? ESI 100 R1
Trial 9.999711930E+01
Trial 2 9.999690324E+01
0 9.999722929E+01
1 9.999712979E+01
2 9.999714256E+01
3 9.999717471E+01
4 9.999689610E+01
5 9.999707279E+01
6 9.999708441E+01
7 9.999701128E+01
8 9.999725316E+01
9 9.999709699E+01
10 9.999697153E+01
11 9.999708017E+01
12 9.999706381E+01
13 9.999702265E+01
14 9.999718259E+01
15 9.999709527E+01
16 9.999691501E+01
17 9.999700524E+01
18 9.999704988E+01
19 9.999709301E+01
Result
Thu Jul 17 15:54:01 2014
99.9971
9.3599e-05
>Exit code: 0
GMH51 19.80
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E1
Item? ESI 100 1-10 P
Trial 9.999553033E+00
Trial 2 9.999557216E+00
0 9.999549003E+00
1 9.999584869E+00
2 9.999556193E+00
3 9.999557771E+00
4 9.999562567E+00
5 9.999546486E+00
6 9.999551302E+00
7 9.999570111E+00
8 9.999564404E+00
9 9.999560671E+00
10 9.999576982E+00
11 9.999560644E+00
12 9.999561386E+00
13 9.999566108E+00
14 9.999562171E+00
15 9.999557495E+00
16 9.999558348E+00
17 9.999548467E+00
18 9.999552776E+00
19 9.999523645E+00
Result
Thu Jul 17 16:00:49 2014
9.9996
1.2468e-05
>Exit code: 0
GMH51 19.89
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E1
Item? ESI 10 R1
Trial 9.999964368E+00
Trial 2 9.999940452E+00
0 9.999942579E+00
1 9.999961826E+00
2 9.999943242E+00
3 9.999926210E+00
4 9.999950548E+00
5 9.999969567E+00
6 9.999945126E+00
7 9.999932827E+00
8 9.999937230E+00
9 9.999955760E+00
10 9.999961918E+00
11 9.999941277E+00
12 9.999954219E+00
13 9.999935945E+00
14 9.999949951E+00
15 9.999942958E+00
16 9.999942906E+00
17 9.999948244E+00
18 9.999950680E+00
19 9.999941952E+00
Result
Thu Jul 17 16:08:50 2014
9.9999
1.05e-05
>Exit code: 0
GMH51 20.06
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E2
Item? ESI 10 R1-10 series
Trial 9.999895003E+01
Trial 2 9.999891885E+01
0 9.999907150E+01
1 9.999911487E+01
2 9.999908492E+01
3 9.999914930E+01
4 9.999904286E+01
5 9.999886680E+01
6 9.999908335E+01
7 9.999895627E+01
8 9.999908324E+01
9 9.999898378E+01
10 9.999896120E+01
11 9.999899297E+01
12 9.999900346E+01
13 9.999883605E+01
14 9.999909381E+01
15 9.999921214E+01
16 9.999924453E+01
17 9.999905933E+01
18 9.999902850E+01
19 9.999906527E+01
Result
Thu Jul 17 16:14:14 2014
99.999
0.00010012
>Exit code: 0
GMH51 20.14
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E5
Item? 100k ishay
Trial 1.000017099E+05
Trial 2 1.000017357E+05
0 1.000017177E+05
1 1.000016704E+05
2 1.000017294E+05
3 1.000017696E+05
4 1.000017417E+05
5 1.000017171E+05
6 1.000017417E+05
7 1.000017366E+05
8 1.000017292E+05
9 1.000017461E+05
10 1.000017244E+05
11 1.000017205E+05
12 1.000017326E+05
13 1.000017152E+05
14 1.000017374E+05
15 1.000017273E+05
16 1.000017183E+05
17 1.000017323E+05
18 1.000017069E+05
19 1.000017188E+05
Result
Thu Jul 17 16:20:40 2014
100001.7267
0.019199
>Exit code: 0
GMH51 20.19
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E5
Item? Vishay 100k repeat
Trial 1.000015411E+05
Trial 2 1.000015510E+05
0 1.000015365E+05
1 1.000015268E+05
2 1.000015517E+05
3 1.000015355E+05
4 1.000015254E+05
5 1.000015401E+05
6 1.000015240E+05
7 1.000015304E+05
8 1.000015454E+05
9 1.000015580E+05
10 1.000015524E+05
11 1.000015321E+05
12 1.000015344E+05
13 1.000015440E+05
14 1.000015456E+05
15 1.000015365E+05
16 1.000015337E+05
17 1.000015335E+05
18 1.000015451E+05
19 1.000015348E+05
Result
Thu Jul 17 16:28:04 2014
100001.5383
0.0092966
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E4
Item? 10k Vishay
Trial 1.000027757E+04
Trial 2 1.000027674E+04
0 1.000027601E+04
1 1.000027719E+04
2 1.000027778E+04
3 1.000027750E+04
4 1.000027408E+04
5 1.000027672E+04
6 1.000027695E+04
7 1.000027494E+04
8 1.000027555E+04
9 1.000027679E+04
10 1.000027615E+04
11 1.000028002E+04
12 1.000027514E+04
13 1.000027444E+04
14 1.000027457E+04
15 1.000027570E+04
16 1.000027637E+04
17 1.000027354E+04
18 1.000027488E+04
19 1.000027540E+04
Result
Thu Jul 17 16:31:49 2014
10000.276
0.0015076
>Exit code: 0
GMH51 20.17
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'COM7', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E2
Item? 100 Vishay
Trial 1.000131660E+02
Trial 2 1.000130719E+02
0 1.000131073E+02
1 1.000134502E+02
2 1.000132129E+02
3 1.000131490E+02
4 1.000130690E+02
5 1.000132104E+02
6 1.000130966E+02
7 1.000131675E+02
8 1.000132009E+02
9 1.000131551E+02
10 1.000133579E+02
11 1.000132767E+02
12 1.000130706E+02
13 1.000131544E+02
14 1.000131324E+02
15 1.000130392E+02
16 1.000131840E+02
17 1.000133201E+02
18 1.000129175E+02
19 1.000130950E+02
Result
Thu Jul 17 16:35:29 2014
100.0132
0.00011932
>Exit code: 0
GMH 51 20.11
>pythonw -u "ACALall3458.py"
HP3458A
mains frequency = 50.0400462E+00 Hz
Temperature = 33.4
ACAL ALL starting
ACAL finished
Time taken = 16.0772499998 minutes
Temperature = 33.4
Thu Jul 31 13:57:24 2014
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? Vishay S106D 1M
Trial 9.999963193E+05
Trial 2 9.999967197E+05
0 9.999965833E+05
1 9.999967235E+05
2 9.999966186E+05
3 9.999964135E+05
4 9.999964585E+05
5 9.999964929E+05
6 9.999966365E+05
7 9.999964942E+05
8 9.999965459E+05
9 9.999962713E+05
10 9.999964026E+05
11 9.999962692E+05
12 9.999963153E+05
13 9.999962428E+05
14 9.999961601E+05
15 9.999962719E+05
16 9.999963667E+05
17 9.999964213E+05
18 9.999962916E+05
19 9.999963838E+05
Result
Thu Jul 31 14:32:49 2014
999996.4182
0.15059
>Exit code: 0
GMH51 20.01
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.2
Range? :1.2E6
Item? Vishay S106D 1M
Trial 9.999965304E+05
Trial 2 9.999966892E+05
0 9.999970253E+05
1 9.999971704E+05
2 9.999969313E+05
3 9.999971360E+05
4 9.999973011E+05
5 9.999971888E+05
6 9.999971570E+05
7 9.999971115E+05
8 9.999971711E+05
9 9.999968658E+05
10 9.999970695E+05
11 9.999969482E+05
12 9.999971898E+05
13 9.999969724E+05
14 9.999971814E+05
15 9.999968986E+05
16 9.999970115E+05
17 9.999970907E+05
18 9.999971375E+05
19 9.999971830E+05
Result
Thu Jul 31 14:37:33 2014
999997.087
0.11704
>Exit code: 0
GMH51 19.95
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E4
Item? SR104 temperature sensor
Trial 9.969186052E+03
Trial 2 9.969185708E+03
0 9.969184643E+03
1 9.969185981E+03
2 9.969186238E+03
3 9.969186144E+03
4 9.969188423E+03
5 9.969189547E+03
6 9.969187924E+03
7 9.969188497E+03
8 9.969186384E+03
9 9.969187898E+03
10 9.969188013E+03
11 9.969187596E+03
12 9.969185831E+03
13 9.969189290E+03
14 9.969186741E+03
15 9.969188574E+03
16 9.969190410E+03
17 9.969191168E+03
18 9.969189748E+03
19 9.969190179E+03
Result
Thu Jul 31 15:40:36 2014
9969.188
0.0017682
>Exit code: 0
GMH51 19.81
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E4
Item? SR104 10k
Trial 9.999970312E+03
Trial 2 9.999971031E+03
0 9.999970112E+03
1 9.999972805E+03
2 9.999968950E+03
3 9.999968510E+03
4 9.999968676E+03
5 9.999969939E+03
6 9.999970711E+03
7 9.999971186E+03
8 9.999970540E+03
9 9.999970725E+03
10 9.999968956E+03
11 9.999969770E+03
12 9.999970224E+03
13 9.999968849E+03
14 9.999968916E+03
15 9.999969152E+03
16 9.999971083E+03
17 9.999967072E+03
18 9.999969479E+03
19 9.999969715E+03
Result
Thu Jul 31 15:49:22 2014
9999.9698
0.0012382
>Exit code: 0
GMH51 19.82
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E4
Item? ESI 100k 1-10 p
Trial 1.000005854E+04
Trial 2 1.000005519E+04
0 1.000005444E+04
1 1.000005470E+04
2 1.000005663E+04
3 1.000005504E+04
4 1.000005708E+04
5 1.000005620E+04
6 1.000005344E+04
7 1.000005340E+04
8 1.000005378E+04
9 1.000005322E+04
10 1.000005472E+04
11 1.000005645E+04
12 1.000005765E+04
13 1.000005537E+04
14 1.000005472E+04
15 1.000005644E+04
16 1.000005663E+04
17 1.000005621E+04
18 1.000005901E+04
19 1.000005702E+04
Result
Thu Jul 31 15:55:34 2014
10000.0556
0.0015706
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.3
Range? :1.2E4
Item? ESI 100k/step, 1-10 p
Trial 1.000005337E+04
Trial 2 1.000005397E+04
0 1.000005361E+04
1 1.000005177E+04
2 1.000005296E+04
3 1.000005285E+04
4 1.000005567E+04
5 1.000005313E+04
6 1.000005635E+04
7 1.000005177E+04
8 1.000005225E+04
9 1.000005360E+04
10 1.000005273E+04
11 1.000005338E+04
12 1.000005544E+04
13 1.000005413E+04
14 1.000005072E+04
15 1.000005368E+04
16 1.000005260E+04
17 1.000005450E+04
18 1.000005296E+04
19 1.000005367E+04
Result
Thu Jul 31 15:58:09 2014
10000.0534
0.00137
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E6
Item? ESI 100k, 1-10 s
Trial 1.000003360E+06
Trial 2 1.000004248E+06
0 1.000003722E+06
1 1.000003822E+06
2 1.000003732E+06
3 1.000003887E+06
4 1.000004086E+06
5 1.000004042E+06
6 1.000004148E+06
7 1.000004002E+06
8 1.000004200E+06
9 1.000004054E+06
10 1.000004334E+06
11 1.000003929E+06
12 1.000004077E+06
13 1.000004348E+06
14 1.000004306E+06
15 1.000004498E+06
16 1.000004333E+06
17 1.000004624E+06
18 1.000004432E+06
19 1.000004480E+06
Result
Thu Jul 31 16:07:25 2014
1000004.1528
0.26189
>Exit code: 0
>pythonw -u "MeasureR.py"
['COM3', 'COM4', 'COM5', 'COM6', 'GPIB0::22']
HP3458A
Temperature 33.4
Range? :1.2E6
Item? ESI 100k/step, 1-10 series
Trial 1.000004733E+06
Trial 2 1.000004482E+06
0 1.000004688E+06
1 1.000005104E+06
2 1.000004909E+06
3 1.000004584E+06
4 1.000004707E+06
5 1.000004897E+06
6 1.000005083E+06
7 1.000004882E+06
8 1.000004765E+06
9 1.000004879E+06
10 1.000004920E+06
11 1.000005022E+06
12 1.000004656E+06
13 1.000004950E+06
14 1.000004856E+06
15 1.000004739E+06
16 1.000004974E+06
17 1.000004938E+06
18 1.000004839E+06
19 1.000004959E+06
Result
Thu Jul 31 16:09:22 2014
1000004.8675
0.14039
>Exit code: 0
GMH51 20.31
Resistor SR104 StdsR SN:109006 measured by HP3458A with ID number 5
mains frequency = | 50.0278
temperature = | 33.3
Range?: | 12000 | Item?: | SR104 StdsR SN:109006
Trail1 | 10000.006 | Trail2 | 10000.0079
at | Mon Feb 20 16:03:37 2023 | SR104 StdsR SN:109006 | reading is:
Resistance_mean: | 10000.0148
Resistance_stdevp: | 0.0018672
Time taken = | 3.3326 | minutes
----------------
Resistor SR104 Temp_sensor SN109006 measured by HP3458A with ID number 5
mains frequency = | 50.0314
temperature = | 33.3
Range?: | 12000 | Item?: | SR104 Temp_sensor SN109006
Trail1 | 9984.959 | Trail2 | 9984.9613
at | Mon Feb 20 16:15:02 2023 | SR104 Temp_sensor SN109006 | reading is:
Resistance_mean: | 9984.9603
Resistance_stdevp: | 0.002167
Time taken = | 3.3168 | minutes
----------------
Resistor ESI 1k S NetworkBars measured by HP3458A with ID number 5
mains frequency = | 49.9639
temperature = | 33.3
Range?: | 10000 | Item?: | ESI 1k S NetworkBars
Trail1 | 9999.6166 | Trail2 | 9999.6179
at | Tue Feb 21 11:25:01 2023 | ESI 1k S NetworkBars | reading is:
Resistance_mean: | 9999.6176
Resistance_stdevp: | 0.002539
Time taken = | 3.3164 | minutes
----------------
Resistor ESI 1k S 2Jumpers measured by HP3458A with ID number 5
mains frequency = | 49.9862
temperature = | 33.4
Range?: | 10000 | Item?: | ESI 1k S 2Jumpers
Trail1 | 9999.6168 | Trail2 | 9999.6182
at | Tue Feb 21 13:09:36 2023 | ESI 1k S 2Jumpers | reading is:
Resistance_mean: | 9999.6182
Resistance_stdevp: | 0.0022938
Time taken = | 3.3171 | minutes
----------------
Resistor ESI 1k S//P NetworkBars measured by HP3458A with ID number 5
mains frequency = | 49.9846
temperature = | 33.5
Range?: | 1.0e+3.0 | Item?: | ESI 1k S//P NetworkBars
Trail1 | 999.96 | Trail2 | 999.9601
at | Tue Feb 21 14:56:41 2023 | ESI 1k S//P NetworkBars | reading is:
Resistance_mean: | 999.9599
Resistance_stdevp: | 0.00034004
Time taken = | 3.3154 | minutes
----------------
Resistor ESI R10 only measured by HP3458A with ID number 5
mains frequency = | 49.9893
temperature = | 33.5
Range?: | 1000 | Item?: | ESI R10 only
Trail1 | 999.9724 | Trail2 | 999.9724
at | Tue Feb 21 15:06:34 2023 | ESI 1k R10 only | reading is:
Resistance_mean: | 999.972
Resistance_stdevp: | 0.00026674
Time taken = | 3.3183 | minutes
----------------
Resistor ESI 1k R1 only measured by HP3458A with ID number 5
mains frequency = | 49.9303
temperature = | 33.4
Range?: | 1000 | Item?: | ESI 1k R1 only
Trail1 | 999.9607 | Trail2 | 999.9605
at | Tue Feb 21 15:12:29 2023 | ESI 1k R1 only | reading is:
Resistance_mean: | 999.9604
Resistance_stdevp: | 0.00026334
Time taken = | 3.3168 | minutes
----------------
Resistor ESI 1k P// only measured by HP3458A with ID number 5
mains frequency = | 49.9985
temperature = | 33.5
Range?: | 100 | Item?: | ESI 1k P// only
Trail1 | 99.9961 | Trail2 | 99.9962
at | Tue Feb 21 15:23:17 2023 | ESI 1k P// only | reading is:
Resistance_mean: | 99.996
Resistance_stdevp: | 0.00019027
Time taken = | 3.3171 | minutes
----------------
Resistor ESI 100 Ohm S network bars measured by HP3458A with ID number 5
mains frequency = | 49.9532
temperature = | 33.5
Range?: | 100 | Item?: | ESI 100 Ohm S network bars SN:123009
Trail1 | 999.9663 | Trail2 | 999.9663
at | Tue Feb 21 16:17:16 2023 | ESI 100 Ohm S network bars | reading is:
Resistance_mean: | 999.9663
Resistance_stdevp: | 0.00022752
Time taken = | 3.3172 | minutes
----------------
Resistor SI 100 Ohm S 2XJumper Sn:123009 measured by HP3458A with ID number 5
mains frequency = | 49.9626
temperature = | 33.4
Range?: | 100 | Item?: | SI 100 Ohm S 2XJumper Sn:123009
Trail1 | 999.9662 | Trail2 | 999.9665
at | Tue Feb 21 16:23:19 2023 | SI 100 Ohm S 2XJumper Sn:123009 | reading is:
Resistance_mean: | 999.9661
Resistance_stdevp: | 0.00063946
Time taken = | 3.3164 | minutes
----------------
Resistor ESI 100 Ohm R1 only measured by HP3458A with ID number 5
mains frequency = | 49.9091
temperature = | 33.4
Range?: | 100 | Item?: | ESI 100 Ohm R1 only
Trail1 | 99.9976 | Trail2 | 99.9974
at | Tue Feb 21 16:30:02 2023 | ESI 100 Ohm R1 only | reading is:
Resistance_mean: | 99.9977
Resistance_stdevp: | 0.00015645
Time taken = | 3.3161 | minutes
----------------
Resistor eSI 100 Ohm P// network bars measured by HP3458A with ID number 5
mains frequency = | 50.0043
temperature = | 33.4
Range?: | 10 | Item?: | eSI 100 Ohm P// network bars
Trail1 | 9.9997 | Trail2 | 9.9996
at | Tue Feb 21 16:39:42 2023 | eSI 100 Ohm P// network bars | reading is:
Resistance_mean: | 9.9996
Resistance_stdevp: | 1.8983e-05
Time taken = | 3.3164 | minutes
----------------
Resistor ESI 10 Ohm R1 only SN:A1-1752030 measured by HP3458A with ID number 5
mains frequency = | 49.9526
temperature = | 33.4
Range?: | 10 | Item?: | ESI 10 Ohm R1 only SN:A1-1752030
Trail1 | 10 | Trail2 | 10.0001
at | Tue Feb 21 16:46:30 2023 | ESI 10 Ohm R1 only SN:A1-1752030 | reading is:
Resistance_mean: | 10.0001
Resistance_stdevp: | 1.6918e-05
Time taken = | 3.3309 | minutes
----------------
Resistor ESI 10 Ohm S network bars measured by HP3458A with ID number 5
mains frequency = | 49.9945
temperature = | 33.4
Range?: | 100 | Item?: | ESI 10 Ohm S network bars
Trail1 | 100.0012 | Trail2 | 100.0009
at | Tue Feb 21 16:55:53 2023 | ESI 10 Ohm S network bars | reading is:
Resistance_mean: | 100.0009
Resistance_stdevp: | 0.00015356
Time taken = | 3.3164 | minutes
----------------
Resistor ESI 10 Ohm S 2 X jumpers measured by HP3458A with ID number 5
mains frequency = | 50.0265
temperature = | 33.4
Range?: | 100 | Item?: | ESI 10 Ohm S 2 X jumpers
Trail1 | 100.0009 | Trail2 | 100.001
at | Tue Feb 21 17:01:07 2023 | ESI 10 Ohm S 2 X jumpers | reading is:
Resistance_mean: | 100.0009
Resistance_stdevp: | 0.00022843
Time taken = | 3.3175 | minutes
----------------
Resistor 100k Vishay measured by HP3458A with ID number 5
mains frequency = | 49.9767
temperature = | 33.4
Range?: | 100000 | Item?: | 100k Vishay
Trail1 | 100003.3326 | Trail2 | 100003.4064
at | Tue Feb 21 17:08:35 2023 | 100k Vishay | reading is:
Resistance_mean: | 100003.3199
Resistance_stdevp: | 0.042298
Time taken = | 3.3851 | minutes
----------------
Resistor 10k Vishay measured by HP3458A with ID number 5
mains frequency = | 49.9718
temperature = | 33.4
Range?: | 12000 | Item?: | 10k Vishay
Trail1 | 10000.3108 | Trail2 | 10000.3371
at | Tue Feb 21 17:12:44 2023 | 10k Vishay | reading is:
Resistance_mean: | 10000.3233
Resistance_stdevp: | 0.0081507
Time taken = | 3.3846 | minutes
----------------
Resistor 100 Vishay measured by HP3458A with ID number 5
mains frequency = | 49.9699
temperature = | 33.4
Range?: | 100 | Item?: | 100 Vishay
Trail1 | 100.0148 | Trail2 | 100.0146
at | Tue Feb 21 17:20:47 2023 | 100 Vishay | reading is:
Resistance_mean: | 100.0148
Resistance_stdevp: | 0.00019899
Time taken = | 3.3178 | minutes
----------------
end of the measurement
Auto Calibration
HP3458A
mains frequency = | 49.9734
temperature = | 33.5
ACAL starting
ACAL finished
Time taken = | 16.0508 | minutes
temperature = | 33.5
Cal done
Tue Feb 21 17:37:30 2023
Auto Calibration
HP3458A
mains frequency = | 49.998
temperature = | 32.9
ACAL starting
ACAL finished
Time taken = | 15.7996 | minutes
temperature = | 32.9
Cal done
Wed Mar 15 08:37:17 2023
Resistor SR1010_10 // P network measured by HP3458A with ID number 5
mains frequency = | 49.9581
temperature = | 32.9
Range?: | 10 | Item?: | SR1010_10 // P network
Trail1 | -1 | Trail2 | -1
Resistor SR1010_10 // P Network measured by HP3458A with ID number 5
mains frequency = | 50.0107
temperature = | 32.9
Range?: | e1 | Item?: | SR1010_10 // P Network
Resistor SR1010_10 // P Network measured by HP3458A with ID number 5
mains frequency = | 49.9808
temperature = | 32.9
Range?: | 10 | Item?: | SR1010_10 // P Network
Trail1 | 0.99999 | Trail2 | 1
at | Wed Mar 15 08:55:43 2023 | SR1010_10 // P Network | reading is:
Resistance_mean: | 1
Resistance_stdevp: | 1.0187e-05
Time taken = | 4.7431 | minutes
----------------
Resistor SR1010_100k // P network measured by HP3458A with ID number 5
mains frequency = | 50.0527
temperature = | 32.9
Range?: | 10000 | Item?: | SR1010_100k // P network
Trail1 | 10000.1282 | Trail2 | 10000.1308
at | Wed Mar 15 09:05:50 2023 | SR1010_100k // P network | reading is:
Resistance_mean: | 10000.1341
Resistance_stdevp: | 0.0018379
Time taken = | 4.7448 | minutes
----------------
end of the measurement
Resistor SR1010_100K S SN 937003 measured by HP3458A with ID number 5
mains frequency = | 49.9596
temperature = | 33
Range?: | 1000000 | Item?: | SR1010_100K S SN 937003
Trail1 | 1000015 | Trail2 | 1000014.877
at | Wed Mar 15 09:47:09 2023 | SR1010_100K S SN 937003 | reading is:
Resistance_mean: | 998424.044
Resistance_stdevp: | 15912.094
Time taken = | 3.5207 | minutes
----------------
Resistor SR104_10k SN 109006 measured by HP3458A with ID number 5
mains frequency = | 49.9277
temperature = | 33
Range?: | 10000 | Item?: | SR104_10k SN 109006
Trail1 | 10000.0087 | Trail2 | 10000.0108
at | Wed Mar 15 09:54:32 2023 | SR104_10k SN 109006 | reading is:
Resistance_mean: | 10000.0123
Resistance_stdevp: | 0.0025697
Time taken = | 4.7421 | minutes
----------------
Resistor SR104 T sensor measured by HP3458A with ID number 5
mains frequency = | 49.9904
temperature = | 33
Range?: | 10000 | Item?: | SR104 T sensor
Trail1 | 9989.3808 | Trail2 | 9989.3761
at | Wed Mar 15 10:01:26 2023 | SR104 T sensor | reading is:
Resistance_mean: | 9989.3076
Resistance_stdevp: | 0.039552
Time taken = | 4.7441 | minutes
----------------
Resistor SN1010_100k S measured by HP3458A with ID number 5
mains frequency = | 49.9644
temperature = | 33
Range?: | 1000000 | Item?: | SN1010_100k S
Trail1 | 1000015.386 | Trail2 | 1000014.961
at | Wed Mar 15 10:08:36 2023 | SN1010_100k S | reading is:
Resistance_mean: | 1000015.8203
Resistance_stdevp: | 0.54195
Time taken = | 3.5195 | minutes
----------------
Resistor Vishay 1M measured by HP3458A with ID number 5
mains frequency = | 50.0745
temperature = | 33
Range?: | 1000000 | Item?: | Vishay 1M
Trail1 | 1000002.511 | Trail2 | 1000002.78
at | Wed Mar 15 10:14:46 2023 | Vishay 1M | reading is:
Resistance_mean: | 1000001.7175
Resistance_stdevp: | 0.45496
Time taken = | 3.5199 | minutes
----------------
end of the measurement

## Comparision
Year |  | 2017 | 2019 | 2023 | delta19-17 | delta23-19 | Delta 23-17 | [5 empty] | 2023-02-20
DUT | sd ppm | [6 empty] | SR104 S/N:109006 |  | Sensor SR104 | 9981.0833095 | 0.004518600918942526 | Sensor SR104 | 9984.9603 | 0.002167
SR104 |  | 0.098999 | 0.12495 | 0.17046 | 0.26214 | 0.36425 | 0.72187 | α23 = +0.18 |  | SR104 | 10000.017241500002 | 0.0017046292234726512 | SR104 | 10000.0148 | 0.0018672
100k esi parallel |  | 0.52249 | 0.11113 | 0.52696 | -0.78731 | 3.7419 | 0.0085637 | β = -0.027
100k esi series-para |  | 0.080698 | 0.072484 | 0.27798 | -0.10179 | 2.8351 | 2.4447
100k esi R10 |  | 0.07234 | 0.083189 | 0.116 | 0.14998 | 0.39447 | 0.60361
100k esi series |  | 0.54059 | 0.090137 | 0.59573 | -0.83326 | 5.6091 | 0.10199 |  |  | temp by the SR 104 box formula
G1 |  | 0.096892 | 0.14669 | 0.13348 | 0.51393 | -0.090051 | 0.3776 | Alpha_23 | 0.18 | 21.1066 |  |  | 21.4946
G2 |  | 0.071488 | 0.12419 | 0.093481 | 0.73727 | -0.2473 | 0.30765 | Beta_23 | -0.027
R4A |  | 0.11869 | 0.12082 | 0.66375 | 0.017957 | 4.4938 | 4.5924 | T | 21
R4B |  | 0.072539 | 0.095025 | 0.16064 | 0.30998 | 0.69053 | 1.2146 | Delta T | -2 | temp by E009p23 formula
R4C |  | 0.16396 | 0.37705 | 1.7798 | 1.2997 | 3.7203 | 9.8554 | A_T_numerator | 0.288
[8 empty] | denominator | 0.532
[8 empty] | Alpha_21 | 0.54135
[8 empty] | Beta_21 | -0.050752
[8 empty] | delta | 0.54164
[8 empty] | 2a | -0.1015
[8 empty] | + | -0.0027802
[8 empty] | - | 10.6694
Calculation of temp by the Temp sensors
From E009-011 P23

## DumpWord
Calibration Log
First quick check on the meter, serial no.2823A21245 (watt bridge meter).
C:\Python27\python.exe C:/Users/k.jones/PycharmProjects/UbridgeCalibration/ACALall3458.py
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
mains frequency = 49.9897804E+00
Hz
Temperature = 35.9
ACAL ALL starting
ACAL finished
Time taken = 16.2833166639 minutes
Temperature = 35.9
Wed May 31 14:29:40 2017
Wait at least 5 minutes before using the meter after doing an ACAL
Process finished with exit code 0
2017-06-01
>pythonw -u "ACALall3458.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
mains frequency = 50.0444640E+00
Hz
Temperature = 35.7
ACAL ALL starting
ACAL finished
Time taken = 16.021266667 minutes
Temperature = 35.7
Thu Jun 01 11:48:26 2017
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.7
Range? :1e4
Item? SR104 thermometer, GMH5 20.58 (in well)
Trial 9.976265574E+03
Trial 2 9.976267493E+03
0 9.976264332E+03
1 9.976267237E+03
2 9.976266812E+03
3 9.976266992E+03
4 9.976267074E+03
5 9.976267454E+03
6 9.976268215E+03
7 9.976266918E+03
8 9.976265816E+03
9 9.976270492E+03
10 9.976265893E+03
11 9.976270749E+03
12 9.976266889E+03
13 9.976267363E+03
14 9.976270895E+03
15 9.976267596E+03
16 9.976273635E+03
17 9.976269898E+03
18 9.976269092E+03
19 9.976268622E+03
Result
Thu Jun 01 11:57:52 2017
9976.2681
0.0021611
>Exit code: 0
Next SR104 itself
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.7
Range? :1e4
Item? SR104 serial 109006, GMH5 reads 20.58
Trial 9.999996623E+03
Trial 2 9.999995857E+03
0 9.999996665E+03
1 9.999995950E+03
2 9.999996199E+03
3 9.999996232E+03
4 9.999996653E+03
5 9.999997037E+03
6 9.999997722E+03
7 9.999996756E+03
8 9.999998462E+03
9 9.999994700E+03
10 9.999997428E+03
11 9.999996476E+03
12 9.999997922E+03
13 9.999997527E+03
14 9.999995886E+03
15 9.999995764E+03
16 9.999998147E+03
17 9.999998328E+03
18 9.999995937E+03
19 9.999996673E+03
Result
Thu Jun 01 12:04:33 2017
9999.9968
0.00098999
>Exit code: 0
Then 100k ESI in parallel
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e4
Item? 100k ESI in parallel as 10k, GMH5 20.7 on bench by box
Trial 1.000009391E+04
Trial 2 1.000009504E+04
0 1.000009390E+04
1 1.000009231E+04
2 1.000009262E+04
3 1.000009478E+04
4 1.000009406E+04
5 1.000009150E+04
6 1.000009222E+04
7 1.000009278E+04
8 1.000009514E+04
9 1.000009416E+04
10 1.000009110E+04
11 1.000009023E+04
12 1.000009213E+04
13 1.000009521E+04
14 1.000009206E+04
15 1.000009362E+04
16 1.000009306E+04
17 1.000009370E+04
18 1.000009321E+04
19 1.000009213E+04
Result
Thu Jun 01 12:11:48 2017
10000.093
0.0013385
>Exit code: 0
Repeat after lunch
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e4
Item? ESI 100k 1-10 in parallel as 10k, GMH5 reads 20.59 on bench
Trial 1.000010336E+04
Trial 2 1.000008327E+04
0 1.000008892E+04
1 1.000009666E+04
2 1.000009674E+04
3 1.000008282E+04
4 1.000009640E+04
5 1.000008502E+04
6 1.000008656E+04
7 1.000008480E+04
8 1.000009614E+04
9 1.000008601E+04
10 1.000009211E+04
11 1.000009476E+04
12 1.000009571E+04
13 1.000008632E+04
14 1.000008842E+04
15 1.000009068E+04
16 1.000009869E+04
17 1.000009544E+04
18 1.000008573E+04
19 1.000009704E+04
Result
Thu Jun 01 12:55:51 2017
10000.0912
0.0052249
>Exit code: 0
Next 100k S-P
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e5
Item? ESI 100 k 1-9 in SP, GMH5 21.08 on bench
Trial 1.000006986E+05
Trial 2 1.000007199E+05
0 1.000007041E+05
1 1.000006963E+05
2 1.000007012E+05
3 1.000007099E+05
4 1.000006864E+05
5 1.000007064E+05
6 1.000007070E+05
7 1.000006901E+05
8 1.000006904E+05
9 1.000007013E+05
10 1.000007061E+05
11 1.000007121E+05
12 1.000007082E+05
13 1.000006952E+05
14 1.000007151E+05
15 1.000007000E+05
16 1.000007020E+05
17 1.000007087E+05
18 1.000007064E+05
19 1.000006906E+05
Result
Thu Jun 01 13:10:18 2017
100000.7019
0.0080699
>Exit code: 0
Now the 10th element
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e5
Item? R10 of ESI 100k, GMH5 20.99
Trial 1.000004829E+05
Trial 2 1.000004928E+05
0 1.000004875E+05
1 1.000004837E+05
2 1.000004916E+05
3 1.000004988E+05
4 1.000005007E+05
5 1.000004828E+05
6 1.000004854E+05
7 1.000004778E+05
8 1.000004885E+05
9 1.000004765E+05
10 1.000004819E+05
11 1.000004808E+05
12 1.000004842E+05
13 1.000004819E+05
14 1.000004875E+05
15 1.000004837E+05
16 1.000005043E+05
17 1.000004875E+05
18 1.000004884E+05
19 1.000004897E+05
Result
Thu Jun 01 13:16:06 2017
100000.4872
0.007234
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? repeat R10, GMH5 20.99 (error beep on run above)
Trial 1.000004581E+05
Trial 2 1.000004790E+05
0 1.000004715E+05
1 1.000004664E+05
2 1.000004589E+05
3 1.000004669E+05
4 1.000004525E+05
5 1.000004504E+05
6 1.000004390E+05
7 1.000004528E+05
8 1.000004556E+05
9 1.000004496E+05
10 1.000004464E+05
11 1.000004645E+05
12 1.000004511E+05
13 1.000004606E+05
14 1.000004480E+05
15 1.000004402E+05
16 1.000004404E+05
17 1.000004487E+05
18 1.000004500E+05
19 1.000004415E+05
Result
Thu Jun 01 13:21:57 2017
100000.4527
0.0094992
>Exit code: 0
Finally, for now, 1 M ohm, 1-10 ESI in series
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e6
Item? ESI 100 k 1-10 in series, GMH5 reads 20.98
Trial 1.000007842E+06
Trial 2 1.000007476E+06
0 1.000008252E+06
1 1.000007469E+06
2 1.000007949E+06
3 1.000007737E+06
4 1.000007814E+06
5 1.000007766E+06
6 1.000008099E+06
7 1.000007901E+06
8 1.000008295E+06
9 1.000007957E+06
10 1.000008156E+06
11 1.000008696E+06
12 1.000008525E+06
13 1.000008879E+06
14 1.000008833E+06
15 1.000009030E+06
16 1.000009032E+06
17 1.000009416E+06
18 1.000008790E+06
19 1.000008830E+06
Result
Thu Jun 01 13:27:36 2017
1000008.3713
0.5406
>Exit code: 0
Now the internal UB standards.
R4B
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? R4B, GMH5 reads 20.31 in bridge
Trial 9.999888028E+04
Trial 2 9.999888658E+04
0 9.999888912E+04
1 9.999889498E+04
2 9.999890208E+04
3 9.999889466E+04
4 9.999889424E+04
5 9.999889108E+04
6 9.999889148E+04
7 9.999889942E+04
8 9.999889246E+04
9 9.999888631E+04
10 9.999889514E+04
11 9.999888229E+04
12 9.999887335E+04
13 9.999889276E+04
14 9.999889298E+04
15 9.999889206E+04
16 9.999889492E+04
17 9.999888841E+04
18 9.999890439E+04
19 9.999890489E+04
Result
Thu Jun 01 13:41:59 2017
99998.8929
0.0072538
>Exit code: 0
Now R4A
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e4
Item? R4A 10k, GMH5 reads 20.44 in ubridge
Trial 1.000054051E+04
Trial 2 1.000053927E+04
0 1.000053910E+04
1 1.000053927E+04
2 1.000053542E+04
3 1.000053705E+04
4 1.000053919E+04
5 1.000053914E+04
6 1.000053741E+04
7 1.000053847E+04
8 1.000053943E+04
9 1.000053832E+04
10 1.000053799E+04
11 1.000053992E+04
12 1.000053957E+04
13 1.000053948E+04
14 1.000053802E+04
15 1.000053877E+04
16 1.000054049E+04
17 1.000053786E+04
18 1.000053763E+04
19 1.000053981E+04
Result
Thu Jun 01 13:53:03 2017
10000.5386
0.0011869
>Exit code: 0
Next R4C with right angle adapters and C joined to P in a pi adapter
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e6
Item? R4C, GMH5 reads 20.52
Trial 1.000064184E+06
Trial 2 1.000065322E+06
0 1.000065078E+06
1 1.000065417E+06
2 1.000065590E+06
3 1.000065629E+06
4 1.000065491E+06
5 1.000065461E+06
6 1.000065435E+06
7 1.000065456E+06
8 1.000065535E+06
9 1.000065511E+06
10 1.000065681E+06
11 1.000065569E+06
12 1.000065410E+06
13 1.000065473E+06
14 1.000065586E+06
15 1.000065439E+06
16 1.000065513E+06
17 1.000065578E+06
18 1.000065945E+06
19 1.000065324E+06
Result
Thu Jun 01 14:00:52 2017
1000065.506
0.16397
>Exit code: 0
Then we come to R2 (lower position)
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? R2, GMH5 reads 20.54
Trial 9.999889488E+04
Trial 2 9.999889159E+04
0 9.999888341E+04
1 9.999888759E+04
2 9.999887405E+04
3 9.999888074E+04
4 9.999887237E+04
5 9.999888553E+04
6 9.999887841E+04
7 9.999887725E+04
8 9.999888686E+04
9 9.999887619E+04
10 9.999886073E+04
11 9.999887784E+04
12 9.999888020E+04
13 9.999886656E+04
14 9.999887949E+04
15 9.999886930E+04
16 9.999887298E+04
17 9.999887200E+04
18 9.999888101E+04
19 9.999886705E+04
Result
Thu Jun 01 14:07:01 2017
99998.8765
0.0071487
>Exit code: 0
And, finally G1, upper position
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? G1, with GMH5 reading 20.51
Trial 9.999836176E+04
Trial 2 9.999834648E+04
0 9.999834301E+04
1 9.999835466E+04
2 9.999835246E+04
3 9.999834782E+04
4 9.999833479E+04
5 9.999832900E+04
6 9.999834780E+04
7 9.999833382E+04
8 9.999834182E+04
9 9.999833914E+04
10 9.999834125E+04
11 9.999832407E+04
12 9.999831739E+04
13 9.999833160E+04
14 9.999832869E+04
15 9.999832822E+04
16 9.999834164E+04
17 9.999834068E+04
18 9.999833286E+04
19 9.999832933E+04
Result
Thu Jun 01 14:11:19 2017
99998.337
0.009689
>Exit code: 0

## DumpScite
>pythonw -u "ACALall3458.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
mains frequency = 50.0444640E+00
Hz
Temperature = 35.7
ACAL ALL starting
ACAL finished
Time taken = 16.021266667 minutes
Temperature = 35.7
Thu Jun 01 11:48:26 2017
Wait at least 5 minutes before using the meter after doing an ACAL
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.7
Range? :1e4
Item? SR104 thermometer
Trial 9.976265574E+03
Trial 2 9.976267493E+03
0 9.976264332E+03
1 9.976267237E+03
2 9.976266812E+03
3 9.976266992E+03
4 9.976267074E+03
5 9.976267454E+03
6 9.976268215E+03
7 9.976266918E+03
8 9.976265816E+03
9 9.976270492E+03
10 9.976265893E+03
11 9.976270749E+03
12 9.976266889E+03
13 9.976267363E+03
14 9.976270895E+03
15 9.976267596E+03
16 9.976273635E+03
17 9.976269898E+03
18 9.976269092E+03
19 9.976268622E+03
Result
Thu Jun 01 11:57:52 2017
9976.2681
0.0021611
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.7
Range? :1e4
Item? SR104 serial 109006, GMH5 reads 20.58
Trial 9.999996623E+03
Trial 2 9.999995857E+03
0 9.999996665E+03
1 9.999995950E+03
2 9.999996199E+03
3 9.999996232E+03
4 9.999996653E+03
5 9.999997037E+03
6 9.999997722E+03
7 9.999996756E+03
8 9.999998462E+03
9 9.999994700E+03
10 9.999997428E+03
11 9.999996476E+03
12 9.999997922E+03
13 9.999997527E+03
14 9.999995886E+03
15 9.999995764E+03
16 9.999998147E+03
17 9.999998328E+03
18 9.999995937E+03
19 9.999996673E+03
Result
Thu Jun 01 12:04:33 2017
9999.9968
0.00098999
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e4
Item? 100k ESI in parallel as 10k, GMH5 20.7 on bench by box
Trial 1.000009391E+04
Trial 2 1.000009504E+04
0 1.000009390E+04
1 1.000009231E+04
2 1.000009262E+04
3 1.000009478E+04
4 1.000009406E+04
5 1.000009150E+04
6 1.000009222E+04
7 1.000009278E+04
8 1.000009514E+04
9 1.000009416E+04
10 1.000009110E+04
11 1.000009023E+04
12 1.000009213E+04
13 1.000009521E+04
14 1.000009206E+04
15 1.000009362E+04
16 1.000009306E+04
17 1.000009370E+04
18 1.000009321E+04
19 1.000009213E+04
Result
Thu Jun 01 12:11:48 2017
10000.093
0.0013385
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e4
Item? ESI 100k 1-10 in parallel as 10k, GMH5 reads 20.59 on bench
Trial 1.000010336E+04
Trial 2 1.000008327E+04
0 1.000008892E+04
1 1.000009666E+04
2 1.000009674E+04
3 1.000008282E+04
4 1.000009640E+04
5 1.000008502E+04
6 1.000008656E+04
7 1.000008480E+04
8 1.000009614E+04
9 1.000008601E+04
10 1.000009211E+04
11 1.000009476E+04
12 1.000009571E+04
13 1.000008632E+04
14 1.000008842E+04
15 1.000009068E+04
16 1.000009869E+04
17 1.000009544E+04
18 1.000008573E+04
19 1.000009704E+04
Result
Thu Jun 01 12:55:51 2017
10000.0912
0.0052249
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e5
Item? ESI 100 k 1-9 in SP, GMH5 21.08 on bench
Trial 1.000006986E+05
Trial 2 1.000007199E+05
0 1.000007041E+05
1 1.000006963E+05
2 1.000007012E+05
3 1.000007099E+05
4 1.000006864E+05
5 1.000007064E+05
6 1.000007070E+05
7 1.000006901E+05
8 1.000006904E+05
9 1.000007013E+05
10 1.000007061E+05
11 1.000007121E+05
12 1.000007082E+05
13 1.000006952E+05
14 1.000007151E+05
15 1.000007000E+05
16 1.000007020E+05
17 1.000007087E+05
18 1.000007064E+05
19 1.000006906E+05
Result
Thu Jun 01 13:10:18 2017
100000.7019
0.0080699
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1e5
Item? R10 of ESI 100k, GMH5 20.99
Trial 1.000004829E+05
Trial 2 1.000004928E+05
0 1.000004875E+05
1 1.000004837E+05
2 1.000004916E+05
3 1.000004988E+05
4 1.000005007E+05
5 1.000004828E+05
6 1.000004854E+05
7 1.000004778E+05
8 1.000004885E+05
9 1.000004765E+05
10 1.000004819E+05
11 1.000004808E+05
12 1.000004842E+05
13 1.000004819E+05
14 1.000004875E+05
15 1.000004837E+05
16 1.000005043E+05
17 1.000004875E+05
18 1.000004884E+05
19 1.000004897E+05
Result
Thu Jun 01 13:16:06 2017
100000.4872
0.007234
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? repeat R10, GMH5 20.99 (error beep on run above)
Trial 1.000004581E+05
Trial 2 1.000004790E+05
0 1.000004715E+05
1 1.000004664E+05
2 1.000004589E+05
3 1.000004669E+05
4 1.000004525E+05
5 1.000004504E+05
6 1.000004390E+05
7 1.000004528E+05
8 1.000004556E+05
9 1.000004496E+05
10 1.000004464E+05
11 1.000004645E+05
12 1.000004511E+05
13 1.000004606E+05
14 1.000004480E+05
15 1.000004402E+05
16 1.000004404E+05
17 1.000004487E+05
18 1.000004500E+05
19 1.000004415E+05
Result
Thu Jun 01 13:21:57 2017
100000.4527
0.0094992
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e6
Item? ESI 100 k 1-10 in series, GMH5 reads 20.98
Trial 1.000007842E+06
Trial 2 1.000007476E+06
0 1.000008252E+06
1 1.000007469E+06
2 1.000007949E+06
3 1.000007737E+06
4 1.000007814E+06
5 1.000007766E+06
6 1.000008099E+06
7 1.000007901E+06
8 1.000008295E+06
9 1.000007957E+06
10 1.000008156E+06
11 1.000008696E+06
12 1.000008525E+06
13 1.000008879E+06
14 1.000008833E+06
15 1.000009030E+06
16 1.000009032E+06
17 1.000009416E+06
18 1.000008790E+06
19 1.000008830E+06
Result
Thu Jun 01 13:27:36 2017
1000008.3713
0.5406
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? R4B, GMH5 reads 20.31 in bridge
Trial 9.999888028E+04
Trial 2 9.999888658E+04
0 9.999888912E+04
1 9.999889498E+04
2 9.999890208E+04
3 9.999889466E+04
4 9.999889424E+04
5 9.999889108E+04
6 9.999889148E+04
7 9.999889942E+04
8 9.999889246E+04
9 9.999888631E+04
10 9.999889514E+04
11 9.999888229E+04
12 9.999887335E+04
13 9.999889276E+04
14 9.999889298E+04
15 9.999889206E+04
16 9.999889492E+04
17 9.999888841E+04
18 9.999890439E+04
19 9.999890489E+04
Result
Thu Jun 01 13:41:59 2017
99998.8929
0.0072538
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e4
Item? R4A 10k, GMH5 reads 20.44 in ubridge
Trial 1.000054051E+04
Trial 2 1.000053927E+04
0 1.000053910E+04
1 1.000053927E+04
2 1.000053542E+04
3 1.000053705E+04
4 1.000053919E+04
5 1.000053914E+04
6 1.000053741E+04
7 1.000053847E+04
8 1.000053943E+04
9 1.000053832E+04
10 1.000053799E+04
11 1.000053992E+04
12 1.000053957E+04
13 1.000053948E+04
14 1.000053802E+04
15 1.000053877E+04
16 1.000054049E+04
17 1.000053786E+04
18 1.000053763E+04
19 1.000053981E+04
Result
Thu Jun 01 13:53:03 2017
10000.5386
0.0011869
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e6
Item? R4C, GMH5 reads 20.52
Trial 1.000064184E+06
Trial 2 1.000065322E+06
0 1.000065078E+06
1 1.000065417E+06
2 1.000065590E+06
3 1.000065629E+06
4 1.000065491E+06
5 1.000065461E+06
6 1.000065435E+06
7 1.000065456E+06
8 1.000065535E+06
9 1.000065511E+06
10 1.000065681E+06
11 1.000065569E+06
12 1.000065410E+06
13 1.000065473E+06
14 1.000065586E+06
15 1.000065439E+06
16 1.000065513E+06
17 1.000065578E+06
18 1.000065945E+06
19 1.000065324E+06
Result
Thu Jun 01 14:00:52 2017
1000065.506
0.16397
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? R2, GMH5 reads 20.54
Trial 9.999889488E+04
Trial 2 9.999889159E+04
0 9.999888341E+04
1 9.999888759E+04
2 9.999887405E+04
3 9.999888074E+04
4 9.999887237E+04
5 9.999888553E+04
6 9.999887841E+04
7 9.999887725E+04
8 9.999888686E+04
9 9.999887619E+04
10 9.999886073E+04
11 9.999887784E+04
12 9.999888020E+04
13 9.999886656E+04
14 9.999887949E+04
15 9.999886930E+04
16 9.999887298E+04
17 9.999887200E+04
18 9.999888101E+04
19 9.999886705E+04
Result
Thu Jun 01 14:07:01 2017
99998.8765
0.0071487
>Exit code: 0
>pythonw -u "MeasureR.py"
(u'ASRL3::INSTR', u'ASRL4::INSTR', u'ASRL5::INSTR', u'ASRL6::INSTR', u'ASRL7::INSTR', u'ASRL10::INSTR', u'GPIB0::22::INSTR')
HP3458A
Temperature 35.8
Range? :1.2e5
Item? G1, with GMH5 reading 20.51
Trial 9.999836176E+04
Trial 2 9.999834648E+04
0 9.999834301E+04
1 9.999835466E+04
2 9.999835246E+04
3 9.999834782E+04
4 9.999833479E+04
5 9.999832900E+04
6 9.999834780E+04
7 9.999833382E+04
8 9.999834182E+04
9 9.999833914E+04
10 9.999834125E+04
11 9.999832407E+04
12 9.999831739E+04
13 9.999833160E+04
14 9.999832869E+04
15 9.999832822E+04
16 9.999834164E+04
17 9.999834068E+04
18 9.999833286E+04
19 9.999832933E+04
Result
Thu Jun 01 14:11:19 2017
99998.337
0.009689
>Exit code: 0